## Collected Ge Detector Sample Re
Some Useful Information Concerning the results:
The Conversion factors for the primordial nuclides are given by: | [9 empty] | 1 Bq 238U/kg = 81 ppb U (81 x 10-9 gU/g) | [5 empty] | The 238U decay chain gammas used are: | [5 empty] | 234Th: 63.29 and 92.59 keV 226Ra: 186.1 keV
214Pb: 295.21 and 351.92 keV
214Bi: 609.31, 1120.29, 1764.49 and 2204.21 keV
The relationships are valid for any daughters in the 238U, 235U or 232Th chain only if the chain is in equilibrium. | [9 empty] | 1 Bq 232Th/kg = 246 ppb Th (246 x 10-9 gTh/g) | [5 empty] | The 232Th decay chain gammas used are: | [5 empty] | 212Pb: 238.63 and 300.09 keV
208Tl: 277.371, 583.19, 860.557 and 2614.53 keV,
228Ac: 209.253, 338.320, 463,004, 911.21, 964.766 and 968.971 keV
[10 empty] | 1 Bq 40K/kg = 32300 ppb K (32300 x 10-6 gK/g) | [5 empty] | The 40K decay chain gamma used is: | [5 empty] | 40K: 1460.83 keV
[10 empty] | 1 Bq 235U/kg = 1.76 ppm U (1.76 x 10-6 gU/g) | [5 empty] | The 235U decay chain gammas used are: | [5 empty] | 235U: 143.76, 163.33 and 205.31 keV
Background Measurements:
If a measurement in the signal region is below the sideband regions then the 90% confidence limit is calculated.
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
Background CW1 |  |  | 38.281 | 121121 121123 121130 121229 | 2012-11-21 | (mBq) | 0.153 | +/- | 0.053 | 0.1 | +- | 0.071 | 0.345 | +- | 0.065 | 0.216 | +/- | 0.097 | 0.073 | +/- | 0.278 | <0.269 |  |  | 0.151 | +/- | 0.276
Completely Empty Detector |  |  |  |  | 2012-12-31 | [19 empty] | 1.9823 |  | 13.1275
Background CW2 |  |  | 23.184 | 130328 130330 130406 130419 | 2013-03-28 | (mBq) | 0.225 | +/- | 0.063 | 0.02 | +- | 0.089 | 0.227 | +- | 0.066 | 0.308 | +/- | 0.098 | 0.759 | +/- | 0.378 | 0.7 | +/- | 0.7 | 0.195 | +/- | 0.332
Completely Empty Detector |  |  |  |  | 2013-04-22 | [19 empty] | 1.7691 |  | 9.0724
Background CW3 |  |  | 10.681 | 130722 130727 130730 130802 | 2013-07-22 | (mBq) | 0.187 | +/- | 0.078 | 0.058 | +- | 0.196 | 0.116 | +- | 0.038 | 0.218 | +/- | 0.099 | <0.518 |  |  | <1.508 |  |  | <0.172
Completely Empty Detector |  |  |  |  | 2013-08-13 | [19 empty] | 0 |  | 33.8021
Background CW4 |  |  | 14.158 | 130929 131111 13111101 | 2013-10-29 | (mBq) | 0.08 | +- | 0.078 | 0.1 | +- | 0.089 | 0.065 | +- | 0.032 | 0.028 | +- | 0.099 | <0.201 |  |  | <1.195 |  |  | <0.391
Completely Empty Detector |  |  |  |  | 2013-11-13 | [19 empty] | 0 |  | 6.541
Background CW5 |  |  | 6.721 | 140814 | 2014-08-18 | (mBq) | 0.08 | +/- | 0.083 | <0.206 |  |  | 0.012 | +- | 0.086 | <0.250 |  |  | 0.186 | +- | 0.759 | <3.159 |  |  | 1.5 | +- | 0.7
Completely Empty Detector |  |  |  |  | 2014-08-25 | [19 empty] | 3.0612 |  | 2.0408
Background CW6 |  |  | 44.129 | 141215 150107 150116 15011601 | 2014-12-15 | (mBq) | 0.178 | +- | 0.039 | 0.1 | +- | 0.059 | 0.051 | +- | 0.018 | 0.164 | +- | 0.054 | 0.23 | +- | 0.245 | <0.23 |  |  | 0.921 | +- | 0.21
Completely Empty Detector |  |  |  |  | 2015-01-29 | [19 empty] | 20.8844 |  | 22.6757
Background CW7 |  |  | 12.386 | 150224 150227 15022701 050301 150308 150309 | 2015-02-24 | (mBq) | 0.132 | +- | 0.07 | <0.100 |  |  | 0.012 | +- | 0.033 | 0.379 | +- | 0.099 | <0.364 |  |  | <2.19 |  |  | 0.65 | +- | 0.37
Completely Empty Detector |  |  |  |  | 2015-03-10 | [19 empty] | 4.748 |  | 7.3046
Background CW8 |  |  | 25.147 | 151218 151219 15121901 151224 151231 15123101 160101 | 2015-12-18 | (mBq) | 0.209 | +- | 0.055 | <0.011 |  |  | <0.021 |  |  | 0.294 | +- | 0.078 | 0.838 | +- | 0.361 | <0.735 |  |  | 0.62 | +- | 0.29
Completely Empty Detector |  |  |  |  | 2016-01-13 | [19 empty] | 7.3722 |  | 11.8906
Background CW9 |  |  | 25.098 | 160608 160819 16062801 | 2016-06-08 | (mBq) | 0.162 | +- | 0.051 | 0.011 | +- | 0.073 | <0.008 |  |  | 0.314 | +- | 0.082 | <0.280 |  |  | 0.688 | +- | 0.844 | 0.74 | +- | 0.25
Completely Empty Detector |  |  |  |  | 2016-07-07 | [19 empty] | 11.84 |  | 16
Background CW10 |  |  | 52.569 | 161202 161213 161215 161229 170111 | 2016-12-02 | (mBq) | 0.123 | +- | 0.035 | 0.03 | +- | 0.051 | 0.044 | +- | 0.015 | 0.211 | +- | 0.054 | 0.092 | +- | 0.22 | <0.187 |  |  | 0.47 | +- | 0.18
Completely Empty Detector |  |  |  |  | 2017-01-25 | [19 empty] | 14.5062 |  | 30.8642
[6 empty] | Additional Backgrounds: |  | 7Be: |  |  | 54Mn:
[6 empty] | (mBq) | <0.198 |  |  | 2.5 | +- | 2.64
Background CW11 |  |  | 20.398 | 180925 180930 181005 181008 | 2018-09-25 | (mBq) | 0.172 | +- | 0.058 | 0.097 | +- | 0.092 | 0.035 | +- | 0.025 | 0.188 | +- | 0.082 | 0.189 | +- | 0.331 | <1.33 |  |  | 1.38 | +- | 0.43
Completely Empty Detector |  |  |  |  | 2018-10-17
[6 empty] | Additional Backgrounds: |  | 7Be: |  |  | 54Mn:
[6 empty] | (mBq) | 1.34 | +- | 1.73 | <4.61
Background CW12 |  |  | 10.726 | 181102 | 2018-11-02 | (mBq) | 0.077 | +- | 0.08 | 0.241 | +- | 0.126 | <0.022 |  |  | 0.124 | +- | 0.11 | 0.712 | +- | 0.559 | <1.72 |  |  | 0.84 | +- | 0.56
Completely Empty Detector |  |  |  |  | 2018-11-13
[6 empty] | Additional Backgrounds: |  | 7Be: |  |  | 54Mn:
[6 empty] | (mBq) | <0.34 |  |  | <12.86
Background CW13 |  |  | 53.002 | 200324 200430 | 2020-03-24 | (mBq) | 1.105 | +- | 0.145 | 0.1276 | +- | 0.07508 | 0.02407 | +- | 0.0093 | 0.3456 | +- | 0.06701 | 0.4596 | +- | 0.1895 | <0.57 |  |  | 0.424 | +- | 0.203
Completely Empty Detector |  |  |  |  | 2020-05-19
[6 empty] | Additional Backgrounds: |  | 7Be: |  |  | 54Mn:
[6 empty] | (mBq) | 1.6297 | +- | 1.081 | 0 | +- | 0.6657
Background CW14 |  |  | 40.956 | 200811 20081101 | 2020-08-11 | (mBq) | 0.1683 | +- | 0.1044 | 0.1355 | +- | 0.7493 | <0.011 |  |  | 0.2076 | +- | 0.06712 | 0.1891 | +- | 0.1831 | <0.87 |  |  | 0.43679 | +- | 0.2013
Completely Empty Detector |  |  |  |  | 2020-09-02
[6 empty] | Additional Backgrounds: |  | 7Be: |  |  | 54Mn: |  |  | 60Co
[6 empty] | (mBq) | <0.75 |  |  | <1.22 |  |  | 1.068 | +- | 2.145
Combined Background | Combined Backgrounds of runs CW1, CW2, CW3, CW4, CW5, CW6, CW7, CW8, CW9, CW10, CW11 & CW12 |  | 283.566 |  |  | (mBq) | 0.23 | +- | 0.04 | 0.044 | +- | 0.027 | 0.023 | +- | 0.008 | 0.752 | +- | 0.216 | 0.249 | +- | 0.108 | <0.024 |  |  | 0.221 | +- | 0.072
Completely Empty Detector
[6 empty] | Additional Backgrounds: |  | 7Be: |  |  | 54Mn:
[6 empty] | (mBq) | <0.54 | . |  | 0.785 | +- | 0.97
Background CW15 | This is the first background run with the Fast Discriminator set to Low Energy | N/A | 33.032 | 210322 210405 | 2021-03-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2021-04-26 | (mBq) | 0.2586 | +- | 0.03125 | 0.5146 | +- | 0.1591 | 0.007358 | +- | 0.005664 | 0.2039 | +- | 0.02909 | 4.161 | +- | 1.116 | 0.043215 | +- | 0.02767 | 0.2246 | +- | 0.06323
RT: 18.4 uSec FT: 1.2 uSec Pole Zero: 2048
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (Mbq) | <1.01 |  |  | <0.17 |  |  | 0.0314 | +- | 0.03578 | 0.2947 | +- | 0.07995
Background CW16 | This is the second background run with the Fast Discriminator set to Low Energy | N/A | 40.479 | 210729 21080501 210812 21081201 | 2021-07-29 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  | 210823 210901 21090101 | 2021-09-10 | (mBq) | 0.02235 | +- | 0.02191 | 1.634 | +- | 0.2491 | <0.0046 |  |  | 0.0504 | +- | 0.0218 | 0.79954 | +- | 1.442 | 0.033305 | +- | 0.02507 | 0.09359 | +- | 0.0582
RT: 18.4 uSec FT: 1.2 uSec Pole Zero: 2048
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (Mbq) | <2.21 |  |  | <0.30 |  |  | <0.032 |  |  | 0.0972 | +- | 0.06547
Background CW17 | This is the third background run with the Fast Discriminator set to Low Energy | N/A | 20.798 | 220114 | 2022-01-14 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2022-02-04 | (mBq) | 0.1141 | +- | 0.03294 | <0.37 |  |  | 0.01238 | +- | 0.00752 | 0.1188 | +- | 0.03279 | 3.8904 | +- | 1.495 | <0.068 |  |  | 0.2056 | +- | 0.07735
RT: 18.4 uSec FT: 1.2 uSec Pole Zero: 2048
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (Mbq) | <1.31 |  |  | 0.063238 | +- | 0.1747 | 0.066494 | +- | 0.04714 | 0.4168 | +- | 0.09282
Background CW18 | This is the forth background run with the Fast Discriminator set to Low Energy | N/A | 103.809 | 220805 220807 220809 220811 220813 |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  | 220815 22081501 |  | (mBq) | 0.1225 | +- | 0.0156 | 0.2433 | +- | 0.07917 | 0.004414 | +- | 0.003372 | 0.1945 | +- | 0.01689 | 7.5605 | +- | 0.9157 | <0.023 |  |  | 0.1334 | +- | 0.02997
| RT: 18.4 uSec FT: 1.2 uSec Pole Zero: 2048 |  |  | 220816 220817 220825 220906
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  | 220919 22091901 220920 |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
|  |  |  | 221014 221021 221024 221025 22102701 |  | (Mbq) | <0.55 |  |  | <0.15 |  |  | <0.033 |  |  | 0.2293 | +- | 0.04286
22110401 221219 230117
Background CW19 |  | N/A | 63.332 | 230419 230613 230614 230620 |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector | [5 empty] | (mBq) | 0.1527 | +- | 0.02125 | 0.1536 | +- | 0.1044 | 0.007561 | +- | 0.00476 | 0.2039 | +- | 0.0227 | 6.8798 | +- | 1.127 | 0.030535 | +- | 0.02649 | 0.1709 | +- | 0.04222
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (Mbq) | <0.65 |  |  | 0.22683 | +- | 0.1124 | <0.041 |  |  | 0.2281 | +- | 0.05763
The measurements of the samples below take into account the background measurements from the background table. If a measurement is below the background then the upper bound shown is the 90% confidence limit.
SNOLAB Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
SNOLAB CW01 |  |  | 17.199 | 130214 130508 13051601 | Feb 14, 2013 –- Feb 19, 2013 | (mBq) | 0.281 | +- | 0.168 | 0.047 | +- | 0.133 | 0.039 | +- | 0.036 | 1.833 | +- | 0.974 | <0.183 |  |  | <1.815 |  |  | 0.598 | +- | 0.76
Empty Cowie Bottle Background |  |  |  |  | May 8, 2013 –- May 21, 2013
KCM CW01a | Sample #1a | 15 g (total) | 7.741 | 180131 | 2018-01-31 | (Bq/kg) | 81.231 | +- | 2.55 | 17.802 | +- | 1.738 | 2.6647 | +- | 2.7887 | 61.013 | +- | 2.75 | 142.25 | +- | 4.57 | 3.25 | +- | 2.72 | <0.65
Water Sample from Kidd Creek Mine Level 7850 | 29.01.2018_KC7850_12299_SLRS | 7.5 g (inside well) |  |  | 2018-02-08 | (ppb or ppm) | 6.58 ppm | +- | 0.21 ppm | 1.44 ppm | +- | 0.14 ppm | 4.69 ppm | +- | 0.49 ppm | 15.01 ppm | +- | 0.68 ppm | 34.99 ppm | +- | 1.12 ppm |  |  |  | <53.03 ppb
KCM CW01b | Sample #1b | 13.5 g | 6.935 | 180215 18021501 | 2018-02-15 | (Bq/kg) | 111.85 | +- | 3.35 | 31.13 | +- | 2.83 | 2.41 | +- | 0.28 | 10.19 | +- | 0.61 | 137.92 | +- | 4.51 | <4.03 |  |  | <1.36
Water Sample from Kidd Creek Mine Level 7850 | 29.01.2018_KC7850_12299_SLRS | 7.5 g (inside well) |  |  | 2018-02-22 | (ppb or ppm) | 9.06 ppm | +- | 0.27 ppm | 2.52 ppm | +- | 0.23 ppm | 4.23 ppm | +- | 0.49 ppm | 2.51 ppm | +- | 0.15 ppm | 33.93 ppm | +- | 1.11 ppm |  |  |  | <110.12 ppb
KCM CW02a | Sample #2a | 15 g | 6.889 | 180208 18020801 180213 | 2018-02-08 | (mBq/kg) | 43.618 | +- | 35.11 | <25.87 |  |  | 7.427 | +- | 6.497 | 26.22 | +- | 23.433 | 120.409 | +- | 56.981 | <658.93 |  |  | 97.067 | +- | 82.533
Service Water Sample from Kidd Creek Mine Level 7850 | 29.01.2018_KC7850_SW_SLRS | 7.5 g (inside well) |  |  | 2018-02-15 | (ppb or ppm) | 3.53 ppb | +- | 2.84 ppb | <2.10 ppb |  |  | 13.07 ppb | +- | 11.43 ppb | 6.45 ppb | +- | 5.76 ppb | 29.62 ppb | +- | 14.02 ppb |  |  |  | 7.86 ppb | +- | 6.69 ppb
KCM CW02b | Sample #2b | 14 .2 g | 11.872 | 180307 180314 | 2018-03-07 | (mBq/kg) | 6.269 | +- | 25.482 | 38 | +- | 21.733 | <6.13 |  |  | 24.097 | +- | 18.663 | 59.126 | +- | 42.817 | <325.20 |  |  | 163.33 | +- | 60
Service Water Sample from Kidd Creek Mine Level 7850 | 29.01.2018_KC7850_SW_SLRS | 7.5 g (inside well) |  |  | 2018-03-19 | (ppb or ppm) | 0.51 ppb | +- | 2.06 ppb | 3.08 ppb | +- | 1.76 ppb | <10.80 ppb |  |  | 5.93 ppb | +- | 4.59 ppb | 14.55 ppb | +- | 10.53 ppb |  |  |  | 13.23 ppb | +- | 4.86 ppb
SNOLAB CW02 | FisherBrand FS585008 8 mL test tube | 4.5 g | 22.604 | 18041101 180420 180425 | 2018-04-11 | (mBq/kg) | 88.593 | +- | 39.61 | <29.78 |  |  | 3.16 | +- | 6.059 | 15.876 | +- | 25.244 | 23.534 | +- | 49.039 | <323.78 |  |  | 44.444 | +- | 58.222
Empty Sample Vial | HDPE 17.5 mm diam 63 mm tall |  |  |  | 2018-05-04 | (ppb or ppm) | 7.18 ppb | +- | 3.21 ppb | <2.41 ppb |  |  | 5.56 ppb | +- | 10.66 ppb | 3.91 ppb | +- | 6.21 ppb | 5.79 ppb | +- | 12.06 ppb |  |  |  | 3.60 ppb | +- | 4.72 ppb
SNOLAB CW03 | Lead sample from Soudan lead | 70.9 g | 17.492 | 18121901 190104 | 2018-12-19 | (mBq/kg) | 23.89 | +- | 21.86 | <121.90 |  |  | 29.85 | +- | 21.41 | 27.89 |  | 25.91 | <87.29 |  |  | <32.54 |  |  | 116020 | +- | 6650
Lead Rectangular Cuboid Soudan Lead #1 | 1” x 7/16” x 7/16” 2 rectangular cuboids counted together | Efficiencies are corrected for self-absorption by the lead using ANGLE4 |  |  | 2019-01-07 | (ppb or ppm) | 1.94 ppb | +- | 1.77 ppb | <9.87 ppb |  |  | 52.54 ppb | +- | 37.68 ppb | 6.86 ppb | +- | 6.37 ppb | <21.47 ppb | [5 empty] | 9.4 ppm | +- | 0.54 ppm
SNOLAB CW04 | Lead sample from Soudan lead | 73.3 g | 14.428 | 190527 190601 | 2019-05-27 | (mBq/kg) | 9.812 | +- | 18.56 | <112.30 |  |  | 30.48 | +- | 13.82 | 21.86 | +- | 21.12 | 28.04 | +- | 63.38 | <38.06 |  |  | 18606 | +- | 1719
Lead Rectangular Cuboid Soudan Lead #2 | 1” x 7/16” x 7/16” 2 rectangular cuboids counted together | Efficiencies are corrected for self-absorption by the lead using ANGLE4 |  |  | 2019-06-11 | (ppb or ppm) | 0.79 ppb | +- | 1.5 ppb | <9.1 ppb |  |  | 53.64 ppb | +- | 24.32 ppb | 5.38 ppb | +- | 5.2 ppb | 6.9 ppb | +- | 15.59 ppb |  |  |  | 1.51 ppm | +- | 0.14 ppm
SNOLAB CW05 | Lead sample from Soudan lead | 72.4 g | 23.761 | 190801 190806 | 2019-08-01 | (mBq/kg) | 12.61 | +- | 16.59 | <72.11 |  |  | 33.66 | +- | 16.39 | 33.48 | +- | 23.9 | <79.14 |  |  | 23.107 | +- | 18.46 | 82250 | +- | 4723
Lead Rectangular Cuboid Soudan Lead #3 | 1” x 7/16” x 7/16” 2 rectangular cuboids counted together | Efficiencies are corrected for self-absorption by the lead using ANGLE4 |  |  | 2019-08-26 | (ppb or ppm) | 1.02 ppb | +- | 1.34 ppb | <5.84 ppb |  |  | 59.24 ppb | +- | 28.85 ppb | 8.24 ppb | +- | 5.88 ppb | <19.47 ppb | [5 empty] | 6.66 ppm | +- | 0.38 ppm
SNOLAB CW06 | Lead sample from Soudan lead | 72.3 g | 30.424 | 191007 191014 19101401 191016 | 2019-10-07 | (mBq/kg) | 50.49 | +- | 19.47 | 29.2 | +- | 171.9 | <11.14 |  |  | <22.97 |  |  | <70.28 |  |  | <18.63 |  |  | 140850 | +- | 7580
Lead Rectangular Cuboid Soudan Lead #4 | 1” x 7/16” x 7/16” 2 rectangular cuboids counted together | Efficiencies are corrected for self-absorption by the lead using ANGLE4 |  |  | 2019-11-08 | (ppb or ppm) | 4.09 ppb | +- | 1.58 ppb | 2.37 ppb | +- | 13.92 ppb | <2.74 ppb |  |  | <5.65 ppb |  |  | <17.29 ppb | [5 empty] | 11.41 ppm | +- | 0.61 ppm
SNOLAB CW07 | Swipe Test Prepared on 2019.01.24 | 0.2 g | 6.746 | 190130 | 2019-01-30 | (mBq) | 444.3 | +- | 1481 | <1149.00 |  |  | <162.20 |  |  | <1569.00 |  |  | 489.1 | +- | 2258 | <18670.00 |  |  | 5969 | +- | 3907
XIA Th-230 Source Swipe Test | 1 swipe placed inside vial with the swiped side facing outwards |  |  |  | 2019-02-06 | (ppb or ppm) | 35.99 ppb | +- | 119.96 ppb | <93.07 ppb |  |  | <285.47 ppb |  |  | <385.97 ppb |  |  | 120.32 ppb | +- | 555.47 ppb |  |  |  | 0.48 ppm | +- | 0.32 ppm
SNOLAB CW08 | ~55cm long strand | 24.399 g | 7.092 | 190625 | 2019-06-25 | (mBq) | 17.32 | +- | 25.97 | <41.32 |  |  | 6.484 | +- | 9.399 | <30.45 |  |  | <65.57 |  |  | <125.30 |  |  | 8045.1 | +- | 444.9
White Silver Sil-Pure Pb-free | CMP Global Ltd |  |  |  | 2019-07-02 | (ppb or ppm) | 1.4 ppb | +- | 2.1 ppb | <3.35 ppb |  |  | 11.41 ppb | +- | 16.54 ppb | <7.49 ppb |  |  | <16.13 ppb | [5 empty] | 651.65 ppb | +- | 36.04 ppb
SNOLAB CW09 | NIST Sample | 3.689 g | 17.674 | 200213 200114 | 2020-02-13 | (mBq) | 14370 | +- | 639.2 | 13300 | +- | 1456 | 874.1 | +- | 32.78 | 16630 | +- | 643.9173 | 6942 | +- | 513 | 12269.5655 | +- | 1547.5403 | 10937 | +- | 920.4
Powdered Ocean Sediment Calibration Sample | Cross Comparison Sample |  |  |  | 2020-03-02 | (ppb or ppm) | 1163.97 ppb | +- | 51.78 ppb | 1077.3 ppb | +- | 117.94 ppb | 1538.42 ppb | +- | 57.69 ppb | 4090.98 ppb | +- | 158.4 ppb | 1707.73 ppb | +- | 126.2 ppb |  |  |  | 885.9 ppb | +- | 74.55 ppb
[8 empty] | 60Co |  |  | 241Am |  |  | 40K
[6 empty] | (mBq) | <2990.00 |  |  | 15468 | +- | 803.2 | 387388.8934 | +- | 90244.1064
[6 empty] | (ppb or ppm) | [6 empty] | 12512.66 ppm | +- | 2914.88 ppm
SNOLAB CW10 | Lucas Cell Coating | 13.0 g | 23.614 | 201218 | 2020-12-18 | (mBq/kg) | <15.69 |  |  | 16.39 | +- | 8.276 | <1.14 |  |  | <14.98 |  |  | <31.90 |  |  | 20.307 | +- | 51.52 | 21.055 | +- | 22.39
ZnS(Ag) Powder | Saint-Gobain |  |  | The sample is also counted on the VDA detector as SNOLAB V33 | 2021-01-11 | (ppb or ppm) | <1.27 ppb |  |  | 1.33 ppb | +- | 0.67 ppb | <0.28 ppb |  |  | <3.69 ppb |  |  | <7.85 ppb | [5 empty] | 1.71 ppb | +- | 1.81 ppb
SNOLAB CW11 | Project 2020-11 | 2.8 g | 13.557 | 220204 | 2022-02-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Regular Activated Charcoal |  |  |  |  | 2022-02-18 | (mBq/kg) | 465.5 | +- | 47.48 | 166.9 | +- | 225.4 | 4.382 | +- | 8.565 | 114.5 | +- | 37.57 | 499650 | +- | 27610 | 183.05 | +- | 118.9 | <164.60
[6 empty] | (ppb or ppm) | 37.71 ppb | +- | 3.85 ppb | 13.52 ppb | +- | 18.26 ppb | 7.71 ppb | +- | 15.07 ppb | 28.17 ppb | +- | 9.24 ppb | 16138.7 ppm | +- | 891.8 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 642.91 | +- | 981.5 | <501.40 |  |  | <117.60 |  |  | <194.50
[6 empty] | (ppb or ppm) | 52.08 ppb | +- | 79.5 ppb | [6 empty] | <47.85 ppb
SNOLAB CW12 | Project 2020-11 | 3.2 g | 13.763 | 220218 | 2022-02-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Regular Activated Charcoal Rinsed with Nitric Acid | Sample dried with N2 gas. |  |  |  | 2022-03-04 | (mBq/kg) | <33.25 |  |  | <205.10 |  |  | <5.38 |  |  | 99.75 | +- | 20.06 | 26256 | +- | 2653 | <162.50 |  |  | <72.30
[6 empty] | (ppb or ppm) | <2.69 ppb |  |  | <16.61 ppb |  |  | <9.47 ppb |  |  | 24.54 ppb | +- | 4.93 ppb | 848.07 ppm | +- | 85.69 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <548.20 |  |  | <93.96 |  |  | <24.32 |  |  | 50.23 | +- | 46.65
[6 empty] | (ppb or ppm) | <44.4 ppb | [8 empty] | 12.36 ppb | +- | 11.48 ppb
SNOLAB CW13 | Project 2020-11 | 4.2 g | 6.89 | 220401 | 2022-04-01 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Regular Activated Charcoal |  |  |  |  | 2022-04-08 | (mBq/kg) | 6867 | +- | 232.4 | 7455 | +- | 543.1 | 287.4 | +- | 17.06 | 6201 | +- | 255.6 | 16764 | +- | 2646 | <60.32 |  |  | <75.14
[6 empty] | (ppb or ppm) | 556.23 ppb | +- | 18.82 ppb | 603.86 ppb | +- | 43.99 ppb | 505.82 ppb | +- | 30.03 ppb | 1525.45 ppb | +- | 62.88 ppb | 541.48 ppm | +- | 85.47 ppm
CarboTech - Activated Carbon – CMS-H2 | This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2764.2 | +- | 884.5 | 261.34 | +- | 255.9 | <83.53 |  |  | 5379 | +- | 260.2
[6 empty] | (ppb or ppm) | 223.9 ppb | +- | 71.64 ppb | [6 empty] | 1323.23 ppb | +- | 64.01 ppb
SNOLAB CW14 | Project 2020-11 | 4.7 g | 10.798 | 220408 | 2022-04-08 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Regular Activated Charcoal Rinsed with Nitric Acid | Sample dried with N2 gas. |  |  |  | 2022-04-19 | (mBq/kg) | 3005 | +- | 107.9 | 5538 | +- | 376.3 | 182.3 | +- | 10.14 | 4350 | +- | 169.6 | 9237.8 | +- | 1489 | <94.92 |  |  | <56.32
[6 empty] | (ppb or ppm) | 243.41 ppb | +- | 8.74 ppb | 448.58 ppb | +- | 30.48 ppb | 320.85 ppb | +- | 17.85 ppb | 1070.1 ppb | +- | 41.72 ppb | 298.38 ppm | +- | 48.09 ppm
CarboTech - Activated Carbon – CMS-H2 | This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1950.1 | +- | 629.5 | <209.10 |  |  | <76.88 |  |  | 2219 | +- | 127.5
[6 empty] | (ppb or ppm) | 157.96 ppb | +- | 50.99 ppb | [6 empty] | 545.87 ppb | +- | 31.37 ppb
SNOLAB CW15 | Sample Rock Bolt from TAD Expansion Area | 131.1 g | 12.735 | 230531 | 2023-05-31 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Rock Bolt (Carbon Steel) |  |  |  |  | 2023-06-13 | (mBq/kg) | 33.55 | +- | 2.037 | 101 | +- | 24.59 | 1.861 | +- | 0.3594 | 53.87 | +- | 3.075 | 877.76 | +- | 90.74 | <5.33 |  |  | <2.68
[6 empty] | (ppb or ppm) | 2.72 ppb | +- | 0.16 ppb | 8.18 ppb | +- | 1.99 ppb | 3.28 ppb | +- | 0.63 ppb | 13.25 ppb | +- | 0.76 ppb | 28.35 ppm | +- | 2.93 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 107.31 | +- | 472.3 | 6.4684 | +- | 7.086 | <2.03 |  |  | 50.54 | +- | 3.961
[6 empty] | (ppb or ppm) | 8.69 ppb | +- | 38.26 ppb | [6 empty] | 12.43 ppb | +- | 0.97 ppb
SNOLAB CW16 | Charcoal Trap for Gas Assays at SNOLAB | 4.95 g | 15.649 | 240209 240910 24021001 | 2024-02-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Blucher Protects Spherical Adsorbents |  |  |  |  | 2024-02-26 | (mBq/kg) | 72.96 | +- | 11.98 | <52.53 |  |  | <3.15 |  |  | 5.056 | +- | 10.02 | 1509 | +- | 559.6 | <85.24 |  |  | <10.57
[6 empty] | (ppb or ppm) | 5.91 ppb | +- | 0.97 ppb | <4.25 ppb |  |  | <5.54 ppb |  |  | 1.24 ppb | +- | 2.46 ppb | 48.74 ppm | +- | 18.08 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <363.70 |  |  | <107.90 |  |  | <14.78 |  |  | 8.061 | +- | 17.71
[6 empty] | (ppb or ppm) | <29.46 ppb | [8 empty] | 1.98 ppb | +- | 4.36 ppb
SNO+:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
SNO+ CW01 |  | --- | 5.776 | 140312 14031302 140316 | 2014-03-12 | (mBq) | 0.406 | +- | 0.332 | 1.908 | +- | 0.333 | 0.368 | +- | 0.088 | 0.589 | +- | 0.199 | <0.219 |  |  | 8.8 | +- | 5.2 | 80.2 | +- | 3.1
Acrylic_G UPW 12C, leached 37 days |  |  |  |  | 2014-03-18
SNO+ CW02 |  | --- | 10.875 | 140327 140330 140404 | 2014-03-27 | (mBq) | 0.224 | +- | 0.222 | <0.181 |  |  | 0.125 | +- | 0.068 | 0.1 | +- | 0.1 | <0.635 |  |  | <0.764 |  |  | 3.1 | +- | 0.6
UPW Assay from 140214 Supernatant Solution |  |  |  |  | 2014-04-07
SNO+ CW03 |  | --- | 15.161 | 140407 140421 | 2014-04-07 | (mBq) | 0.194 | +- | 0.16 | 0.24 | +- | 0.089 | 0.012 | +- | 0.068 | 0.297 | +- | 0.1 | 0.121 | +- | 0.447 | <1.338 |  |  | 2.6 | +- | 0.5
Water Assay with Dowex, set #2, 140214 |  |  |  |  | 2014-04-23
SNO+ CW04 |  | --- | 4.928 | 140423 | 2014-04-23 | (mBq/kg) | <0.050 |  |  | 0.062 | +- | 0.197 | <0.055 |  |  | 0.297 | +- | 0.199 | 0.398 | +- | 0.735 | 3.3 | +- | 3.5 | 37.1 | +- | 2.1
EDTA 16 Days Leaching A_D 010414 |  |  |  |  | 2014-04-28 | (ppb or ppm)
SNO+ CW05 | Set 1, 010414 | --- | 7.903 | 140428 140505 | 2014-04-28 | (mBq/kg) | <0.231 |  |  | 0.12 | +- | 0.179 | 0.095 | +- | 0.086 | 0.37 | +- | 0.198 | <0.224 |  |  | <2.620 |  |  | 48.9 | +- | 1.9
Cavity Water Assay Supernate |  |  |  |  | 2014-05-06 | (ppb or ppm)
SNO+ CW06 | Set 2, 010414 | --- | 7.826 | 140506 | 2014-05-06 | (mBq/kg) | <0.209 |  |  | <0.206 |  |  | 0.127 | +- | 0.069 | 0.213 | +- | 0.198 | 0.912 | +- | 0.853 | <2.620 |  |  | 2.2 | +- | 0.6
Cavity Water Assay Supernate |  |  |  |  | 2014-05-14 | (ppb or ppm)
SNO+ CW07 | Prepared On: 270414 | --- | 17.936 | 140514 140519 140530 | 2014-05-14 | (mBq/kg) | 0.693 | +- | 0.215 | <0.042 |  |  | 0.239 | +- | 0.068 | 0.107 | +- | 0.1 | <0.635 |  |  | <1.082 |  |  | 3.1 | +- | 0.5
A_D UPW Leach 25 days |  |  |  |  | 2014-06-02 | (ppb or ppm)
SNO+ CW08 | Prepared On: 270514 |  | 7.306 | 140602 | 2014-06-02 | (mBq/kg) | <0.051 |  |  | 0.2 | +- | 0.141 | 0.191 | +- | 0.069 | <0.149 |  |  | <0.280 |  |  | <2.877 |  |  | 2.2 | +- | 0.6
HCL QSIL Wash |  |  |  |  | 2014-06-09 | (ppb or ppm)
SNO+ CW09 | Prepared On: 270514 |  | 7.686 | 140609 | 2014-06-09 | (mBq/kg) | 0.076 | +- | 0.229 | 0.023 | +- | 0.166 | 0.113 | +- | 0.048 | 0.053 | +- | 0.199 | <0.592 |  |  | <2.749 |  |  | 10.4 | +- | 1
Qsil A_G LAB |  |  |  |  | 2014-06-17 | (ppb or ppm)
SNO+ CW10 | Prepared On: 270514 |  | 5.84 | 140708 | 2014-07-08 | (mBq/kg) | 0.349 | +- | 0.294 | 0.308 | +- | 0.166 | <0.11 |  |  | 0.297 | +- | 0.199 | <0.94 |  |  | <3.67 |  |  | 4.6 | +- | 0.8
LAB A_G 5 mL |  |  |  |  | 2014-07-15 | (ppb or ppm)
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
SNO+ CW11 | Prepared On: 270514 |  | 8.896 | 140714 140722 | 2014-07-14 | (mBq/kg) | <0.32 |  |  | 0.054 | +- | 0.166 | 0.023 | +- | 0.044 | 0.081 | +- | 0.197 | <0.55 |  |  | <2.36 |  |  | 2.4 | +- | 0.6
HCL QSIL Wash 2 mLSS |  |  |  |  | 2014-07-23 | (ppb or ppm)
SNO+ CW12 | Prepared On: 270514 |  | 10.5 | 140807 | 2014-08-07 | (mBq/kg) | <0.025 |  |  | <0.211 |  |  | 0.063 | +- | 0.044 | 0.093 | +- | 0.198 | <0.197 |  |  | <1.979 |  |  | 2.3 | +- | 0.7
LAB A_G Qsil 10 mL |  |  |  |  | 2014-08-18 | (ppb or ppm)
SNO+ CW13 | Prepared On: 070714 |  | 7.91 | 140723 | 2014-07-23 | (mBq/kg) | 0.714 | +- | 0.285 | 0.1 | +- | 0.166 | 0.003 | +- | 0.048 | 0.196 | +- | 0.198 | 0.374 | +- | 0.594 | <2.62 |  |  | 9.6 | +- | 0.9
EDTA Leaching A_D 25C |  |  |  |  | 2014-07-31 | (ppb or ppm)
SNO+ CW14 |  |  | 6.451 | 140731 | 2014-07-31 | (mBq/kg) | <0.264 |  |  | <0.114 |  |  | 0.064 | +- | 0.087 | 1.576 | +- | 0.295 | <0.505 |  |  | <3.26 |  |  | 16.4 | +- | 1.3
A_B EDTA Leach 0.027M, 12 C |  |  |  |  | 2014-08-07 | (ppb or ppm)
SNO+ CW15 | Prepared On: 110814 |  | 6.95 | 140825 140830 | 2014-08-25 | (mBq) | <0.212 |  |  | 0.2 | +- | 0.447 | <0.086 |  |  | 4.076 | +/- | 0.392 | 0.071 | +- | 0.77 | <3.005 |  |  | 14.8 | +- | 1.3
AV Tube EDTA Leaching 17/25 |  |  |  |  | 2014-09-02
SNO+ CW16 | Prepared On: 140814 |  | 5.861 | 140902 | 2014-09-02 | (mBq) | 0.175 | +- | 0.235 | 5.844 | +- | 0.591 | 0.331 | +- | 0.088 | 1.995 | +/- | 0.297 | 2.001 | +- | 0.876 | <3.646 |  |  | 7.7 | +- | 1.1
AD Leaching in TeLS 12C, 23 days |  |  |  |  | 2014-09-08
SNO+ CW17 | Prepared On: 140901 |  | 6.812 | 140915 140918 | 2014-09-15 | (mBq) | <0.226 |  |  | 0.069 | +- | 0.186 | 0.109 | +- | 0.088 | 1.137 | +/- | 0.296 | 0.22 | +- | 0.55 | <3.01 |  |  | 4.3 | +- | 0.8
LWA 140801 Supernate, Set 1 |  |  |  |  | 2014-09-22
SNO+ CW18 | Prepared On: 140901 |  | 6.874 | 140922 14092201 | 2014-09-22 | (mBq) | 0.148 | +/- | 0.297 | 0.392 | +- | 0.455 | 0.154 | +- | 0.088 | 5.275 | +/- | 0.393 | < 0.41 |  |  | <3.01 |  |  | 2.6 | +- | 0.9
LWA 140801 Supernate, Set 2 |  |  |  |  | 2014-09-29
SNO+ CW19 | Prepared On: Nov. 2014 |  | 18.569 | 14112604 141205 141209 | 2014-11-26 | (mBq) | 0.223 | +- | 0.154 | 1.175 | +- | 0.166 | 0.085 | +- | 0.033 | 0.34 | +- | 0.137 | <0.56 |  |  | <1.54 |  |  | 3.324 | +- | 0.489
A_D PRS 12C, 30 Days, Nov 2014 |  |  |  |  | 2014-12-15
SNO+ CW20 | Prepared On: January 27, 2015 |  | 25.076 | 150129 150130 150209 15020901 150211 | 2015-01-29 | (mBq) | 0.277 | +- | 0.149 | 0.126 | +- | 0.166 | 0.091 | +- | 0.031 | 0.333 | +- | 0.126 | 0.188 | +- | 0.282 | <1.13 |  |  | 922.818 | +- | 47.153
AV LWA Supernate at 50% |  |  |  |  | 2015-02-23
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
SNO+ CW21 |  |  | 5.826 | 150310 150313 | 2015-03-10 | (mBq) | 1.151 | +- | 0.347 | 0.608 | +- | 0.278 | 0.207 | +- | 0.068 | 0.458 | +- | 0.234 | <3.74 |  |  | <5.86 |  |  | 102.736 | +- | 6.073
Radon Box #1 |  |  |  |  | 2015-03-16
SNO+ CW22 |  |  | 4.951 | 150316 | 2015-03-16 | (mBq) | 0.487 | +- | 0.355 | 0.3 | +- | 0.247 | 0.019 | +- | 0.056 | 1.02 | +- | 0.267 | 0.862 | +- | 0.539 | <6.92 |  |  | 14.288 | +- | 1.571
Radon Box #2 |  |  |  |  | 2015-03-21
SNO+ CW23 |  |  | 4.91 | 150321 150323 | 2015-03-21 | (mBq) | <0.34 |  |  | 0.331 | +- | 0.198 | 0.073 | +- | 0.053 | 0.353 | +- | 0.21 | 0.139 | +- | 0.5 | <6.98 |  |  | 3.883 | +- | 0.878
Radon Box #3 |  |  |  |  | 2015-03-26
SNO+ CW24 |  |  | 7.722 | 150408 | 2015-04-08 | (mBq) | 0.978 | +- | 0.37 | 0.578 | +- | 0.25 | 0.092 | +- | 0.056 | 0.726 | +- | 0.233 | <1.39 |  |  | 6.512 | +- | 4.203 | 4.548 | +- | 0.888
Radon Box #4 |  |  |  |  | 2015-04-16
SNO+ CW25 |  |  | 3.889 | 150416 | 2015-04-16 | (mBq) | 2.498 | +- | 0.76 | 2.472 | +- | 0.59 | 0.228 | +- | 0.117 | 2.329 | +- | 0.462 | 0.898 | +- | 1.021 | <8.34 |  |  | 63.865 | +- | 4.672
Radon Box #5_1 |  |  |  |  | 2015-04-20
SNO+ CW26 |  |  | 9.868 | 150220 | 2015-04-20 | (mBq) | 0.68 | +- | 0.301 | 0.308 | +- | 0.187 | 0.096 | +- | 0.047 | 0.605 | +- | 0.178 | <1.07 |  |  | <3.42 |  |  | 15.101 | +- | 1.304
Radon Box #5_2 |  |  |  |  | 2015-04-30
SNO+ CW27 |  |  | 6.792 | 150430 150505 15050504 | 407366 | (mBq) | <0.092 |  |  | <0.12 |  |  | 0.012 | +- | 0.043 | 0.335 | +- | 0.207 | 0.464 | +- | 0.468 | <5.01 |  |  | 10.49 | +- | 1.15
Radon Box #6 |  |  |  |  | 2015-05-07
SNO+ CW28 |  |  | 6.903 | 150507 | 2015-05-07 | (mBq) | 0.067 | +- | 0.258 | <0.091 |  |  | <0.039 |  |  | 0.2 | +- | 0.173 | <0.75 |  |  | <4.93 |  |  | 7.196 | +- | 0.938
Radon Box #7 |  |  |  |  | 2015-05-14
SNO+ CW29 |  |  | 6.889 | 150326 | 2015-03-26 | (mBq) | 0.533 | +- | 0.301 | <0.071 |  |  | 0.016 | +- | 0.052 | 2.495 | +- | 0.338 | 0.271 | +- | 0.512 | <4.94 |  |  | 2.112 | +- | 0.733
Cavity Water Assay #1, 22 litres |  |  |  |  | 2015-04-02
SNO+ CW30 |  |  | 5.868 | 150402 | 2015-04-02 | (mBq) | 0.286 | +- | 0.292 | <0.16 |  |  | 0.043 | +- | 0.052 | 2.044 | +- | 0.319 | <0.90 |  |  | <5.82 |  |  | 1.71 | +- | 0.7
Cavity Water Assay #2, Dowex Eluted |  |  |  |  | 2015-04-08
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
SNO+ CW31 |  | 18.5 g | 9.965 | 160205 160206 16020601 16020604 160210 160214 160215 160216 | 2016-02-05 | (mBq/kg) | <17.35 |  |  | 6.093 | +- | 8.197 | 2.974 | +- | 2.318 | 61.058 | +- | 11.926 | 29.087 | +- | 20.3421 | 82.865 | +- | 115.351 | 182.324 | +- | 34.054
SNO+ Water Assay From 160203 |  |  |  |  | 2016-02-16 | (ppb or ppm) | <1.41 ppb |  |  | 0.49 ppb | +- | 0.66 ppb | 5.23 ppb | +- | 4.08 ppb | 15.02 ppb | +- | 2.93 ppb | 7.16 ppb | +- | 5.00 ppb |  |  |  | 14.77 ppb | +- | 2.76 ppb
SNO+ CW32 |  | 13.4 g | 5.776 | 16021601 160217 16021702 16021705 16021707 16021708 16021713 160218 16021802 16021803 16021805 | 2016-02-16 | (mBq/kg) | <15.52 |  |  | <10.99 |  |  | 0.821 | +- | 3.951 | <21.27 |  |  | <86.27 |  |  | <441.12 |  |  | 233.28 | +- | 62.61
SNO+ Assay Cleanup from 160128 |  |  |  |  | 2016-02-23 | (ppb or ppm) | <1.26 ppb |  |  | <0.89 ppb |  |  | 1.44 ppb | +- | 6.95 ppb | <5.23 ppb |  |  | <21.22 ppb | [5 empty] | 18.90 ppb | +- | 5.07 ppb
SNO+ CW33 | Sample made on 15.02.16 |  | 5.972 | 160223 | 2016-02-23 | (mBq/kg) | 0.057 | +- | 0.251 | <0.21 |  |  | <0.086 |  |  | 0.152 | +- | 0.179 | 0.971 | +- | 0.475 | <5.72 |  |  | 2.436 | +- | 0.709
A_D Leaching is TeBD_LS |  |  |  |  | 2016-02-29 | (ppb or ppm) | 0.005 ppb | +- | 0.020 ppb | <0.017 ppb |  |  | <0.15 ppb |  |  | 0.038 ppb | +- | 0.044 ppb | 0.24 ppb | +- | 0.12 ppb
SNO+ CW34 | Sample made on 08.02.16 | 37.0 g | 16.118 | 160229 16022908 160301 160303 160316 | 2016-02-29 | (mBq/kg) | <3.49 |  |  | 4.366 | +- | 2.876 | <1.19 |  |  | 0.928 | +- | 3.404 | <20.14 |  |  | <55.81 |  |  | 58.892 | +- | 11.811
Supernatant Sample |  |  |  |  | 2016-03-17 | (ppb or ppm) | <0.28 ppb |  |  | 0.35 ppb | +- | 0.23 ppb | <2.09 ppb |  |  | 0.20 ppb | +- | 0.84 ppb | <4.95 ppb
SNO+ CW35 | Sample made on 160208 | 18.1 g | 7.66 | 160331 | 2016-03-31 | (mBq/kg) | 10.494 | +- | 17.15 | <19.88 |  |  | <3.48 |  |  | 44.018 | +- | 12.771 | <37.98 |  |  | <245.03 |  |  | 112.21 | +- | 49.006
Supernatant LWA |  |  |  |  | 2016-04-08 | (ppb or ppm) | 0.85 ppb | +- | 1.39 ppb | <1.61 ppb |  |  | <6.13 ppb |  |  | 10.83 ppb | +- | 3.04 ppb | <9.34 ppb | [5 empty] | 9.09 ppb | +- | 3.97 ppb
SNO+ CW36 | Sample made on 160218B | 17.1 g | 6.319 | 160408 | 2016-04-08 | (mBq/kg) | 24.407 | +- | 16.862 | 5.682 | +- | 10.193 | 5.313 | +- | 2.679 | 5.397 | +- | 10.123 | 24.916 | +- | 25.792 | <315.56 |  |  | 246.082 | +- | 45.556
Supernatant LWA |  |  |  |  | 2016-04-14 | (ppb or ppm) | 1.98 ppb | +- | 1.37 ppb | 0.46 ppb | +- | 0.83 ppb | 9.35 ppb | +- | 4.71 ppb | 1.33 ppb | +- | 2.49 ppb | 6.13 ppb | +- | 6.35 ppb |  |  |  | 19.93 ppb | +- | 3.69 ppb
SNO+ CW37 | Sample made on 160309 | 18.6 g | 6.554 | 160414 160415 16041501 | 2016-04-14 | (mBq/kg) | <16.08 |  |  | 7.276 | +- | 9.295 | 1.351 | +- | 2.359 | 23.801 | +- | 10.106 | 14.583 | +- | 22.303 | <279.46 |  |  | 98.763 | +- | 39.785
Supernatant LWA |  |  |  |  | 2016-04-21 | (ppb or ppm) | <1.30 ppb |  |  | 0.59 ppb | +- | 0.75 ppb | 2.38 ppb | +- | 4.15 ppb | 5.86 ppb | +- | 2.49 ppb | 3.59 ppb | +- | 5.49 ppb |  |  |  | 8.00 ppb | +- | 3.22 ppb
SNO+ CW38 | Sample made on 160330 | 20.0 g | 6.528 | 160421 160426 | 2016-04-21 | (mBq/kg) | 17.007 | +- | 14.439 | <11.24 |  |  | 1.325 | +- | 2.301 | <15.35 |  |  | 0.233 | +- | 20.923 | <261.05 |  |  | 103.75 | +- | 30.5
Supernatant LWA | Blank |  |  |  | 2016-04-28 | (ppb or ppm) | 1.38 ppb | +- | 1.17 ppb | <0.91 ppb |  |  | 2.33 ppb | +- | 4.05 ppb | <3.78 ppb |  |  | 0.06 ppb | +- | 5.15 ppb |  |  |  | 8.40 ppb | +- | 2.47 ppb
SNO+ CW39 | Sample made on 160404 | 20.5 g | 13.799 | 160428 | 2016-04-28 | (mBq/kg) | 3.733 | +- | 10.311 | <6.77 |  |  | 0.191 | +- | 1.608 | 17.985 | +- | 7.009 | 14.043 | +- | 14.853 | <118.24 |  |  | 60.293 | +- | 20.585
CF-Scint-Assay Supernatant for Pb210 |  |  |  |  | 2016-05-12 | (ppb or ppm) | 0.30 ppb | +- | 0.84 ppb | <0.55 ppb |  |  | 0.34 ppb | +- | 2.83 ppb | 4.42 ppb | +- | 1.72 ppb | 3.46 ppb | +- | 3.65 ppb |  |  |  | 4.88 ppb | +- | 1.67 ppb
SNO+ CW40 | AV Set 5 160308 | 22.4 g | 6.312 | 160512 160514 160517 | 2016-05-12 | (mBq/kg) | <21.12 |  |  | <5.66 |  |  | 0.308 | +- | 2.409 | 154.401 | +- | 17.856 | <33.13 |  |  | 113.259 | +- | 150 | 315.804 | +- | 46.473
Supernatant LWA |  |  |  |  | 2016-05-19 | (ppb or ppm) | <1.71 ppb |  |  | <0.46 ppb |  |  | 0.54 ppb | +- | 4.24 ppb | 37.98 ppb | +- | 4.39 ppb | <8.15 ppb | [5 empty] | 25.58 ppb | +- | 3.76 ppb
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
SNO+ CW41 | Sample Prepared on 160512 | -- | 11.743 | 160519 160520 | 2016-05-19 | (mBq) | <0.24 |  |  | 0.253 | +- | 0.133 | 0.033 | +- | 0.034 | 0.294 | +- | 0.143 | 0.432 | +- | 0.332 | 2.742 | +- | 2.339 | 1.795 | +- | 0.499
A_D Leach TeBD 56 days |  |  |  |  | 2016-05-31 | (ppb or ppm)
SNO+ CW42 | Sample Prepared on 160517 | 21.5 g | 8.213 | 160531 160603 | 2016-05-31 | (mBq/kg) | 0.509 | +- | 9.748 | 0.105 | +- | 10.405 | <2.84 |  |  | 15.859 | +- | 8.453 | 1.317 | +- | 20.781 | <192.00 |  |  | 17896.372 | +- | 939.116
Supernatant LWA AV Assay |  |  |  |  | 2016-06-08 | (ppb or ppm) | 0.05 ppb | +- | 0.79 ppb | 0.01 ppb | +- | 0.84 ppb | <4.99 ppb |  |  | 3.90 ppb | +- | 2.08 ppb | 0.32 ppb | +- | 5.11 ppb |  |  |  | 1.45 ppm | +- | 0.08 ppm
SNO+ CW43 | Sample Prepared on 170330 | 18.9 g | 13.59 | 170511 170518 | 2017-05-11 | (mBq/kg) | <10.90 |  |  | <11.69 |  |  | 1.073 | +- | 1.887 | 0.092 | +- | 7.095 | <32.43 |  |  | <112.33 |  |  | 58.571 | +- | 25.767
Supernatant LWA |  |  |  |  | 2017-05-25 | (ppb or ppm) | <0.88 ppb |  |  | <0.95 ppb |  |  | 1.89 ppb | +- | 3.32 ppb | 0.02 ppb | +- | 1.75 ppb | <7.98 ppb | [5 empty] | 4.74 ppb | +- | 2.09 ppb
SNO+ CW44 | Sample Prepared on 170613 | 30.5 g | 7.792 | 170711 17071101 | 2017-07-11 | (mBq/kg) | 6.027 | +- | 8.981 | 5.26 | +- | 6.165 | 1.852 | +- | 1.548 | 1.814 | +- | 5.209 | <19.21 |  |  | <142.89 |  |  | 4125.508 | +- | 230.754
Supernatant LWA |  |  |  |  | 2017-07-19 | (ppb or ppm) | 0.49 ppb | +- | 0.73 ppb | 0.43 ppb | +- | 0.50 ppb | 3.26 ppb | +- | 2.73 ppb | 0.45 ppb | +- | 1.28 ppb | <4.73 ppb | [5 empty] | 334.17 ppb | +- | 18.69 ppb
SNO+ CW45 | Sample Prepared on 160830 | 19.9 g | 16.75 | 161017 161020 | 2016-10-17 | (mBq/kg) | 10.472 | +- | 9.51 | 12.121 | +- | 5.372 | 1.356 | +- | 1.471 | .6.16 | +- | 6.38 | 8.512 | +- | 13.417 | <99.75 |  |  | 251.407 | +- | 28.392
Supernatant LWA |  |  |  |  | 2016-11-03 | (ppb or ppm) | 0.85 ppb | +- | 0.77 ppb | 0.98 ppb | +- | 0.44 ppb | 2.39 ppb | +- | 2.59 ppb | 1.52 ppb | +- | 1.57 ppb | 2.09 ppb | +- | 3.30 ppb |  |  |  | 20.36 ppb | +- | 2.30 ppb
SNO+ CW46 |  |  | 18.694 | 170719 170720 170804 170807 | 2017-07-19 | (mBq) | 0.051 | +- | 0.132 | <0.23 |  |  | 0.064 | +- | 0.028 | 0.006 | +- | 0.124 | 0.349 | +- | 0.251 | <1.53 |  |  | 22.865 | +- | 1.446
Acrylic Leaching DDA-LAB |  |  |  |  | 2017-08-09 | [7 empty] | .
SNO+ CW47 | Sample Prepared on 170802 & 03 | 21.1 g | 8.826 | 170815 170818 | 2017-08-15 | (mBq/kg) | 21.931 | +- | 13.428 | 0.745 | +- | 6.977 | <2.37 |  |  | 7.939 | +- | 7.618 | 46.918 | +- | 19.882 | <181.80 |  |  | 48.483 | +- | 28.104
Supernatant LWA |  |  |  |  | 2017-08-24 | (ppb or ppm) | 1.78 ppb | +- | 1.09 ppb | 0.06 ppb | +- | 0.56 ppb | <4.17 ppb |  |  | 1.95 ppb | +- | 1.87 ppb | 11.54 ppb | +- | 4.89 ppb |  |  |  | 3.93 ppb | +- | 2.28 ppb
SNO+ CW48 | Sample Prepared on 171116, sample 1 | 15.7 g | 6.727 | 171206 171213 | 2017-12-06 | (mBq/kg) | 518.41 | +- | 58.811 | <15.48 |  |  | 6.451 | +- | 5.483 | 1.013 | +- | 15.574 | 82.306 | +- | 49.584 | <322.42 |  |  | 208.981 | +- | 60.255
Supernatant LWA |  |  |  |  | 2017-12-13 | (ppb or ppm) | 41.99 ppb | +- | 4.44 ppb | <1.25 ppb |  |  | 11.35 ppb | +- | 9.65 ppb | 0.25 ppb | +- | 3.83 ppb | 20.25 ppb | +- | 12.20 ppb |  |  |  | 16.93 ppb | +- | 4.88 ppb
SNO+ CW49 | Sample Prepared on 171116, sample 2 | 15.8 g | 4.797 | 17121301 | -618008 | (mBq/kg) | 30.116 | +- | 24.493 | <30.57 |  |  | 3.689 | +- | 3.635 | 8.693 | +- | 14.641 | 71.429 | +- | 36.351 | <451.84 |  |  | 489.747 | +- | 76.076
Supernatant LWA |  |  |  |  | 2017-12-18 | (ppb or ppm) | 2.44 ppb | +- | 1.98 ppb | <2.48 ppb |  |  | 6.49 ppb | +- | 6.40 ppb | 2.14 ppb | +- | 3.60 ppb | 17.57 ppb | +- | 8.94 ppb |  |  |  | 39.67 ppb | +- | 6.16 ppb
SNO+ CW50 | Sample Prepared on 171207 | 12.9 g | 1.021 | 180104 18010401 | 2018-01-04 | (mBq/kg) | <108.68 |  |  | <116.51 |  |  | 17.541 | +- | 16.426 | 31.458 | +- | 29.168 | 27.103 | +- | 168.468 | <2633.80 |  |  | 334496.124 | +- | 17403.101
Supernatant | This sample is from SNO+ AV |  |  |  | 2018-01-05 | (ppb or ppm) | <8.80 ppb |  |  | <9.44 ppb |  |  | 30.87 ppb | +- | 28.91 ppb | 7.74 ppb | +- | 7.18 ppb | 6.67 ppb | +- | 41.44 ppb |  |  |  | 27.09 ppm | +- | 1.41 ppm
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
SNO+ CW51 | Sample Labeled AV-171207 | 7.5 g | 10.011 | 180319 | 2018-03-19 | (mBq/kg) | 16.463 | +- | 30.791 | 8.813 | +- | 16.604 | 13.311 | +- | 5.343 | 17.572 | +- | 20.295 | 39.706 | +- | 46.279 | <449.47 |  |  | 4536.267 | +- | 293.867
AV Assay | Sample Prepared on March 15, 2018 by D.C. |  |  |  | 2018-03-29 | (ppb or ppm) | 1.33 ppb | +- | 2.49 ppb | 0.71 ppb | +- | 1.35 ppb | 23.43 ppb | +- | 9.40 ppb | 4.32 ppb | +- | 4.99 ppb | 9.77 ppb | +- | 11.39 ppb |  |  |  | 367.44 ppb | +- | 23.80 ppb
SNO+ CW52 | Sample LWA-180321 | 25.7 g | 10.99 | 180329 180411 | 2018-03-29 | (mBq/kg) | 33.35 | +- | 11.734 | <8.17 |  |  | <2.30 |  |  | 12.873 | +- | 6.214 | <29.30 |  |  | <119.18 |  |  | 150.973 | +- | 23.658
Supernatant 180322 | Sample prepared on 180322 |  |  |  | 2018-04-11 | (ppb or ppm) | 2.70 ppb | +- | 0.95 ppb | <0.66 ppb |  |  | <4.04 ppb |  |  | 3.17 ppb | +- | 1.53 ppb | <7.21 ppb | [5 empty] | 12.23 ppb | +- | 1.92 ppb
SNO+ CW53 | Sample prepared on 180419 by Ian Lam | 8.8 g | 13.861 | 180516 180517 | 2018-05-16 | (mBq/kg) | 82.8 | +- | 26.279 | 0.15 | +- | 11.684 | 8.478 | +- | 3.86 | 9.889 | +- | 15.68 | <137.39 |  |  | <274.21 |  |  | 37.045 | +- | 41.023
US AV 171116-2 |  |  |  |  | 2018-05-30 | (ppb or ppm) | 6.71 ppb | +- | 2.13 ppb | 0.01 ppb | +- | 0.95 ppb | 14.92 ppb | +- | 6.79 ppb | 2.43 ppb | +- | 3.90 ppb | <33.80 ppb | [5 empty] | 3.01 ppb | +- | 3.32 ppb
SNO+ CW54 | Sample Prepared on 160420 by L. Anselmo | 20.5 g | 3.924 | 180504 | 2018-05-04 | (mBq/kg) | 46.183 | +- | 28.631 | <98.73 |  |  | <6.88 |  |  | <19.27 |  |  | <110.39 |  |  | <427.07 |  |  | 491219.512 | +- | 25058.537
Supernatant 180416-US |  |  |  |  | 2018-05-08 | (ppb or ppm) | 3.74 ppb | +- | 2.32 ppb | <8.00 ppb |  |  | <12.11 ppb |  |  | <4.74 ppb |  |  | <27.16 ppb | [5 empty] | 39.79 ppm | +- | 2.03 ppm
SNO+ CW55 | Sample Prepared on 160420 by L. Anselmo | 21.8 g | 7.931 | 180508 18050801 | 2018-05-08 | (mBq/kg) | 24.361 | +- | 16.378 | <38.35 |  |  | 0.681 | +- | 2.285 | 13.055 | +- | 8.354 | <104.95 |  |  | <196.33 |  |  | 11282.339 | +- | 602.844
Supernatant 180416-DS |  |  |  |  | 2018-05-16 | (ppb or ppm) | 1.97 ppb | +- | 1.33 ppb | <3.11 ppb |  |  | 1.20 ppb | +- | 4.02 ppb | 3.21 ppb | +- | 2.06 ppb | <25.82 ppb | [5 empty] | 913.06 ppb | +- | 48.83 ppb
SNO+ CW56 | Sample prepared on 180530 by Tom Sonley | 8.5 g | 11.91 | 180530 180604 180608 180612 | 2018-05-30 | (mBq/kg) | 73.186 | +- | 28.268 | 5.087 | +- | 16.743 | 3.661 | +- | 4.33 | 42.262 | +- | 18.087 | 32.584 | +- | 39.056 | <286.12 |  |  | 12782.824 | +- | 698.941
US 180416 |  |  |  |  | 2018-06-13 | (ppb or ppm) | 5.93 ppb | +- | 2.29 ppb | 0.41 ppb | +- | 1.36 ppb | 6.44 ppb | +- | 7.62 ppb | 10.40 ppb | +- | 4.45 ppb | 8.02 ppb | +- | 9.61 ppb |  |  |  | 1.04 ppm | +- | 0.06 ppm
SNO+ CW57 | Sample prepared on 180530 by Tom Sonley | 8.2 g | 13.968 | 180613 180619 18061901 18061902 180625 | 2018-06-13 | (mBq/kg) | 195.641 | +- | 34.865 | 10.689 | +- | 14.457 | 10.385 | +- | 4.686 | 29.691 | +- | 17.527 | 68.7 | +- | 40.148 | 201.829 | +- | 239.878 | 170.854 | +- | 50.854
DS 180416 |  |  |  |  | 2018-06-28 | (ppb or ppm) | 14.85 ppb | +- | 2.82 ppb | 0.87 ppb | +- | 1.17 ppb | 18.28 ppb | +- | 8.25 ppb | 7.30 ppb | +- | 4.31 ppb | 16.90 ppb | +- | 9.88 ppb |  |  |  | 13.84 ppb | +- | 4.12 ppb
SNO+ CW58 | Sample Prepared on 180923 | 7.11 g | 15.232 | 171017 171018 | 2018-10-17 | (mBq/kg) | <413.80 |  |  | 2210 | +- | 1293 | <36.51 |  |  | 256300 | +- | 14490 | <1134.00 |  |  | <946.50 |  |  | 23107 | +- | 1818
Dowex Resin |  |  |  |  | 2018-11-02 | (ppb or ppm) | <33.52 ppb |  |  | 179.01 ppb | +- | 104.73 ppb | <64.26 ppb |  |  | 63.05 ppm | +- | 35.65 ppm | <278.96 ppb | [5 empty] | 1871.67 ppb | +- | 147.26 ppb
SNO+ CW59 | Sample Prepared on 181205 | 1.1 g | 13.606 | 181205 181219 | 2018-12-05 | (mBq/kg) | <728.80 |  |  | <4085.00 |  |  | <219.60 |  |  | 535300 | +- | 27000 | <2226.00 |  |  | <4256.00 |  |  | 36046 | +- | 4051
Dowex Resin |  |  |  |  | 2018-12-19 | (ppb or ppm) | <59.03 ppb |  |  | <330.89 ppb |  |  | <386.5 ppb |  |  | 131.68 ppm | +- | 66.42 ppm | <547.6 ppb | [5 empty] | 2919.73 ppb | +- | 328.13 ppb
SNO+ CW60 | Eight meters of umibilical was soaked in boiling water. | 1.5 g | 8.75 | 190107 | 2019-01-07 | (mBq/kg) | 1124 | +- | 298.3 | 1821 | +- | 262.9 | 234.1 | +- | 46.86 | 586.2 | +- | 220.1 | <592.90 |  |  | <1923.00 |  |  | 1366.5 | +- | 592
Powdered Umbillical Leachate | The sample was evaporated to become a powder. |  |  |  | 2019-01-16 | (ppb or ppm) | 91.04 ppb | +- | 24.16 ppb | 147.5 ppb | +- | 21.29 ppb | 412.02 ppb | +- | 82.47 ppb | 144.21 ppb | +- | 54.14 ppb | <145.85 ppb | [5 empty] | 110.69 ppb | +- | 47.95 ppb
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
SNO+ CW61 | Sample Prepared on 190206 | 8.013 g | 7.044 | 190206 | 2019-02-06 | (mBq/kg) | 179.7 | +- | 98.1 | <91.74 |  |  | <20.21 |  |  | 15680 | +- | 776.5 | <239.10 |  |  | <322.50 |  |  | <370.50
Dowex Effluent |  |  |  |  | 2019-02-13 | (ppb or ppm) | 14.56 ppb | +- | 7.95 ppb | <7.43 ppb |  |  | <35.57 ppb |  |  | 3857.28 ppb | +- | 191.02 ppb | <58.82 ppb | [5 empty] | <30.01 ppb
SNO+ CW62 | Sample Prepared on 190301 | ?? g | 6.878 | 190301 | 2019-03-01 | (mBq/kg) | <30.20 |  |  | <18.04 |  |  | 3.081 | +- | 8.822 | 1469 | +- | 113.9 | <110.50 |  |  | <457.90 |  |  | <141.20
Dowex Effluent |  | Assume 8g |  |  | 2019-03-08 | (ppb or ppm) | <2.45 ppb |  |  | <1.46 ppb |  |  | 5.42 ppb | +- | 15.53 ppb | 361.37 ppb | +- | 28.02 ppb | <27.18 ppb | [5 empty] | <11.44 ppb
SNO+ CW63 | 190225-BLK | 8..8 g | 9.935 | 190308 | 2019-03-08 | (mBq/kg) | 59.53 | +- | 33.65 | <29.05 |  |  | 3.861 | +- | 5.787 | <36.55 |  |  | <41.52 |  |  | <289.10 |  |  | 267.13 | +- | 78.18
SNO+ Supernatant Sample |  |  |  |  | 2019-03-18 | (ppb or ppm) | 4.82 ppb | +- | 2.73 ppb | <2.35 ppb |  |  | 6.8 ppb | +- | 10.19 ppb | <8.99 ppb |  |  | <10.21 ppb | [5 empty] | 21.64 ppb | +- | 6.33 ppb
SNO+ CW64 | Sample Prepared by D.C and B.H on 190418 | 8.653 g | 6.878 | 190418 | 2019-04-18 | (mBq/kg) | <76.66 |  |  | <43.30 |  |  | <12.10 |  |  | 8436 | +- | 440.8 | <88.43 |  |  | <305.80 |  |  | <176.00
Dowex Effluent |  |  |  |  | 2019-04-25 | (ppb or ppm) | <6.21 ppb |  |  | <3.51 ppb |  |  | <21.3 ppb |  |  | 2075.26 ppb | +- | 108.44 ppb | <21.75 ppb | [5 empty] | <14.26 ppb
SNO+ CW65 | Sample Prepared on 190708 | 13.5 g | 7.966 | 190826 | 2019-08-26 | (mBq/kg) | <20.68 |  |  | <22.16 |  |  | 10.17 | +- | 4.068 | <21.39 |  |  | 7.498 | +- | 34.02 | <201.90 |  |  | 215.53 | +- | 51.38
Supernatant-LWA | Sample Labeled LWA-190708-Blank |  |  |  | 2019-09-03 | (ppb or ppm) | <1.68 ppb |  |  | <1.79 ppb |  |  | 17.9 ppb | +- | 7.16 ppb | <5.26 ppb |  |  | 1.84 ppb | +- | 8.37 ppb |  |  |  | 17.46 ppb | +- | 4.16 ppb
SNO+ CW66 | Sample Prepared on 190719 | 13.3 g | 3.023 | 190923 | 2019-09-23 | (mBq/kg) | <30.59 |  |  | 59.61 | +- | 38.21 | 1.193 | +- | 3.039 | <27.94 |  |  | <80.59 |  |  | <1402.00 |  |  | 34539 | +- | 1883
Dowex Resin | Prepared after the EDTA elution for the PSUP assay |  |  |  | 2019-09-26 | (ppb or ppm) | <2.48 ppb |  |  | 4.83 ppb | +- | 3.1 ppb | 2.1 ppb | +- | 5.35 ppb | <6.87 ppb |  |  | <19.83 ppb | [5 empty] | 2797.66 ppb | +- | 152.52 ppb
SNO+ CW67 | Sample Prepared on 190722 | 15.2 g | 6.781 | 190916 | 2019-09-16 | (mBq/kg) | 6.513 | +- | 22.7 | <24.88 |  |  | <1.55 |  |  | <24.81 |  |  | <40.58 |  |  | <244.40 |  |  | 11672 | +- | 643.4
Supernatant-LWA-190719 | Sample from PSUP using HtiO Technique |  |  |  | 2019-09-23 | (ppb or ppm) | 0.53 ppb | +- | 1.84 ppb | <2.02 ppb |  |  | <2.73 ppb |  |  | <6.1 ppb |  |  | <9.98 ppb | [5 empty] | 945.43 ppb | +- | 52.12 ppb
SNO+ CW68 | Sample was eluted from QuadraSil using DMSA, eluate was boiled down and transferred into the sample vial. | 8.896 g | 6.954 | 191108 | 2019-11-08 | (mBq/kg) | 71.15 | +- | 44.2 | <29.05 |  |  | <4.15 |  |  | <34.45 |  |  | <55.57 |  |  | <407.30 |  |  | 1625 | +- | 162.8
DMSA Eluent Pb-210 (liquid with precipitate) |  | Prep. On 2019-10-29 |  |  | 2019-11-15 | (ppb or ppm) | 5.76 ppb | +- | 3.58 ppb | <2.35 ppb |  |  | <7.3 ppb |  |  | <8.47 ppb |  |  | <13.67 ppb | [5 empty] | 131.63 ppb | +- | 13.19 ppb
SNO+ CW69 | Sample was eluted from QuadraSil using 1M HCl, eluate was boiled down and transferred into the sample vial. | 8.816 g | 6.926 | 191115 | 2019-11-15 | (mBq/kg) | 164.6 | +- | 52.46 | 188.4 | +- | 42.56 | 9.306 | +- | 3.769 | 80.82 | +- | 40.81 | 119 | +- | 78.86 | <412.60 |  |  | 6961.5 | +- | 440.1
1M HCl eluent (liquid/acidic/ Precipitate) |  | Prep. On 2019-10-29 |  |  | 2019-11-22 | (ppb or ppm) | 13.33 ppb | +- | 4.25 ppb | 15.26 ppb | +- | 3.45 ppb | 16.38 ppb | +- | 6.63 ppb | 19.88 ppb | +- | 10.04 ppb | 29.27 ppb | +- | 19.4 ppb |  |  |  | 563.88 ppb | +- | 35.65 ppb
SNO+ CW70 | Sample removed from column | 8.914 g | 5.699 | 191122 191128 | 2019-11-22 | (mBq/kg) | 178.4 | +- | 55.72 | 20.69 | +- | 29.67 | <4.67 |  |  | <36.68 |  |  | <80.67 |  |  | <495.40 |  |  | 795.86 | +- | 122.4
QuadraSil-Top after extraction (beads) |  | Prep. On 2019-10-24 |  |  | 2019-11-29 | (ppb or ppm) | 14.45 ppb | +- | 4.51 ppb | 1.68 ppb | +- | 2.4 ppb | <8.22 ppb |  |  | <9.02 ppb |  |  | <19.84 ppb | [5 empty] | 64.46 ppb | +- | 9.91 ppb
SNO+ CW71 | Sample was removed from column and transferred into vial | 9.475 g | 11.836 | 191129 | 2019-11-29 | (mBq/kg) | 39.62 | +- | 26.92 | 33.81 | +- | 17.34 | <1.67 |  |  | 21.18 | +- | 29.36 | <53.65 |  |  | <226.00 |  |  | 50.425 | +- | 54.07
QuadraSil- Bottom after extraction (Beads) | Prep. on 2019-10-27 |  |  |  | 2019-12-11 | (ppb or ppm) | 3.21 ppb | +- | 2.18 ppb | 2.74 ppb | +- | 1.4 ppb | <2.94 ppb |  |  | 5.21 ppb | +- | 7.22 ppb | <13.2 ppb | [5 empty] | 4.08 ppb | +- | 4.38 ppb
SNO+ CW72 | System flushed and sample collected in vial. | 7.920 g | 7.876 | 191211 | 2019-12-11 | (mBq/kg) | <19.68 |  |  | 28 | +- | 23.87 | <5.10 |  |  | <29.09 |  |  | <91.30 |  |  | <404.30 |  |  | 101.44 | +- | 63.95
UPW Supply via V-563L to MHAD | Prep on 2019-10-16 |  |  |  | 2019-12-19 | (ppb or ppm) | <1.59 ppb |  |  | 2.27 ppb | +- | 1.93 ppb | <8.98 ppb |  |  | <7.16 ppb |  |  | <22.46 ppb | [5 empty] | 8.22 ppb | +- | 5.18 ppb
SNO+ CW73 | After preparing the sample in the clean lab on surface, the outside of the counting vial was methanol wiped. Prep on 2019-10-31 | 8.2 g | 7.752 | 191219 | 2019-12-19 | (mBq/kg) | <44.66 |  |  | 5.788 | +- | 23.68 | <4.46 |  |  | <52.21 |  |  | 8.486 | +- | 55.18 | 196.1 | +- | 241.4 | 727.09 | +- | 111.1
Dowex after the EDTA elution (solid sample) |  |  |  |  | 2020-01-06 | (ppb or ppm) | <3.62 ppb |  |  | 0.47 ppb | +- | 1.92 ppb | <7.85 ppb |  |  | <12.84 ppb |  |  | 2.09 ppb | +- | 13.57 ppb |  |  |  | 58.89 ppb | +- | 9 ppb
SNO+ CW74 | Prepared in the clean lab on surface. The outside of the vial was methanol wiped. | 15.4 g | 6.878 | 200106 | 2020-01-06 | (mBq/kg) | 21.12 | +- | 22.86 | 10.49 | +- | 13.65 | <1.61 |  |  | <23.42 |  |  | 24.59 | +- | 31.44 | <204.70 |  |  | 104.77 | +- | 47.38
Supernatant sample from a blank assay (191031-BLK) | Prep on 2019-10-31 |  |  |  | 2020-01-13 | (ppb or ppm) | 1.71 ppb | +- | 1.85 ppb | 0.85 ppb | +- | 1.11 ppb | <2.83 ppb |  |  | <5.76 ppb |  |  | 6.05 ppb | +- | 7.73 ppb |  |  |  | 8.49 ppb | +- | 3.84 ppb
SNO+ CW75 | Prep on 2019-12-02 | 13.0 g | 6.878 | 200113 | 2020-01-13 | (mBq/kg) | <11.39 |  |  | <25.45 |  |  | 0.713 | +- | 1.801 | <20.19 |  |  | <32.53 |  |  | <281.80 |  |  | 1541.8 | +- | 132.3
Supernatant (in solid form) from PSUP Assay |  |  |  |  | 2020-01-20 | (ppb or ppm) | <0.92 ppb |  |  | <2.06 ppb |  |  | 1.25 ppb | +- | 3.17 ppb | <4.97 ppb |  |  | <8 ppb | [5 empty] | 124.89 ppb | +- | 10.72 ppb
SNO+ CW76 | Prep on 2019-11-28 | 8.472 g | 6.871 | 200120 | 2020-01-20 | (mBq/kg) | 18 | +- | 34.42 | <40.41 |  |  | <2.83 |  |  | <30.72 |  |  | <81.30 |  |  | <432.80 |  |  | 9989.6 | +- | 606.3
Dowex after EDTA from PSUP Assay |  |  |  |  | 2020-01-27 | (ppb or ppm) | 1.46 ppb | +- | 2.79 ppb | <3.27 ppb |  |  | <4.98 ppb |  |  | <7.56 ppb |  |  | <20 ppb | [5 empty] | 809.16 ppb | +- | 49.11 ppb
SNO+ CW77 | Prep on 2020-01-16 | 8.4 g | 7.883 | 200127 | 2020-01-27 | (mBq/kg) | 11.92 | +- | 27.98 | 14.49 | +- | 21.29 | 1.093 | +- | 2.458 | <34.18 |  |  | 47.12 | +- | 51.73 | <380.80 |  |  | 550.92 | +- | 94.35
Dowex after EDTA elution | From a blank assay on January 16, 2020 |  |  |  | 2020-02-04 | (ppb or ppm) | 0.97 ppb | +- | 2.27 ppb | 1.17 ppb | +- | 1.72 ppb | 1.92 ppb | +- | 4.33 ppb | <8.41 ppb |  |  | 11.59 ppb | +- | 12.73 ppb |  |  |  | 44.62 ppb | +- | 7.64 ppb
SNO+ CW78 | Prep on 2020-01-16 | 14.5 g | 8.833 | 200204 | 2020-02-04 | (mBq/kg) | 28.87 | +- | 19.43 | 14.17 | +- | 11.3 | 1.31 | +- | 1.619 | <20.37 |  |  | <24.61 |  |  | <197.10 |  |  | 167.85 | +- | 47.48
Supernatant Solution from Blank assay on 2020-01-16 | From a blank assay on January 16, 2020 |  |  |  | 2020-02-13 | (ppb or ppm) | 2.34 ppb | +- | 1.57 ppb | 1.15 ppb | +- | 0.92 ppb | 2.31 ppb | +- | 2.85 ppb | <5.01 ppb |  |  | <6.05 ppb | [5 empty] | 13.6 ppb | +- | 3.85 ppb
SNO+ CW79 | Prep on 2020-03-06 by D.C. | 7.1 g | 21.287 | 200519 | 2020-05-19 | (mBq/kg) | 441 | +- | 53.33 | <21.03 |  |  | <2.23 |  |  | <38.63 |  |  | <76.04 |  |  | 41.593 | +- | 104 | 7730.4 | +- | 444
Dowex Resin after elution with EDTA |  |  |  |  | 2020-06-10 | (ppb or ppm) | 35.72 ppb | +- | 4.32 ppb | <1.7 ppb |  |  | <3.92 ppb |  |  | <9.5 ppb |  |  | <18.71 ppb | [5 empty] | 626.16 ppb | +- | 35.96 ppb
SNO+ CW80 | Prep on 2020-03-13 by S.R. | 11.5 g | 19.647 | 200610 200624 | 2020-06-10 | (mBq/kg) | <17.95 |  |  | 35.95 | +- | 16.04 | 0.094 | +- | 1.347 | <28.61 |  |  | <18.69 |  |  | 0 | +- | 69.26 | 33766 | +- | 1765
Supernatant PSUP Assay |  |  |  |  | 2020-06-30 | (ppb or ppm) | <1.45 ppb |  |  | 2.91 ppb | +- | 1.3 ppb | 0.17 ppb | +- | 2.37 ppb | <7.04 ppb |  |  | <4.6 ppb | [5 empty] | 2735.05 ppb | +- | 142.97 ppb
SNO+ CW81 | Prep on 2020-12-02 by D.F. | 10.7 g | 9.582 | 210517 210524 | 2021-05-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Supernatant PSUP Assay | For Assay completed on 2020-11-27 |  |  |  | 2021-05-27 | (mBq/kg) | <8.42 |  |  | 68.71 | +- | 38.7 | <1.59 |  |  | 9.657 | +- | 5.613 | 1183.6 | +- | 310 | <35.07 |  |  | <17.57
[6 empty] | (ppb or ppm) | <0.68 ppb |  |  | 5.57 ppb | +- | 3.13 ppb | <2.8 ppb |  |  | 2.38 ppb | +- | 1.38 ppb | 38.23 ppm | +- | 10.01 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 8070.8 | +- | 566.7 | <66.11 |  |  | <5.91 |  |  | <21.96
[6 empty] | (ppb or ppm) | 653.73 ppb | +- | 45.9 ppb | [6 empty] | <5.4 ppb
SNO+ CW82 | Prep on 2021-05-31 by D.F. | 12.6 g | 11.69 | 210611 | 2021-06-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Supernatant Blank Assay | For Blank Assay completed on 2021-05-25 |  |  |  | 2021-06-23 | (mBq/kg) | <2.65 |  |  | 121.2 | +- | 41.48 | 1.384 | +- | 0.843 | 0.7439 | +- | 4.429 | 1133.3 | +- | 326.8 | 2.8854 | +- | 21.48 | <13.60
[6 empty] | (ppb or ppm) | <0.21 ppb |  |  | 9.82 ppb | +- | 3.36 ppb | 2.44 ppb | +- | 1.48 ppb | 0.18 ppb | +- | 1.09 ppb | 36.61 ppm | +- | 10.56 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <189.80 |  |  | <23.15 |  |  | <38.21 |  |  | <10.89
[6 empty] | (ppb or ppm) | <15.37 ppb | [8 empty] | <2.68 ppb
SNO+ CW83 | Prep on 2021-06-21 by S.H. | 5.3 g | 6.922 | 210702 | 2021-07-02 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Supernatant Blank Assay | For Blank Assay completed on 2021-06-18 |  |  |  | 2021-07-09 | (mBq/kg) | <7.49 |  |  | 484.3 | +- | 145.1 | <2.71 |  |  | <10.27 |  |  | <1345.00 |  |  | <73.54 |  |  | 14.4 | +- | 28.83
[6 empty] | (ppb or ppm) | <0.61 ppb |  |  | 39.23 ppb | +- | 11.75 ppb | <4.77 ppb |  |  | <2.53 ppb |  |  | <43.44 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1961.9 | +- | 923.1 | <143.70 |  |  | <22.62 |  |  | <13.17
[6 empty] | (ppb or ppm) | 158.91 ppb | +- | 74.77 ppb | [6 empty] | <3.24 ppb
SNO+ CW84 | Prep on 2021-09-20 by S.R. | 14.8 g | 33.383 | 210924 211001 211004 | 2021-09-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Supernatant PSUP Assay | For PSUP Assay completed on 2021-09-17 |  |  |  | 2021-10-29 | (mBq/kg) | 46.83 | +- | 3.342 | <25.78 |  |  | <1.20 |  |  | 11.12 | +- | 2.698 | 779.95 | +- | 142.8 | <28.00 |  |  | <4.04
[6 empty] | (ppb or ppm) | 3.79 ppb | +- | 0.27 ppb | <2.09 ppb |  |  | <2.11 ppb |  |  | 2.74 ppb | +- | 0.66 ppb | 25.19 ppm | +- | 4.61 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1659 | +- | 159.3 | 6.78 | +- | 16.37 | <3.80 |  |  | 16.06 | +- | 4.719
[6 empty] | (ppb or ppm) | 134.38 ppb | +- | 12.9 ppb | [6 empty] | 3.95 ppb | +- | 1.16 ppb
SNO+ CW85 | Prep on 2022-02-07 by Chanel and Jessica | 7.0 g | 12.976 | 220304 220311 | 2022-03-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Supernatant PSUP Assay | For PSUP assay completed on 2202-02-04 |  |  |  | 2022-03-18 | (mBq/kg) | <9.18 |  |  | <48.63 |  |  | 2.143 | +- | 1.335 | <12.30 |  |  | 1732.4 | +- | 406.3 | <57.53 |  |  | <12.85
[6 empty] | (ppb or ppm) | <0.74 ppb |  |  | <3.94 ppb |  |  | 3.77 ppb | +- | 2.35 ppb | <3.03 ppb |  |  | 55.96 ppm | +- | 13.12 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2359.4 | +- | 383.5 | <55.12 |  |  | <17.06 |  |  | 19.59 | +- | 13.43
[6 empty] | (ppb or ppm) | 191.11 ppb | +- | 31.06 ppb | [6 empty] | 4.82 ppb | +- | 3.3 ppb
SNO+ CW86 | Prep on 2022-02-28 by Chanel and Jessica | 10.0 g | 13.813 | 220318 | 2022-03-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Supernatant Blank Assay | For Blank assay completed on 2202-02-25 |  |  |  | 2022-04-01 | (mBq/kg) | <3.94 |  |  | <47.06 |  |  | <2.20 |  |  | 4.529 | +- | 4.499 | 829.44 | +- | 240.3 | <38.67 |  |  | <5.94
[6 empty] | (ppb or ppm) | <0.32 ppb |  |  | <3.81 ppb |  |  | <3.87 ppb |  |  | 1.11 ppb | +- | 1.11 ppb | 26.79 ppm | +- | 7.76 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <160.20 |  |  | 38.834 | +- | 30.28 | <7.82 |  |  | 14.23 | +- | 9.498
[6 empty] | (ppb or ppm) | <12.98 ppb | [8 empty] | 3.5 ppb | +- | 2.34 ppb
SNO+ CW87 | Prep on 2022-06-08 by Jessica | 12.2 g | 12.22 | 220628 220701 220707 22070701 22070706 | 2022-06-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Supernatant UPW Plant Assay | For UPW Plant assay completed on 2202-06-03 |  |  |  | 2022-07-12 | (mBq/kg) | 8.192 | +- | 4.018 | <28.48 |  |  | <1.24 |  |  | 5.808 | +- | 4.098 | 1128.6 | +- | 253 | <33.23 |  |  | <6.07
[6 empty] | (ppb or ppm) | 0.66 ppb | +- | 0.33 ppb | <2.31 ppb |  |  | <2.18 ppb |  |  | 1.43 ppb | +- | 1.01 ppb | 36.45 ppm | +- | 8.17 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 61.252 | +- | 196.1 | 40.478 | +- | 25.71 | <2.87 |  |  | 20.79 | +- | 8.274
[6 empty] | (ppb or ppm) | 4.96 ppb | +- | 15.88 ppb | [6 empty] | 5.11 ppb | +- | 2.04 ppb
SNO+ CW88 | Prep on 2022-06-09 by Chanel | 12.3 g | 12.724 | 220615 220621 | 2022-06-15 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Supernatant Blank Assay | For Blank assay completed on 2202-06-08 |  |  |  | 2022-06-28 | (mBq/kg) | 3.124 | +- | 3.843 | <18.81 |  |  | 0.642 | +- | 0.7719 | 6.79 | +- | 4.146 | 1323.1 | +- | 263.6 | <34.67 |  |  | <4.71
[6 empty] | (ppb or ppm) | 0.25 ppb | +- | 0.31 ppb | <1.52 ppb |  |  | 1.13 ppb | +- | 1.36 ppb | 1.67 ppb | +- | 1.02 ppb | 42.74 ppm | +- | 8.51 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <173.20 |  |  | <34.48 |  |  | <9.33 |  |  | <14.32
[6 empty] | (ppb or ppm) | <14.03 ppb | [8 empty] | <3.52 ppb
SNO+ CW89 | Prep on 2023-10-18 by Sharayah/Sahima | 9.29 g | 20.442 | 231026 23102601 231028 231029 231030 231031 231101 231102 231103 | 2023-10-26 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Supernatant Blank Assay | For Blank assay completed on 2023-10-17 |  |  |  | 2023-11-16 | (mBq/kg) | <4.44 |  |  | <28.97 |  |  | 0.3792 | +- | 0.8067 | 6.093 | +- | 4.507 | 105.54 | +- | 164.7 | <44.29 |  |  | <5.75
[6 empty] | (ppb or ppm) | <0.36 ppb |  |  | <2.35 ppb |  |  | 0.67 ppb | +- | 1.42 ppb | 1.5 ppb | +- | 1.11 ppb | 3.41 ppm | +- | 5.32 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <314.50 |  |  | <53.13 |  |  | <10.45 |  |  | 20.8 | +- | 8.77
[6 empty] | (ppb or ppm) | <25.47 ppb | [8 empty] | 5.12 ppb | +- | 2.16 ppb
SNO+ CW90 | Prep on 2023-10-30 by Sahima | 12.1 g | 12.752 | 231116 | 2023-11-16 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Supernatant UG UPW Water System | For UG UPW assay completed on 2023-10-27 |  |  |  | 2023-11-29 | (mBq/kg) | <3.81 |  |  | <26.83 |  |  | <1.32 |  |  | 8.861 | +- | 4.318 | 471.24 | +- | 183.2 | <31.58 |  |  | <7.21
[6 empty] | (ppb or ppm) | <0.31 ppb |  |  | <2.17 ppb |  |  | <2.32 ppb |  |  | 2.18 ppb | +- | 1.06 ppb | 15.22 ppm | +- | 5.92 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1696.8 | +- | 208.4 | <61.60 |  |  | <7.38 |  |  | 12.05 | +- | 7.881
[6 empty] | (ppb or ppm) | 137.44 ppb | +- | 16.88 ppb | [6 empty] | 2.96 ppb | +- | 1.94 ppb
SNO+ CW91 | Prep on 2024-02-09 by D.C. | 7.934 g | 7.778 | 240328 | 2024-03-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Dowex after EDTA elution | For Assay completed on 2024-02-09 |  |  |  | 2024-04-05 | (mBq/kg) | <4.22 |  |  | <59.24 |  |  | <1.87 |  |  | <13.46 |  |  | 606.42 | +- | 349.4 | <45.96 |  |  | <20.75
[6 empty] | (ppb or ppm) | <0.34 ppb |  |  | <4.8 ppb |  |  | <3.29 ppb |  |  | <3.31 ppb |  |  | 19.59 ppm | +- | 11.29 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 9109.2 | +- | 651.6 | <48.09 |  |  | <14.81 |  |  | 21.46 | +- | 13.99
[6 empty] | (ppb or ppm) | 737.85 ppb | +- | 52.78 ppb | [6 empty] | 5.28 ppb | +- | 3.44 ppb
SNO+ CW92 | Prep on 2024-02-09 by Sahima and Sharayah | 13.6 g | 5.87 | 240405 | 2024-04-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Supernatant UG UPW Water System | For UG UPW assay completed on 2024-02-09 |  |  |  | 2024-04-11 | (mBq/kg) | <7.96 |  |  | <59.09 |  |  | 1.695 | +- | 1.233 | <8.70 |  |  | 683.9 | +- | 273.7 | <41.76 |  |  | <11.91
[6 empty] | (ppb or ppm) | <0.64 ppb |  |  | <4.79 ppb |  |  | 2.98 ppb | +- | 2.17 ppb | <2.14 ppb |  |  | 22.09 ppm | +- | 8.84 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2304.5 | +- | 279.8 | <57.47 |  |  | <3.57 |  |  | <21.49
[6 empty] | (ppb or ppm) | 186.66 ppb | +- | 22.66 ppb | [6 empty] | <5.29 ppb
PICO Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
PICO CW01 | Small Diameter Piezos from December Production | 10.6 g | 15.532 | 140107 14010701 14010703 14010801 14010802 14010804 140115 | 2014-01-07 | (mBq/kg) | 1360.17 | +- | 68.28 | 5.35 | +- | 45.45 | 21.55 | +- | 3.76 | 242.89 | +- | 28.02 | 141.82 | +- | 82.07 | <63.85 |  |  | 1783.02 | +- | 103.77
Piezo Acoustic Sensors |  | Also PICO 3 |  |  | 2014-01-23 | (ppb or ppm) | 110.17 ppb | +- | 5.53 ppb | 0.43 ppb | +- | 3.68 ppb | 37.94 ppb | +- | 6.62 ppb | 59.75 ppb | +- | 6.89 ppb | 34.89 ppb | +- | 20.19 ppb
PICO CW02 |  | 14.8 g | 8.528 | 140617 | 2014-06-17 | (mBq/kg) | 30.86 | +- | 25.73 | <15.05 |  |  | 4.44 | +- | 5.9 | 126.66 | +- | 20.01 | 17.82 | +- | 44.79 | <118.15 |  |  | 513.51 | +- | 60.81
Small Diameter Piezo Acoustic Sensors (11) | May/June 2014 Production (6/16/2014) | Also PICO 10 |  |  | 2014-06-26 | (ppb or ppm) | 2.50 ppb | +- | 2.08 ppb | 1.22 ppb |  |  | 7.82 ppb | +- | 10.39 ppb | 31.26 ppb | +- | 4.92 ppb | 4.38 ppb | +- | 11.02 ppb
PICO CW03 | These piezos are unelectroded | 10.6 g | 5.958 | 140626 | 2014-06-26 | (mBq/kg) | <49.93 |  |  | 2.83 | +- | 17.9 | 6.13 | +- | 8.26 | 10.11 | +- | 18.75 | 119.25 | +- | 74.03 | <331.87 |  |  | 528.3 | +- | 84.91
Small Diameter Piezo Acoustic Sensors (8) | May/June 2014 Production (6/16/2014) |  |  |  | 2014-07-02 | (ppb or ppm) | <3.48 ppb |  |  | 0.23 ppb | +- | 1.45 ppb | 10.79 ppb | +- | 14.53 ppb | 2.49 ppb | +- | 4.61 ppb | 29.34 ppb | +- | 18.21 ppb
PICO CW04 |  | 4.0 g | 5.924 | 140702 | 2014-07-02 | (mBq/kg) | 111.38 | +- | 78.54 | 32.69 | +- | 41.6 | 25.51 | +- | 12.07 | 45.78 | +- | 49.62 | 25.96 | +- | 155.38 | <632.00 |  |  | 75 | +- | 125
Distilled LAB |  |  |  |  | 2014-07-08 | (ppb or ppm) | 9.02 ppb | +- | 6.36 ppb | 2.65 ppb | +- | 3.37 ppb | 44.89 ppb | +- | 21.34 ppb | 11.26 ppb | +- | 12.21 ppb | 6.39 ppb | +- | 38.22 ppb
PICO CW05 | Blue | 19.2 g | 2.861 | 150514 | 2015-05-14 | (mBq/kg) | 781.913 | +- | 103.064 | 867.094 | +- | 95.4 | 80.071 | +- | 16.503 | 2870.069 | +- | 173.144 | 1325.247 | +- | 166.498 | <453.85 |  |  | 703.281 | +- | 180.469
PVC-Clad Wire (Hook up wire) | Alpha Wire |  |  |  | 2015-05-17 | (ppb or ppm) | 63.34 ppb | +- | 8.35 ppb | 70.24 ppb | +- | 7.72 ppb | 140.93 ppb | +- | 29.04 ppb | 706.04 ppb | +- | 42.59 ppb | 326.01 ppb | +- | 40.96 ppb |  |  |  | 56.97 ppb | +- | 14.62 ppb
PICO CW06 | White | 19.3 g | 2.979 | 150517 | 2015-05-17 | (mBq/kg) | 696.445 | +- | 94.875 | 869.438 | +- | 98.115 | 95.459 | +- | 16.574 | 2896.707 | +- | 178.268 | 1511.974 | +- | 171.319 | <599.22 |  |  | 758.171 | +- | 177.668
PVC-Clad Wire (Hook up wire) | Alpha Wire |  |  |  | 2015-05-20 | (ppb or ppm) | 56.41 ppb | +- | 7.69 ppb | 70.42 ppb | +- | 7.95 ppb | 168.01 ppb | +- | 29.17 ppb | 712.59 ppb | +- | 43.85 ppb | 371.95 ppb | +- | 42.15 ppb |  |  |  | 61.49 ppb | +- | 14.39 ppb
PICO CW07 | Orange | 20.0 g | 2.049 | 150520 | 2015-05-20 | (mBq/kg) | 910.606 | +- | 108.617 | 644.843 | +- | 104.649 | 75.155 | +- | 18.816 | 2869.188 | +- | 185.901 | 1609.232 | +- | 202.302 | <725.80 |  |  | 1236.8 | +- | 214.7
PVC-Clad Wire (Hook up wire) | Alpha Wire |  |  |  | 2015-05-22 | (ppb or ppm) | 73.76 ppb | +- | 8.80 ppb | 52.23 ppb | +- | 8.58 ppb | 132.27 ppb | +- | 33.12 ppb | 705.82 ppb | +- | 45.73 ppb | 395.87 ppb | +- | 49.77 ppb |  |  |  | 100.18 ppb | +- | 17.39 ppb
PICO CW08 | Red | 18.9 g | 2.826 | 150522 | 2015-05-22 | (mBq/kg) | 650.437 | +- | 83.781 | 608.418 | +- | 69.844 | 65.688 | +- | 13.53 | 1288.585 | +- | 98.353 | 596.94 | +- | 118.575 | <645.29 |  |  | 1521.534 | +- | 182.593
PVC-Clad Wire (Hook up wire) | Alpha Wire |  |  |  | 2015-05-25 | (ppb or ppm) | 52.69 ppb | +- | 6.79 ppb | 49.28 ppb | +- | 5.66 ppb | 115.61 ppb | +- | 23.82 ppb | 316.99 ppb | +- | 24.20 ppb | 146.85 ppb | +- | 29.17 ppb |  |  |  | 123.24 ppb | +- | 14.79 ppb
PICO CW09 | Green | 19.1 g | 2 | 150525 150527 | 2015-05-25 | (mBq/kg) | 1118.856 | +- | 134.798 | 990.364 | +- | 118.132 | 176.697 | +- | 23.426 | 2540.166 | +- | 178.119 | 771.024 | +- | 185.835 | <778.59 |  |  | 2101.466 | +- | 274.974
PVC-Clad Wire (Hook up wire) | Alpha Wire |  |  |  | 2015-05-27 | (ppb or ppm) | 90.63 ppb | +- | 10.92 ppb | 80.22 ppb | +- | 9.57 ppb | 310.99 ppb | +- | 41.23 ppb | 624.88 ppb | +- | 43.82 ppb | 189.67 ppb | +- | 45.72 ppb |  |  |  | 170.22 ppb | +- | 22.27 ppb
PICO CW10 | Black | 19.0 g | 1.979 | 15052701 | 2015-05-27 | (mBq/kg) | 885.118 | +- | 121.2 | 956.865 | +- | 111.45 | 94.44 | +- | 19.786 | 2675.292 | +- | 180.311 | 1273.23 | +- | 194.822 | <790.95 |  |  | 1561.053 | +- | 232.789
PVC-Clad Wire (Hook up wire) | Alpha Wire |  |  |  | 2015-05-29 | (ppb or ppm) | 71.70 ppb | +- | 9.82 ppb | 77.51 ppb | +- | 9.03 ppb | 166.21 ppb | +- | 34.82 ppb | 658.12 ppb | +- | 44.36 ppb | 313.22 ppb | +- | 47.93 ppb |  |  |  | 126.44 ppb | +- | 18.86 ppb
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
PICO CW11 | Brown | 19.0 g | 9.847 | 150529 150530 | 2015-05-29 | (mBq/kg) | 668.418 | +- | 53.158 | 778.875 | +- | 67.499 | 82.636 | +- | 9.593 | 2454.283 | +- | 134.919 | 1311.822 | +- | 98.362 | <229.05 |  |  | 925.211 | +- | 107.263
PVC-Clad Wire (Hook up wire) | Alpha Wire |  |  |  | 2015-06-08 | (ppb or ppm) | 54.14 ppb | +- | 4.31 ppb | 62.85 ppb | +- | 5.47 ppb | 145.44 ppb | +- | 16.88 ppb | 603.875 ppb | +- | 33.19 ppb | 322.71 ppb | +- | 24.20 ppb |  |  |  | 74.94 ppb | +- | 8.69 ppb
PICO CW12 |  | 3.3 g | 4.069 | 150608 | 2015-06-08 | (mBq/kg) | 90.69 | +- | 114.299 | <77.89 |  |  | <30.00 |  |  | 96.053 | +- | 65.482 | <386.36 |  |  | <2556.67 |  |  | 2428.485 | +- | 362.121
Cable Inner Wires | ECOGen Cables by Alpha Wire |  |  |  | 2015-06-12 | (ppb or ppm) | 7.35 ppb | +- | 9.26 ppb | <6.31 ppb |  |  | <52.80 ppb |  |  | 23.63 ppb | +- | 16.11 ppb | <95.05 ppb | [5 empty] | 196.71 ppb | +- | 29.33 ppb
PICO CW13 | Grey plastic-type outer shell | 3.6 g | 4.883 | 150612 | 2015-06-12 | (mBq/kg) | 1624.352 | +- | 236.317 | 995.099 | +- | 171.711 | 116.812 | +- | 32.073 | 2634.546 | +- | 233.45 | 1577.029 | +- | 309.676 | 440.833 | +- | 1217.5 | 1845.556 | +- | 393.889
Cable Outer Shield | ECOGen Cables by Alpha Wire |  |  |  | 2015-06-17 | (ppb or ppm) | 131.57 ppb | +- | 19.14 ppb | 80.60 ppb | +- | 13.91 ppb | 205.60 ppb | +- | 56.45 ppb | 648.10 ppb | +- | 57.43 ppb | 387.95 ppb | +- | 76.18 ppb |  |  |  | 149.49 ppb | +- | 31.91 ppb
PICO CW14 | Metal Insulator / Grounding Jacket | 5.4 g | 5.542 | 150617 150618 15061801 150619 | 2015-06-17 | (mBq/kg) | 32.003 | +- | 69.062 | 487.736 | +- | 71.158 | 54.83 | +- | 13.481 | 99.562 | +- | 39.389 | 113.105 | +- | 90.774 | <1142..04 |  |  | 938.889 | +- | 179.074
Cable Outer insulator | ECOGen Cables by Alpha Wire |  |  |  | 2015-06-23 | (ppb or ppm) | 2.59 ppb | +- | 5.59 ppb | 39.51 ppb | +- | 5.76 ppb | 96.50 ppb | +- | 23.73 ppb | 24.49 ppb | +- | 9.69 ppb | 27.82 ppb | +- | 22.33 ppb |  |  |  | 76.05 ppb | +- | 14.51 ppb
PICO CW15 |  | 4.6 g | 4.861 | 150623 | 2015-06-23 | (mBq/kg) | 252.386 | +- | 86.209 | 60.32 | +- | 49.537 | 31.194 | +- | 14.29 | 184.425 | +- | 58.689 | 25.09 | +- | 117.576 | <1531.52 |  |  | 516.087 | +- | 171.522
Thick Diameter Cable Inner Wire | ECOGen Cables by Alpha Wire |  |  |  | 2015-06-28 | (ppb or ppm) | 20.44 ppb | +- | 6.98 ppb | 4.89 ppb | +- | 4.01 ppb | 54.90 ppb | +- | 25.12 ppb | 45.37 ppb | +- | 14.44 ppb | 6.17 ppb |  | 28.92 ppb |  |  |  | 41.80 ppb | +- | 13.89 ppb
PICO CW16 |  | 6.1 g | 6 | 150702 | 2015-07-02 | (mBq/kg) | 24.131 | +- | 48.571 | 171.679 | +- | 45.227 | 31.665 | +- | 10.483 | <48.69 |  |  | <116.72 |  |  | <932.46 |  |  | 7896.066 | +- | 535.574
Thick Diameter Cable Outer insulator | ECOGen Cables by Alpha Wire |  |  |  | 2015-07-08 | (ppb or ppm) | 1.96 ppb | +- | 3.93 ppb | 13.91 ppb | +- | 3.66 ppb | 55.73 ppb | +- | 18.45 ppb | <11.98 ppb |  |  | <28.71 ppb | [5 empty] | 639.58 ppb | +- | 43.38 ppb
PICO CW17 |  | 15.8 g | 3.875 | 150628 | 2015-06-28 | (mBq/kg) | 151.766 | +- | 37.997 | 68.357 | +- | 25.227 | 20.28 | +- | 4.408 | 94.182 | +- | 21.743 | <258.92 |  |  | <561.14 |  |  | 639.177 | +- | 101.582
Thick Diameter Cable Outer Shield | ECOGen Cables by Alpha Wire |  |  |  | 2015-07-02 | (ppb or ppm) | 12.29 ppb | +- | 3.10 ppb | 5.54 ppb | +- | 2.04 ppb | 35.69 ppb | +- | 7.76 ppb | 23.17 ppb | +- | 5.35 ppb | <63.70 ppb | [5 empty] | 51.77 ppb | +- | 8.23 ppb
PICO CW18 | Part No. A06000320625S | 5.3 g | 6.91 | 160317 16031701 | 2016-03-17 | (mBq/kg) | 33.429 | +- | 48.234 | 8.354 | +- | 28.495 | 16.438 | +- | 8.951 | <47.17 |  |  | 10.517 | +- | 81.166 | <929.62 |  |  | 41.132 | +- | 101.698
Springs | Lot No. 1076186 Associated Spring Raymond |  |  |  | 2016-03-24 | (ppb or ppm) | 2.71 ppb | +- | 3.91 ppb | 0.68 ppb | +- | 2.31 ppb | 28.93 ppb | +- | 15.75 ppb | <11.60 ppb |  |  | 2.59 ppb | +- | 19.97 ppb |  |  |  | 3.33 ppb | +- | 8.24 ppb
PICO CW19 | Lot No.: 286245-6 8-32 x ½ washer face thumb screw | 17.10 g | 6.924 | 160324 | 2016-03-24 | (mBq/kg) | <13.18 |  |  | 13.495 | +- | 9.725 | 1.181 | +- | 2.508 | <9.94 |  |  | 52.959 | +- | 26.092 | <244.74 |  |  | 67.052 | +- | 29.827
Thumb Screws | Richard Manno Company |  |  |  | 2016-03-31 | (ppb or ppm) | <1.07 ppb |  |  | 1.09 ppb | +- | 0.79 ppb | 2.08 ppb | +- | 4.41 ppb | <2.45 ppb |  |  | 13.03 ppb | +- | 6.42 ppb |  |  |  | 5.43 ppb | +- | 2.42 ppb
PICO CW20 |  | 0.5 g | 12.794 | 160923 160930 | 2016-09-23 | (mBq) | 440.193 | +- | 405.628 | 118.147 | +- | 201.172 | 43.193 | +- | 62.33 | <380.00 |  |  | 198.067 | +- | 567.438 | <5240.00 |  |  | 1882 | +- | 716
Kapton |  |  |  |  | 2016-10-06 | (ppb or ppm) | 35.66 ppb | +- | 32.86 ppb | 9.57 ppb | +- | 16.30 ppb | 76.02 ppb | +- | 109.70 ppb | <93.48 ppb |  |  | 73.33 ppb | +- | 139.59 ppb |  |  |  | 155.44 ppb | +- | 58.00 ppb
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
PICO CW21 |  | 0.4 g | 10.306 | 161006 161010 | 2016-10-06 | (mBq/kg) | <525.00 |  |  | <289.41 |  |  | 23.505 | +- | 84.48 | 116.073 | +- | 337.228 | <4210.00 |  |  | <7052.50 |  |  | 1705 | +- | 962.5
Liquid Crystal Polymer |  |  |  |  | 2016-10-17 | (ppb or ppm) | <42.53 ppb |  |  | <23.44 ppb |  |  | 41.37 ppb | +- | 148.69 ppb | 28.55 ppb | +- | 82.96 ppb | <1.04 ppm | [5 empty] | 138.11 ppb | +- | 77.96 ppb
PICO CW22 | Jar to Flange Seal | 3.2 g | 12.768 | 170125 | 2017-01-25 | (mBq/kg) | 54.99 | +- | 52.649 | <29.15 |  |  | <12.19 |  |  | 10.63 | +- | 42.423 | 100.926 | +- | 92.179 | <820.31 |  |  | 208.75 | +- | 109.688
PTFE Gasket | Hennig Gaskets and Seals |  |  |  | 2017-02-07 | (ppb or ppm) | 4.45 ppb | +- | 4.27 ppb | <2.36 ppb |  |  | <21.45 ppb |  |  | 2.62 ppb | +- | 10.44 ppb | 24.83 ppb | +- | 22.67 ppb |  |  |  | 16.91 ppb | +- | 8.89 ppb
PICO CW23 | PN: 1416-1701-1-ND | 0.4 g | 21.438 | 170315 170316 | 2017-03-15 | (mBq/kg) | 3109.834 | +- | 648.123 | 1076.145 | +- | 348.826 | 98.755 | +- | 83.897 | <510.00 |  |  | 1556.219 | +- | 732.267 | <3325.00 |  |  | 2817.5 | +- | 920
LED Luxeon Rebel Deep Red SMD | Digi-Key | 10 LEDs |  |  | 2017-04-06 | (ppb or ppm) | 251.90 ppb | +- | 54.50 ppb | 87.17 ppb | +- | 28.26 ppb | 173.81 ppb | +- | 147.66 ppb | <125.46 ppb |  |  | 382.83 ppb | +- | 180.14 ppb |  |  |  | 227.41 ppb | +- | 74.52 ppb
PICO CW24 | PN: XPEPHR-L1-0000-00801CT-ND | 0.4 g | 20.826 | 170406 170413 | 2017-04-06 | (mBq/kg) | 404.184 | +- | 442.549 | 684.493 | +- | 309.946 | 200.349 | +- | 80.584 | <442.50 |  |  | 126.357 | +- | 625.641 | <3962.50 |  |  | 2650 | +- | 820
LED Photo Red 650NM 700MA | Digi-Key | 10 LEDs |  |  | 2017-04-27 | (ppb or ppm) | 32.74 ppb | +- | 35.85 ppb | 55.44 ppb | +- | 25.11 ppb | 352.62 ppb | +- | 141.83 ppb | <108.86 ppb |  |  | 31.08 ppb | +- | 153.91 ppb |  |  |  | 214.65 ppb | +- | 66.42 ppb
PICO CW25 | NB PZT Piezeos 5mm diameter and 1 mm thick | 1.8 g | 13.438 | 170427 170429 170502 170506 | 2017-04-27 | (mBq/kg) | 465.903 | +- | 159.808 | <74.44 |  |  | 23.807 | +- | 21.082 | 80.895 | +- | 84.881 | <422.22 |  |  | <1383.89 |  |  | 1286.111 | +- | 322.778
Piezo Acoustic Sensors | April 2017 Batch, 12 piezos |  |  |  | 2017-05-11 | (ppb or ppm) | 37.74 ppb | +- | 12.94 ppb | <6.03 ppb |  |  | 41.90 ppb | +- | 37.10 ppb | 19.90 ppb | +- | 20.88 ppb | <103.87 ppb | [5 empty] | 104.18 ppb | +- | 26.15 ppb
PICO CW26 | NB PZT Piezeos 3 mm diameter and 8.7 mm tall | 10.6 g | 10.722 | 170525 170601 | 2017-05-25 | (mBq/kg) | 6.317 | +- | 22.643 | <18.68 |  |  | 3.832 | +- | 4.387 | <15.66 |  |  | 32.285 | +- | 40.036 | 57.075 | +- | 187.08 | 278.019 | +- | 67.736
Piezo Acoustic Sensors | April 2017 Batch, 17 piezos |  |  |  | 2017-06-05 | (ppb or ppm) | 0.51 ppb | +- | 1.83 ppb | <1.51 ppb |  |  | 6.74 ppb | +- | 7.72 ppb | <3.85 ppb |  |  | 7.94 ppb | +- | 9.85 ppb |  |  |  | 22.52 ppb | +- | 5.49 ppb
PICO CW27 | NB PZT Piezeos 5mm diameter and 1 mm thick | 4.1 g | 13.674 | 170627 | 2017-06-27 | (mBq/kg) | 122.019 | +- | 58.287 | <31.46 |  |  | 8.248 | +- | 8.557 | 67.253 | +- | 35.552 | 98.694 | +- | 81.869 | <596.83 |  |  | 528.049 | +- | 114.634
Piezo Acoustic Sensors | June 2017 Batch 25 piezos |  |  |  | 2017-07-11 | (ppb or ppm) | 9.88 ppb | +- | 2.72 ppb | <2.55 ppb |  |  | 14.52 ppb | +- | 15.06 ppb | 16.54 ppb | +- | 8.75 ppb | 24.28 ppb | +- | 20.14 ppb |  |  |  | 42.77 ppb | +- | 9.29 ppb
PICO CW28 | 3M Scotchlite 8725 | 24.5 g | 5.431 | 170809 170814 | 2017-08-09 | (mBq/kg) | 1409.536 | +- | 93.877 | 0.658 | +- | 51.727 | 97.047 | +- | 13.416 | 599.537 | +- | 51.555 | 561.205 | +- | 112.148 | <221.02 |  |  | 120.653 | +- | 145.755
Reflector Material | 1” wide |  |  |  | 2017-08-15 | (ppb or ppm) | 114.17 ppb | +- | 7.60 ppb | 0.05 ppb | +- | 4.19 ppb | 170.80 ppb | +- | 23.61 ppb | 147.49 ppb | +- | 12.68 ppb | 138.06 ppb | +- | 27.59 ppb |  |  |  | 9.77 ppb | +- | 11.81 ppb
|  | Additional Activities: |  |  |  | (mBq/kg) |  |  | 40K: | 131.61 | +- | 37.82
[6 empty] | (ppb or ppm) |  |  |  | 4.25 ppm | +- | 1.22 ppm
PICO CW29 | 3M Scotchlite 8830 | 14.7 g | 21.276 | 170824 17082401 170826 17082601 17062602 170901 170904 | 2017-08-24 | (mBq/kg) | 815.493 | +- | 49.324 | 69.301 | +- | 26.2 | 54.608 | +- | 6.553 | 300.764 | +- | 26.26 | 248.695 | +- | 52.832 | 41.633 | +- | 130.272 | 327.211 | +- | 79.796
Reflector Material | Also Counted as PICO 70 |  |  |  | 2017-09-15 | (ppb or ppm) | 66.06 ppb | +- | 4.00 ppb | 5.61 ppb | +- | 2.12 ppb | 96.11 ppb | +- | 11.53 ppb | 73.99 ppb | +- | 6.46 ppb | 61.18 ppb | +- | 13.00 ppb |  |  |  | 26.50 ppb | +- | 6.46 ppb
|  | Additional Activities: |  |  |  | (mBq/kg) |  |  | 40K: | 342.89 | +- | 85.28
[6 empty] | (ppb or ppm) |  |  |  | 11.08 ppm | +- | 2.75 ppm
PICO CW30 | NB PZT Piezeos 5mm diameter and 8.7 mm high |  | 11.899 | 171020 | 2017-10-20 | (mBq/kg) | 67.939 | +- | 17.798 | 17.529 | +- | 11.143 | 8.044 | +- | 2.921 | 19.144 | +- | 10.524 | 10.752 | +- | 24.647 | <140.73 |  |  | 584.533 | +- | 58.267
Piezo Acoustic Sensors | Sept 2017 Batch 11 piezos |  |  |  | 2017-11-01 | (ppb or ppm) | 5.50 ppb | +- | 1.44 ppb | 1.42 ppb | +- | 0.90 ppb | 14.16 ppb | +- | 5.14 ppb | 4.71 ppb | +- | 2.60 ppb | 2.65 ppb | +- | 6.06 ppb |  |  |  | 47.35 ppb | +- | 4.72 ppb
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
PICO CW31 | Solder No-Clean 18 AWG 63/37 1 LB Part No: KE1402-ND Mfg PN: 24-6337-8814 | 15.0 g | 4.95 | 180222 | 2018-02-22 | (mBq/kg) | <38.93 |  |  | <67.93 |  |  | ,22.13 |  |  | <31.07 |  |  | <101.60 |  |  | <461.00 |  |  | 22246.533 | +- | 1213.8
Solder Wire | Kester Solder (vendor Digi-Key) |  |  |  | 2018-02-27 | (ppb or ppm) | <3.15 ppb |  |  | <5.50 ppb |  |  | <38.96 ppb |  |  | <7.64 ppb |  |  | <24.99 ppb | [5 empty] | 1.80 ppm | +- | 0.10 ppm
PICO CW32 | Wire Bus Bar 16AWG Part No.: 295 SV005-ND | 14.8 g | 7.81 | 180227 080305 | 2018-02-27 | (mBq/kg) | 14.177 | +- | 16.087 | 12.079 | +- | 13.848 | <4.05 |  |  | 14.852 | +- | 11.339 | <101.62 |  |  | <252.91 |  |  | 2998.581 | +- | 194.595
Tin coated copper wire | Alpha Wire (vendor Digi-Key) |  |  |  | 2018-03-07 | (ppb or ppm) | 1.15 ppb | +- | 1.30 ppb | 0.98 ppb | +- | 1.12 ppb | <7.14 ppb |  |  | 3.65 ppb | +- | 2.79 ppb | <25.00 ppb | [5 empty] | 242.89 ppb | +- | 15.76 ppb
PICO CW33 | 304 Pieces of Fiducial Marks for PICO 40L | 0.122 g | 14.69 | 180628 180629 180710 | 2018-06-28 | (mBq/kg) | 804.567 | +- | 1894.444 | <1270.49 |  |  | 191.347 | +- | 268.835 | 2928.195 | +- | 1220.77 | <7344.26 |  |  | <13122.95 |  |  | 1680.328 | +- | 2713.115
Fiducial Marks |  | 401 ug per mark |  |  | 2018-07-13 | (ppb or ppm) | 65.17 ppb | +- | 153.45 ppb | <102.91 ppb |  |  | 336.77 ppb | +- | 473.15 ppb | 720.34 ppb | +- | 300.31 ppb | <1806.69 ppb | [5 empty] | 136.11 ppb | +- | 219.76 ppb
PICO CW34 | To be used to suspend the copper coils and retro-reflector | 14.1 g | 14.622 | 180713 180717 180730 18073001 | 2018-07-13 | (mBq/kg) | 6.84 | +- | 23.2 | 13.26 | +- | 9.675 | 2.163 | +- | 2.977 | <28.97 |  |  | <34.89 |  |  | <236.70 |  |  | 111.96 | +- | 29.83
316 Stainless Steel Bar | McMaster Carr |  |  |  | 2018-07-30 | (ppb or ppm) | 0.55 ppb | +- | 1.88 ppb | 1.07 ppb | +- | 0.78 ppb | 3.81 ppb | +- | 5.24 ppb | <7.13 ppb |  |  | <8.58 ppb | [5 empty] | 9.07 ppb | +- | 2.42 ppb
PICO CW35 |  | 52.3 g | 26.05 | 18082801 180829 180830 180904 18090401 180905 180920 | 2018-08-28 | (mBq/kg) | 8.387 | +- | 3.865 | 4.153 | +- | 3.086 | 2.396 | +- | 0.8108 | <7.08 |  |  | 4.527 | +- | 6.562 | <18.05 |  |  | 319.39 | +- | 21.44
Piezo Acoustic Sensors | Batch: 2018-07-15 40 piezos | 1.3075 g / Piezo |  |  | 2018-09-25 | (ppb or ppm) | 0.68 ppb | +- | 0.31 ppb | 0.34 ppb | +- | 0.25 ppb | 4.22 ppb | +- | 1.43 ppb | <1.74 ppb |  |  | 1.11 ppb | +- | 1.61 ppb |  |  |  | 25.87 ppb | +- | 1.74 ppb
PICO CW36 |  | 18.1 g | 21.313 | 181113 181124 | 2018-11-13 | (mBq/kg) | 615.8 | +- | 41.17 | <8.80 |  |  | 25.93 | +- | 4.171 | 114.4 | +- | 18.57 | 56.59 | +- | 32.33 | 80.801 | +- | 61.36 | 2121.2 | +- | 128.6
Piezo Acoustic Sensors | Batch: 2018-11-04 14 piezos | 1.293 g / Piezo |  |  | 2018-12-05 | (ppb or ppm) | 49.88 ppb | +- | 3.33 ppb | <0.71 ppb |  |  | 45.64 ppb | +- | 7.34 ppb | 28.14 ppb | +- | 4.57 ppb | 13.92 ppb | +- | 7.95 ppb |  |  |  | 171.82 ppb | +- | 10.42 ppb
PICO CW37 |  | 6.7g | 14.348 | 190402 190413 | 2019-04-02 | (mBq/kg) | <50.00 |  |  | <16.79 |  |  | <7.16 |  |  | <59.61 |  |  | <53.12 |  |  | <264.80 |  |  | 96.93 | +- | 67.03
Acetal Plastic Rod |  |  |  |  | 2019-04-18 | (ppb or ppm) | <4.05 ppb |  |  | <1.36 ppb |  |  | <12.6 ppb |  |  | <14.66 ppb |  |  | <13.07 ppb | [5 empty] | 7.85 ppb | +- | 5.43 ppb
PICO CW38 | 50 Piezos | 69.3 g | 14.677 | 190702 | 2019-07-02 | (mBq/kg) | 15.16 | +- | 4.759 | 2.67 | +- | 3.583 | 3.587 | +- | 0.9855 | 5.351 | +- | 4.099 | <10.97 |  |  | 12.258 | +- | 15.23 | 269.92 | +- | 20.03
Piezo Acoustic Sensors | June 2019 Batch | 1.386 g / piezo |  |  | 2019-07-17 | (ppb or ppm) | 1.23 ppb | +- | 0.39 ppb | 0.22 ppb | +- | 0.29 ppb | 6.31 ppb | +- | 1.73 ppb | 1.32 ppb | +- | 1.01 ppb | <2.7 ppb | [5 empty] | 21.86 ppb | +- | 1.62 ppb
PICO CW39 | Loctite/Henkel | 3.0 g | 7.75 | 220509 | 2022-05-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Lead-free Solder Wire | ID: C511 97SC 3C |  |  |  | 2022-05-17 | (mBq/kg) | <19.57 |  |  | 179.5 | +- | 215.3 | 6.796 | +- | 7.752 | <25.84 |  |  | 352.1 | +- | 977.9 | <138.60 |  |  | <108.30
| Diameter: 0.56 mm |  |  |  |  | (ppb or ppm) | <1.59 ppb |  |  | 14.54 ppb | +- | 17.44 ppb | 11.96 ppb | +- | 13.64 ppb | <6.36 ppb |  |  | 11.37 ppm | +- | 31.59 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 15815 | +- | 1419 | 70.713 | +- | 114.1 | <63.81 |  |  | 82.29 | +- | 67.84
[6 empty] | (ppb or ppm) | 1281.02 ppb | +- | 114.94 ppb | [6 empty] | 20.24 ppb | +- | 16.69 ppb
PICO CW40 | Kester Solder (vendor Digi-Key) | 5.8 g | 13.618 | 220928 22100601 221012 | 2022-09-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Lead-free Solder Wire | Sn: 95.5, Ag: 3.0, Cu: 0.4 |  |  |  | 2022-10-14 | (mBq/kg) | 10.64 | +- | 8.976 | 230.6 | +- | 236.8 | <9.37 |  |  | <23.78 |  |  | 108.14 | +- | 108.6 | <45.48 |  |  | <10.81
| 24-7068-6403 Lot No: 02136567L |  |  |  |  | (ppb or ppm) | 0.86 ppb | +- | 0.73 ppb | 18.68 ppb | +- | 19.18 ppb | <16.49 ppb |  |  | <5.85 ppb |  |  | 3.49 ppm | +- | 3.51 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 375250 | +- | 21320 | <82.19 |  |  | <5.16 |  |  | 39.59 | +- | 15.96
[6 empty] | (ppb or ppm) | 30.4 ppm | +- | 1.73 ppm | [6 empty] | 9.74 ppb | +- | 3.93 ppb
PICO CW41 | Exocor | 65.7 g | 13.566 | 221109 221121 | 2022-11-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
TIG Filler Rod | Executive 316L |  |  |  | 2022-11-23 | (mBq/kg) | <2.65 |  |  | <21.11 |  |  | 0.3264 | +- | 0.2677 | 10.65 | +- | 1.524 | <23.15 |  |  | <5.64 |  |  | <1.26
| 2-36” Welding electrodes cut into 8 cm sections |  |  |  |  | (ppb or ppm) | <0.21 ppb |  |  | <1.71 ppb |  |  | 0.57 ppb | +- | 0.47 ppb | 2.62 ppb | +- | 0.37 ppb | <0.75 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | <356.90 |  |  | 6.5468 | +- | 5.548 | 3.4212 | +- | 1.054 | 5.23 | +- | 2.023 | 0.8592 | +- | 0.68 | 2.5378 | +- | 0.945
[6 empty] | (ppb or ppm) | <28.91 ppb | [8 empty] | 1.29 ppb | +- | 0.5 ppb
PICO CW42 | Lincoln | 71.3 g | 6.841 | 221123 221128 | 2022-11-23 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
TIG Filler Rod | 316L |  |  |  | 2022-11-30 | (mBq/kg) | <1.78 |  |  | <25.37 |  |  | 0.9225 | +- | 0.3649 | 2.173 | +- | 1.562 | 115.33 | +- | 63.56 | <7.68 |  |  | <4.60
| 2-36” Welding electrodes cut into 8 cm sections |  |  |  |  | (ppb or ppm) | <0.14 ppb |  |  | <2.05 ppb |  |  | 1.62 ppb | +- | 0.64 ppb | 0.53 ppb | +- | 0.38 ppb | 3.73 ppm | +- | 2.05 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | <951.70 |  |  | <9.71 |  |  | 4.2869 | +- | 2.302 | 9.204 | +- | 2.853 | 1.331 | +- | 0.8092 | 2.4025 | +- | 1.622
[6 empty] | (ppb or ppm) | <77.09 ppb | [8 empty] | 2.26 ppb | +- | 0.7 ppb
PICO CW43 | Flux Coated | 96.3 g | 6.862 | 221130 221201 | 2022-11-30 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
TIG Filler Rod | 316L |  |  |  | 2022-12-07 | (mBq/kg) | 14870 | +- | 330.6 | 40070 | +- | 1910 | 1051 | +- | 24.73 | 8635 | +- | 244.9 | 79438 | +- | 4206 | <12.70 |  |  | <38.73
| 2-36” Welding electrodes cut into 8 cm sections |  |  |  |  | (ppb or ppm) | 1204.47 ppb | +- | 26.78 ppb | 3245.67 ppb | +- | 154.71 ppb | 1849.76 ppb | +- | 43.52 ppb | 2124.21 ppb | +- | 60.25 ppb | 2565.85 ppm | +- | 135.85 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 73510 | +- | 4452 | <123.50 |  |  | 105.72 | +- | 34.27 | 12470 | +- | 319.2 | <25.81 |  |  | <15.37
[6 empty] | (ppb or ppm) | 5954.31 ppb | +- | 360.61 ppb | [6 empty] | 3067.62 ppb | +- | 78.52 ppb
PICO CW44 | Exocor | 97.1 g | 14.578 | 230404 | 2023-04-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
TIG Filler Rod | Executive 316/316L |  |  |  | 2023-04-19 | (mBq/kg) | 3.538 | +- | 0.9308 | <27.13 |  |  | 0.3841 | +- | 0.2371 | 77.402 | +- | 1.142 | <44.24 |  |  | <6.56 |  |  | <2.36
| 3-36” Welding filler rod cut into 8 cm sections |  |  |  |  | (ppb or ppm) | 0.29 ppb | +- | 0.08 ppb | <2.2 ppb |  |  | 0.68 ppb | +- | 0.42 ppb | 19.04 ppb | +- | 0.28 ppb | <1.43 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 223.3 | +- | 1570 | <4.40 |  |  | 5.11 | +- | 1.538 | 6.773 | +- | 1.883 | <0.75 |  |  | 1.1542 | +- | 1.029
[6 empty] | (ppb or ppm) | 18.09 ppb | +- | 127.17 ppb | [6 empty] | 1.67 ppb | +- | 0.46 ppb
PICO CW45 | PICO Heating Element Circuit Boards | 2.2 g | 13.742 | 230517 | 2023-05-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Polyimide Circuit Boards |  |  |  |  | 2023-05-31 | (mBq/kg) | <16.91 |  |  | <258.90 |  |  | 4.845 | +- | 4.701 | 15.34 | +- | 22.89 | 12716 | +- | 1914 | <193.80 |  |  | <40.27
[6 empty] | (ppb or ppm) | <1.37 ppb |  |  | <20.97 ppb |  |  | 8.53 ppb | +- | 8.27 ppb | 3.77 ppb | +- | 5.63 ppb | 410.73 ppm | +- | 61.82 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <1587.00 |  |  | <124.00 |  |  | <24.31 |  |  | <95.26
[6 empty] | (ppb or ppm) | <128.55 ppb | [8 empty] | <23.43 ppb
PICO CW46 | SAC305 No-Clean T3 Solder Paste | 35.721 g | 7.956 | 230719 | 2023-07-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Chip Quick Solder Paste | Type: SMD291SNL50T3 |  |  |  | 2023-07-27 | (mBq/kg) | <7.47 |  |  | <220.60 |  |  | <2.87 |  |  | 8.063 | +- | 6.137 | 217.09 | +- | 106.9 | <13.26 |  |  | <3.15
| Lot No.: 520005-19986 |  |  |  |  | (ppb or ppm) | <0.61 ppb |  |  | <17.87 ppb |  |  | <5.05 ppb |  |  | 1.98 ppb | +- | 1.51 ppb | 7.01 ppm | +- | 3.45 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 535680 | +- | 32320 | <25.21 |  |  | <4.50 |  |  | 8.487 | +- | 4.466
[6 empty] | (ppb or ppm) | 43.39 ppm | +- | 2.62 ppm | [6 empty] | 2.09 ppb | +- | 1.1 ppb
PICO CW47 | T4 – Halogen & Pb Free | 22.90 g | 27.113 | 231228 240105 240117 | 2023-12-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Loctite Solder Paste | PN: GC 10 SAC305T4 885V 52K |  |  |  | 2024-01-25 | (mBq/kg) | <2.99 |  |  | <140.70 |  |  | 2.569 | +- | 1.605 | <5.70 |  |  | <177.30 |  |  | <18.55 |  |  | <5.80
| Batch No: A733040271 |  |  |  |  | (ppb or ppm) | <0.24 ppb |  |  | <11.4 ppb |  |  | 4.52 ppb | +- | 2.82 ppb | <1.4 ppb |  |  | <5.73 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 251960 | +- | 14500 | <29.52 |  |  | <2.24 |  |  | 6.488 | +- | 3.779
[6 empty] | (ppb or ppm) | 20.41 ppm | +- | 1.17 ppm | [6 empty] | 1.6 ppb | +- | 0.93 ppb
PICO CW48 | Type: EP21TDCS-LO Part A | 5.971 g | 16.479 | 240226 240313 | 2024-02-26 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
MasterBond Epxoy Part A | Lot: 225967 |  |  |  | 2024-03-15 | (mBq/kg) | <8.96 |  |  | <266.70 |  |  | 4.385 | +- | 2.119 | 3.69 | +- | 9.168 | 797.35 | +- | 407.9 | <64.07 |  |  | <24.95
[6 empty] | (ppb or ppm) | <0.73 ppb |  |  | <21.6 ppb |  |  | 7.72 ppb | +- | 3.73 ppb | 0.91 ppb | +- | 2.26 ppb | 25.75 ppm | +- | 13.18 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <9206.00 |  |  | <112.80 |  |  | <10.88 |  |  | 38.57 | +- | 16.4
[6 empty] | (ppb or ppm) | <745.69 ppb | [8 empty] | 9.49 ppb | +- | 4.03 ppb
PICO CW49 | Type: EP21TDCS-LO Part B | 6.900 g | 12.657 | 230315 23031801 | 2024-03-15 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
MasterBond Epxoy Part B | Lot: 225967 |  |  |  | 2024-03-28 | (mBq/kg) | <13.84 |  |  | <138.70 |  |  | <2.97 |  |  | <23.88 |  |  | 639.33 | +- | 371.7 | <62.81 |  |  | <18.71
[6 empty] | (ppb or ppm) | <1.12 ppb |  |  | <11.23 ppb |  |  | <5.23 ppb |  |  | <5.87 ppb |  |  | 20.65 ppm | +- | 12.01 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2161.7 | +- | 7968 | 33.314 | +- | 45.04 | <19.84 |  |  | <36.37
[6 empty] | (ppb or ppm) | 175.1 ppb | +- | 645.41 ppb | [6 empty] | <8.95 ppb
PICO CW50 | Type: TMP1826DGKR | 0.253 g | 13.467 | 240411 240419 240422 | 2024-04-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Texas Instruments Temperature Sensors | One-Wire temperature sensors |  |  |  | 2024-04-25 | (mBq/kg) | <129.40 |  |  | <1935.00 |  |  | 32.21 | +- | 36.2 | <251.00 |  |  | 165.01 | +- | 9182 | <1703.00 |  |  | 379.2 | +- | 388.6
| 10 units |  |  |  |  | (ppb or ppm) | <10.48 ppb |  |  | <156.74 ppb |  |  | 56.69 ppb | +- | 63.71 ppb | <61.75 ppb |  |  | 5.33 ppm | +- | 296.58 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <11930.00 |  |  | <2127.00 |  |  | <485.50 |  |  | 757.2 | +- | 385.7
[6 empty] | (ppb or ppm) | <0.97 ppm | [8 empty] | 186.27 ppb | +- | 94.88 ppb
DAMIC Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
DAMIC CW01 | Epotek 301-2FL two-part epoxy | 4.5 g | 9.042 | 140318 | 2014-03-18 | (mBq/kg) | 1.5 | +- | 50.24 | 11.11 | +- | 31.43 | 13.29 | +- | 9.25 | 24.47 | +- | 43.98 | <63.49 |  |  | <496.84 |  |  | <269.47
Epoxy | Epoxy Technology, Billerica, Mass, USA | Also DAMIC 7 |  |  | 2014-03-27 | (ppb or ppm) | 0.12 ppb | +- | 4.07 ppb | 0.90 ppb | +- | 2.45 ppb | 23.39 ppb | +- | 16.28 ppb | 6.02 ppb | +- | 10.82 ppb | <15.62 ppb
DAMIC CW02 | Ametek Engineered Materials SDN BHD | 11.6 g | 13.575 | 190513 180523 19052301 190527 | 2019-05-13 | (mBq/kg) | 5.237 | +- | 19.83 | 240.9 | +- | 31.4 | 24.7 | +- | 5.169 | 205.9 | +- | 30.81 | 55.94 | +- | 38.67 | <161.40 |  |  | 92.457 | +- | 50.45 | 40K: | <9865.00
Aluminum Rod (with 1% Silicon) |  |  |  |  | 2019-05-27 | (ppb or ppm) | 0.42 ppb | +- | 1.61 ppb | 19.51 ppb | +- | 2.54 ppb | 43.47 ppb | +- | 9.1 ppb | 50.65 ppb | +- | 7.58 ppb | 13.76 ppb | +- | 9.51 ppb |  |  |  | 7.49 ppb | +- | 4.09 ppb |  | <318.64 ppm | #VALUE!
DEAP Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
DEAP CW01 | Prepared On: 140904 | 3.1 g | 6.847 | 140908 | 2014-09-08 | (mBq/kg) | 38.986 | +/- | 92.992 | <55.47 |  |  | 57.941 | +- | 27.941 | 354.312 | +- | 64.129 | 49.06 | +- | 139.04 | <969.36 |  |  | 419.36 | +- | 193.548
Acrylic Vaporization Sample |  |  |  |  | 2014-09-15 |  | 3.16 ppb | +- | 7.69 ppb | <4.49 ppb |  |  | 101.98 ppb | +- | 49.18 ppb | 87.16 ppb | +- | 15.78 ppb | 12.07 ppb | +- | 34.20 ppb
DEAP CW02 | Prepared On: 140826 | 0.3 g | 21.535 | 150708 150729 17072905 | 2015-07-08 | (mBq/kg) | <616.67 |  |  | 466.184 | +- | 288.155 | <93.33 |  |  | 196.907 | +- | 390.315 | 1496.996 | +- | 797.65 | <5106.67 |  |  | 22833.333 | +- | 1986.667
Acrylic Vaporization Sample |  |  |  |  | 2015-07-30 | (ppb or ppm) | <49.85 ppb |  |  | 37.76 ppb | +- | 23.34 ppb | <164.27 ppb |  |  | 48.44 ppb | +- | 96.02 ppb | 368.26 ppb | +- | 196.22 ppb |  |  |  | 1.85 ppm | +- | 0.16 ppm
DEAP CW03 | Solder used for 22Na source encapsulation | 1.4 g | 8.083 | 151028 | 2015-10-28 | (mBq/kg) | <366.43 |  |  | 393.357 | +- | 418.386 | <69.29 |  |  | <174.29 |  |  | 547.94 | +- | 599.519 | <2582.14 |  |  | 152650 | +- | 8272.857
Solder |  |  |  |  | 2015-11-05 | (ppb or ppm) | <29.68 ppb |  |  | 31.86 ppb | +- | 33.89 ppb | <121.94 ppb |  |  | <42.87 ppb |  |  | 134.79 ppb | +- | 147.48 ppb |  |  |  | 12.36 ppm | +- | 0.67 ppm
DEAP CW04 | Foil is 0.0015” thick | 1.5 g | 32.097 | 151115 151125 151209 151212 15121201 151217 | 2015-11-15 | (mBq/kg) | 164.232 | +- | 92.801 | <50.89 |  |  | 7.085 | +- | 13.595 | <117.86 |  |  | <708.86 |  |  | <539.29 |  |  | 513.571 | +- | 168.571
Copper Foil for 22Na Inner Encapsulation |  |  |  |  | 2015-12-18 | (ppb or ppm) | 13.30 ppb | +- | 7.52 ppb | <4.14 ppb |  |  | 12.47 ppb | +- | 23.93 ppb | <28.99 ppb |  |  | <174.13 ppb | [5 empty] | 41.60 ppb | +- | 13.65 ppb
DEAP CW05 | Prepared by Joshua Bonatt on 23 Aug 2016 | --- | 7.823 | 160915 160916 160919 190920 | 2016-09-15 | (mBq) | 0.383 | +- | 0.25 | 0.151 | +- | 0.24 | 0.075 | +- | 0.048 | 0.178 | +- | 0.157 | <0.80 |  |  | <3.74 |  |  | 627.967 | +- | 32.593
Acrylic Vaporization Sample | Labelled “Efficiency 4” |  |  |  | 2016-09-23
DEAP CW06 | Prepared by Joshua Bonatt on 02 Nov 2016 | --- | 10.792 | 161103 161104 161111 16111101 | 407919 | (mBq) | 0.278 | +- | 0.244 | <0.098 |  |  | 0.117 | +- | 0.044 | 2.794 | +- | 0.304 | 0.21 | +- | 0.43 | <2.33 |  |  | 426.29 | +- | 22.155
Acrylic Vaporization Sample | Labelled “Efficiency 5” |  |  |  | 2016-11-14
DEAP CW07 | Prepared by Joshua Bonatt on 03 Nov 2016 | --- | 17.716 | 161114 | 2016-11-14 | (mBq) | 0.246 | +- | 0.169 | <0.124 |  |  | 0.051 | +- | 0.029 | 0.059 | +- | 0.117 | 0.115 | +- | 0.264 | 0.287 | +- | 1.211 | 6.257 | +- | 0.599
Acrylic Vaporization Sample | Labelled ”Blank” |  |  |  | 2016-12-02
DEAP CW08 | Silver Zeolite: Ionex Type AG400 | 6.8 g | 13.808 | 201204 | 2020-12-04 | (mBq/kg) | 3783 | +- | 215.8 | 2066 | +- | 187.5 | 216.4 | +- | 11.78 | 3515 | +- | 213.2 | 1686 | +- | 182.5 | <508.60 |  |  | 742.57 | +- | 163.2
Silver Zeolite | Lot #: 1911002-0062 Molecular Products Inc. |  |  | The sample is also counted on the PGT detector as DEAP 138 | 2020-12-18 | (ppb or ppm) | 306.42 ppb | +- | 17.48 ppb | 167.35 ppb | +- | 15.19 ppb | 380.86 ppb | +- | 20.73 ppb | 864.69 ppb | +- | 52.45 ppb | 414.76 ppb | +- | 44.9 ppb |  |  |  | 60.15 ppb | +- | 13.22 ppb
EXO and nEXO Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
3 samples have been counted for nEXO, results are only available to EXO and nEXO collaboration members.
NEWS-G Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
NEWS-G CW01 | 400 micron UV/Vis Item No: 6-103288 | 0.8 g | 19.756 | 170207 170214 170216 | 2017-02-07 | (mBq/kg) | 79.415 | +- | 209.193 | <140.78 |  |  | 39.948 | +- | 35.079 | <165.00 |  |  | 62.728 | +- | 297.703 | 296.25 | +- | 1361.25 | 1365 | +- | 413.75
In-Vacuum multimode fibre optic cable, Polymide Buffer | Accu-Glass Products Inc., Valenca | 79 pieces |  |  | 2017-02-27 | (ppb or ppm) | 6.43 ppb | +- | 16.94 ppb | <11.40 ppb |  |  | 70.31 ppb | +- | 61.74 ppb | <40.59 ppb |  |  | 15.43 ppb | +- | 73.24 ppb |  |  |  | 110.57 ppb | +- | 33.51 ppb
NEWS-G CW02 | 400 micron UV/Vis Item No: 6-103289 | 1.4 g | 15.75 | 170227 | 2017-02-27 | (mBq/kg) | <134.29 |  |  | <94.42 |  |  | 31.9 | +- | 20.32 | <103.57 |  |  | 225.581 | +- | 195.716 | <1510.71 |  |  | 750.714 | +- | 257.143
In-Vacuum multimode fibre optic cable, Aluminum Buffer | Accu-Glass Products Inc., Valenca | 79 pieces |  |  | 2017-03-15 | (ppb or ppm) | <10.88 ppb |  |  | <7.65 ppb |  |  | 56.15 ppb | +- | 35.76 ppb | <25.28 ppb |  |  | 55.49 ppb | +- | 48.15 ppb |  |  |  | 60.81 ppb | +- | 20.83 ppb
NEWS-G CW03 | Arubis C10100 – Cu-OFE (OF-Cu) | 23.634 g | 14.281 | 171005 171009 | 2017-10-05 | (mBq) | 0.196 | +- | 0.291 | 0.582 | +- | 0.402 | 0.072 | +- | 0.068 | 0.421 | +- | 0.282 | <0.67 |  |  | <1.06 |  |  | 11.7 | +- | 5.71 | 57Co:
Copper Rod | Cleaned with 1% citric acid |  |  |  | 2017-10-20 | (mBq/kg) | 8.277 | +- | 12.319 | 24.644 | +- | 16.997 | 3.048 | +- | 2.859 | 17.8 | +- | 11.914 | <28.39 |  |  | <44.82 |  |  | 494.89 | +- | 241.47 | 0.041 mBq | +- | 0.027 mBq
|  | Efficiencies are corrected for self-absorption by the copper using ANGLE4 |  |  |  | (ppb or ppm) | 0.67 ppb | +- | 1.00 ppb | 2.00 ppb | +- | 1.38 ppb | 5.37 ppb | +- | 5.03 ppb | 4.39 ppb | +- | 2.93 ppb | <6.98 ppb | [5 empty] | 40.09 ppb | +- | 19.56 ppb | 1.74 mBq/kg | +- | 1.16 mBq/kg
NEWS-G CW04 | Arubis C10100 – Cu-OFE (OF-Cu) | 23.724 g | 15.475 | 171101 171106 17110601 17110602 171114 171115 | 2017-11-01 | (mBq) | <0.038 |  |  | 0.14 | +- | 0.359 | 0.063 | +- | 0.07 | 0.069 | +- | 0.275 | 0.219 | +- | 0.626 | <0.63 |  |  | 8.52 | +- | 5.95 | 57Co:
Copper Rod | Etched with 5% nitric acid |  |  |  | 2017-11-17 | (mBq/kg) | <1.62 |  |  | 5.904 | +- | 15.123 | 2.649 | +- | 2.966 | 2.898 | +- | 11.611 | 9.25 | +- | 26.384 | <26.42 |  |  | 359.198 | +- | 250.669 | 0.033 mBq | +- | 0.027 mBq
|  | Efficiencies are corrected for self-absorption by the copper using ANGLE4 |  |  |  | (ppb or ppm) | <0.13 ppb |  |  | 0.48 ppb | +- | 1.23 ppb | 4.66 ppb | +- | 5.22 ppb | 0.71 ppb | +- | 2.86 ppb | 2.28 ppb | +- | 6.49 ppb |  |  |  | 29.10 ppb | +- | 20.30 ppb | 1.37 mBq/kg | +- | 1.12 mBq/kg
NEWS-G 005 | This sample is also being counted for CUTE | 119.1 g | 5.4 | 171201 | 2017-12-01 | (mBq) | <3.72 |  |  | 71.495 | +- | 41.852 | <4.58 |  |  | 2.783 | +- | 5.64 | <5.32 |  |  | <5.85 |  |  | 38865.3 | +- | 2067.18
Lead Rod from Marshield | Manufactured by Marshield Ltd |  |  |  | 2017-12-06 | (mBq/kg) | <31.19 |  |  | 600.294 | +- | 351.404 | <38.46 |  |  | 23.368 | +- | 47.359 | <44.65 |  |  | <49.08 |  |  | 326324.927 | +- | 17356.7
| Cylinder is 34 mm x 10 mm | Efficiencies are corrected for self-absorption by the lead using ANGLE4 |  |  |  | (ppb or ppm) | <2.53 ppb |  |  | 48.62 ppb | +- | 28.46 ppb | <67.69 ppb |  |  | 5.75 ppb | +- | 11.65 ppb | <10.98 ppb | [5 empty] | 26.43 ppm | +- | 1.41 ppm
NEWS-G CW06 | Extraordinary Adsorbants | 8.9 g | 8.928 | 210426 210429 | 2021-04-26 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Silver Zeolite | Ag-ETS-10 |  |  |  | 2021-05-05 | (mBq/kg) | 987.9 | +- | 47.62 | 2022 | +- | 390.1 | 49.81 | +- | 8.217 | 601.4 | +- | 43.04 | 579760 | +- | 30440 | <127.20 |  |  | <147.20
| (15 g 16-30 mesh) |  |  |  |  | (ppb or ppm) | 80.02 ppb | +- | 3.86 ppb | 163.78 ppb | +- | 31.6 ppb | 87.67 ppb | +- | 14.46 ppb | 147.94 ppb | +- | 10.59 ppb | 1.87 % | +- | 0.1 %
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 3437.9 | +- | 2156 | <174.70 |  |  | <62.09 |  |  | 398.9 | +- | 76.7
[6 empty] | (ppb or ppm) | 278.47 ppb | +- | 174.64 ppb | [6 empty] | 98.13 ppb | +- | 18.87 ppb
NEWS-G CW07 | Extraordinary Adsorbents | 7.2 g | 20.083 | 210709 | 2021-07-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Sodium Zeolite Absorbent | Na-ETS-10 |  |  |  | 2021-07-29 | (mBq/kg) | 204.9 | +- | 32.13 | 1473 | +- | 390.1 | 15.45 | +- | 7.424 | 48 | +- | 31.02 | 5113706.08 | + | 14163.44 | 27.789 | +- | 57.3 | 178.2 | +- | 137.9
| (16-30 Mesh) |  |  |  |  | (ppb or ppm) | 16.6 ppb | +- | 2.6 ppb | 119.31 ppb | +- | 31.6 ppb | 27.19 ppb | +- | 13.07 ppb | 11.81 ppb | +- | 7.63 ppb | 16.52 % | +- | 0.05 %
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 5380 | +- | 2560 | <162.60 |  |  | <89.79 |  |  | <74.57
[6 empty] | (ppb or ppm) | 435.78 ppb | +- | 207.36 ppb | [6 empty] | <18.34 ppb
CUTE Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
CUTE CW01 | FA 412 Lead (low activity lead) 1” x 5/8” d | 56.5 g | 13.124 | 19011602 19011702 19012102 | 2019-01-16 | (mBq/kg) | 52.53 | +- | 24.6 | <132.60 |  |  | 17.05 | +- | 16.33 | 17.91 | +- | 27.34 | <59.10 |  |  | <45.20 |  |  | 38748 | +- | 2903
Lead Cylinder | Fonderie de Gentilly | Efficiencies are corrected for self-absorption by the lead using ANGLE4 |  |  | 2019-01-30 | (ppb or ppm) | 4.25 ppb | +- | 1.99 ppb | <10.74 ppb |  |  | 30.01 ppb | +- | 28.74 ppb | 4.41 ppb | +- | 6.73 ppb | <14.54 ppb | [5 empty] | 3.14 ppm | +- | 0.24 ppm
CUTE CW02 | TFA 393 Lead (very low activity lead) 1” x 5/8” d | 56.6 g | 15.538 | 190213 190218 | 2019-02-13 | (mBq/kg) | 12.24 | +- | 20.3 | <109.90 |  |  | 8.799 | +- | 10.88 | 33.81 | +- | 20.17 | <59.77 |  |  | <38.20 |  |  | 6294.8 | +- | 1107
Lead Cylinder | Fonderie de Gentilly | Efficiencies are corrected for self-absorption by the lead using ANGLE4 |  |  | 2019-03-01 | (ppb or ppm) | 0.99 ppb | +- | 1.64 ppb | <8.9 ppb |  |  | 15.49 ppb | +- | 19.15 ppb | 8.32 ppb | +- | 4.96 ppb | <14.7 ppb | [5 empty] | 509.88 ppb | +- | 89.67 ppb
SENSEI Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
SENSEI CW01 | TFA Lead 1” x 5/8” d | 56.5 g | 10.754 | 190926 | 2019-09-26 | (mBq/kg) | 47.14 | +- | 21.98 | 74.58 | +- | 116.6 | <6.94 |  |  | 19.95 | +- | 25.22 | 42.26 | +- | 75.16 | <47.38 |  |  | 4238.9 | +- | 999.8
Lead Cylinder | Fonderie de Gentilly | Efficiencies are corrected for self-absorption by the lead using ANGLE4 |  |  | 2019-10-07 | (ppb or ppm) | 3.82 ppb | +- | 1.78 ppb | 6.04 ppb | +- | 9.44 ppb | <12.21 ppb |  |  | 4.91 ppb | +- | 6.2 ppb | 10.4 ppb | +- | 18.49 ppb |  |  |  | 343.35 ppb | +- | 80.98 ppb
SBC Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | 40K:
SBC CW01 | SiPMs labelled 96 & 97 | 2.6 g | 21.516 | 200302 | 2020-03-02 | (mBq/kg) | 45060 | +- | 2019 | 28170 | +- | 2292 | 2328 | +- | 72.4 | 10940 | +- | 708.1 | 6009 | +- | 568.3 | 222.88 | +- | 593.6 | 20544 | +- | 611.8 | 68019 | +- | 49610
Hamamatsu VUV4 Quad SiPM (MEG-II Style) |  | (2 SiPMs) |  |  | 2020-03-24 | (ppb or ppm) | 3649.86 ppb | +- | 163.54 ppb | 2281.77 ppb | +- | 185.65 ppb | 4097.28 ppb | +- | 127.42 ppb | 2691.24 ppb | +- | 174.19 ppb | 1478.21 ppb | +- | 139.8 ppb |  |  |  | 1664.06 ppb | +- | 49.56 ppb | 2197.01 ppm | +- | 1602.4 ppm
[6 empty] | (mBq / SiPM) | 58.578 | +- | 2.6247 | 36.621 | +- | 2.9796 | 3.0264 | +- | 0.09412 | 14.222 | +- | 0.92053 | 7.8117 | +- | 0.73879 | 0.28974 | +- | 0.77168 | 26.7072 | +- | 0.79534 | 88.4247 | +- | 64.493
SBC CW02 | PCB Size: 15 mm x 15 mm x 1.5 mm | 21.2 g | 15.552 | 200722 | 2020-07-22 | (mBq/kg) | 8740 | +- | 337.3 | 7803 | +- | 615.6 | 498.8 | +- | 14.69 | 15710 | +- | 688.2 | 8831 | +- | 349 | <299.80 |  |  | 4824.2 | +- | 324.5
PCBs for the SBC SiPMs | Quantity: 30 | 0.7 g/ PCB |  |  | 2020-08-11 | (ppb or ppm) | 707.94 ppb | +- | 27.32 ppb | 632.04 ppb | +- | 49.86 ppb | 877.89 ppb | +- | 25.85 ppb | 3864.66 ppb | +- | 169.3 ppb | 2172.43 ppb | +- | 85.85 ppb |  |  |  | 390.76 ppb | +- | 26.28 ppb
|  | Additional Activities: |  |  |  | (mBq/kg) |  |  | 40K: | 11541 | +- | 11560 | 227Ac: |  |  | 1233 | +- | 92.76 |  | 54Mn: |  | 212.23 | +- | 493.2 |  |  | 60Co: | 1005 | +- | 678.2
[6 empty] | (ppb or ppm) |  |  |  | 372.77 ppm | +- | 373.39 ppm |  |  |  | 2170.08 ppb | +- | 163.26 ppb
SBC CW03 | Kester Lead-Free Solder | 10.6 g | 8.887 | 210623 210627 210630 | 2021-06-23 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tin-Silver Solder | Solder: Sn 96.5% Sg: 3.0% Cu: 05% |  |  |  | 2021-07-02 | (mBq/kg) | <7.54 |  |  | 528.2 | +- | 364.4 | 5.067 | +- | 5.36 | <13.30 |  |  | <329.40 |  |  | <36.93 |  |  | <7.49
| Lot: W201038 Solder diameter: 0.8 mm |  |  |  |  | (ppb or ppm) | <0.61 ppb |  |  | 42.78 ppb | +- | 29.52 ppb | 8.92 ppb | +- | 9.43 ppb | <3.27 ppb |  |  | <10.64 ppm
Also counted as PGT SBC P09 | This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
| Kester Itaska, IL, USA |  |  |  |  | (mBq/kg) | 10329 | +- | 7231 | 24.899 | +- | 39.77 | <7.34 |  |  | <11.50
[6 empty] | (ppb or ppm) | 836.65 ppb | +- | 585.71 ppb | [6 empty] | <2.83 ppb
SBC CW04 | Stainless Steel | 53.5 g | 11.486 | 210505 | 2021-05-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Vessel Body |  |  |  | 2021-05-17 | (mBq/kg) | 1.448 | +- | 1.427 | <25.48 |  |  | <0.31 |  |  | <3.24 |  |  | 15.501 | +- | 47.23 | <11.46 |  |  | <6.62
| PV Drawing Item #3 |  |  |  |  | (ppb or ppm) | 0.12 ppb | +- | 0.12 ppb | <2.06 ppb |  |  | <0.55 ppb |  |  | <0.8 ppb |  |  | 0.5 ppm | +- | 1.53 ppm
Also counted as VDA SBC-V01 | [5 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 838.27 | +- | 907.9 | <11.73 |  |  | 9.3909 | +- | 2.711 | <3.67 |  |  | 0.7408 | +- | 0.7333 | <2.53
[6 empty] | (ppb or ppm) | 67.9 ppb | +- | 73.54 ppb | [6 empty] | <0.9 ppb
SBC CW05 | Units: 10 | 2.85 g | 6.589 | 211029 211101 211102 | 2021-10-29 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Four-Pin RTD Connectors |  |  |  |  | 2021-11-05 | (mBq/kg) | 4972 | +- | 196.5 | 9792 | +- | 758.6 | 282.6 | +- | 19.92 | 9605 | +- | 372 | 12476 | +- | 2337 | <180.70 |  |  | <135.20
[6 empty] | (ppb or ppm) | 402.73 ppb | +- | 15.92 ppb | 793.15 ppb | +- | 61.45 ppb | 497.38 ppb | +- | 35.06 ppb | 2.36 ppm | +- | 0.09 ppm | 402.97 ppm | +- | 75.49 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 4722.5 | +- | 1342 | <603.20 |  |  | <174.10 |  |  | 8948 | +- | 372.8
[6 empty] | (ppb or ppm) | 382.52 ppb | +- | 108.7 ppb | [6 empty] | 2.2 ppm | +- | 0.09 ppm
SBC CW06 | Epoxy colour is blue | 6.25 g | 27.383 | 211105 | 2021-11-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stycast Epoxy |  |  |  |  | 2021-12-03 | (mBq/kg) | 160.8 | +- | 9.868 | 2658 | +- | 163 | 48.19 | +- | 2.454 | 119.8 | +- | 9.485 | 473.77 | +- | 225 | <59.64 |  |  | <14.21
[6 empty] | (ppb or ppm) | 13.02 ppb | +- | 0.8 ppb | 215.3 ppb | +- | 13.2 ppb | 84.81 ppb | +- | 4.32 ppb | 29.47 ppb | +- | 2.33 ppb | 15.3 ppm | +- | 7.27 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <545.10 |  |  | <65.98 |  |  | <16.26 |  |  | 75.3 | +- | 13.88
[6 empty] | (ppb or ppm) | <44.15 ppb | [8 empty] | 18.52 ppb | +- | 3.41 ppb
SBC CW07 | Kapton Insulated single conductor wire | 8.45 g | 13.587 | 211203 211207 211213 | 2021-12-03 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
RTD Wire | Kapton, Silver plated copper |  |  |  | 2021-12-17 | (mBq/kg) | <5.92 |  |  | 19.85 | +- | 45.18 | 1.327 | +- | 1.095 | <7.68 |  |  | 1045.9 | +- | 734.3 | <527.90 |  |  | <12.42
[6 empty] | (ppb or ppm) | <0.48 ppb |  |  | 1.61 ppb | +- | 3.66 ppb | 2.34 ppb | +- | 1.93 ppb | <1.89 ppb |  |  | 33.78 ppm | +- | 23.72 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <660.40 |  |  | 553.13 | +- | 570.4 | <112.80 |  |  | 13.44 | +- | 10.47
[6 empty] | (ppb or ppm) | <53.49 ppb | [8 empty] | 3.31 ppb | +- | 2.58 ppb
SBC CW08 | Circuit Board Material: Arlon 85N | 1.85 g | 27.571 | 211217 220103 | 2021-12-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
PCB Materials | OnBoardCircuits Inc. |  |  |  | 2022-01-14 | (mBq/kg) | 8782 | +- | 268.4 | 15040 | +- | 924.1 | 506.9 | +- | 18.63 | 12300 | +- | 399.7 | 14881 | +- | 1851 | <164.10 |  |  | <118.60
[6 empty] | (ppb or ppm) | 711.34 ppb | +- | 21.74 ppb | 1218.24 ppb | +- | 74.85 ppb | 892.14 ppb | +- | 32.79 ppb | 3025.8 ppb | +- | 98.33 ppb | 480.66 ppm | +- | 59.79 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 7378 | +- | 1067 | <454.80 |  |  | 177.64 | +- | 78.8 | 14050 | +- | 425.3
[6 empty] | (ppb or ppm) | 597.62 ppb | +- | 86.43 ppb | [6 empty] | 3456.3 ppb | +- | 104.62 ppb
SBC CW09 | FBK VUV HD4 1x1cm | 13.5 g | 14.68 | 240125 | 2024-01-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
SiPMs | Foundazione Bruno Kessler, Italy |  |  |  | 2024-02-09 | (mBq/kg) | <2.03 |  |  | <15.80 |  |  | <1.24 |  |  | <4.35 |  |  | 344.6 | +- | 191.5 | <31.10 |  |  | <7.92
| Quantity: 30 |  |  |  |  | (ppb or ppm) | <0.16 ppb |  |  | <1.28 ppb |  |  | <2.18 ppb |  |  | <1.07 ppb |  |  | 11.13 ppm | +- | 6.19 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <263.40 |  |  | <21.85 |  |  | <8.89 |  |  | <6.48
[6 empty] | (ppb or ppm) | <21.34 ppb | [8 empty] | <1.59 ppb
QBITS-CUTE Measurements:
QBITS-CUTE 01 | Homemade IR filters (CNTs) | 68.8 g | 13.408 | 220712 220720 220721 22072101 | 2022-07-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
IR Filter |  | 3 Pieces |  |  | 2022-07-26 | (mBq/kg) | 3.531 | +- | 1.612 | 193.3 | +- | 37.37 | 5.379 | +- | 0.6095 | 9.631 | +- | 2.043 | 57.837 | +- | 34.61 | <0.89 |  |  | <3.62
[6 empty] | (ppb or ppm) | 0.29 ppb | +- | 0.13 ppb | 15.66 ppb | +- | 3.03 ppb | 9.47 ppb | +- | 1.07 ppb | 2.37 ppb | +- | 0.5 ppb | 1.87 ppm | +- | 1.12 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 753.24 | +- | 876.4 | 0.86986 | +- | 0.5936 | <0.15 |  |  | 11.46 | +- | 3.133 | <0.063 |  |  | <0.34
[6 empty] | (ppb or ppm) | 61.01 ppb | +- | 70.99 ppb | [6 empty] | 2.82 ppb | +- | 0.77 ppb
QBITS-CUTE 02 |  | 6.2 g | 8.708 | 220726 220728 220729 22072901 | 2022-07-26 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
NbTI Coaxial Cable with Connectors |  |  |  | 220731 220803 | 2022-08-05 | (mBq/kg) | 56.6 | +- | 13.82 | 17470 | +- | 1043 | 352.2 | +- | 13.31 | 509.4 | +- | 33.78 | 137.77 | +- | 146.4 | <58.40 |  |  | <9.25
[6 empty] | (ppb or ppm) | 4.58 ppb | +- | 1.12 ppb | 1415.07 ppb | +- | 84.48 ppb | 619.87 ppb | +- | 23.43 ppb | 125.31 ppb | +- | 8.31 ppb | 4.45 ppm | +- | 4.73 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 11055 | +- | 4231 | <150.20 |  |  | <16.52 |  |  | 527 | +- | 49.22
[6 empty] | (ppb or ppm) | 895.46 ppb | +- | 342.71 ppb | [6 empty] | 129.64 ppb | +- | 12.11 ppb
QBITS-CUTE CW03 |  | 10.0 g | 6.711 | 230210 | 2023-02-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Used Sample Holder Circuit Boards |  | 5 Units |  |  | 2023-02-17 | (mBq/kg) | 270.4 | +- | 18.18 | 419.3 | +- | 94.71 | 15.22 | +- | 2.701 | 220.5 | +- | 18.03 | 3593.4 | +- | 686.7 | <39.71 |  |  | 9.724 | +- | 15.07
[6 empty] | (ppb or ppm) | 21.9 ppb | +- | 1.47 ppb | 33.96 ppb | +- | 7.67 ppb | 26.79 ppb | +- | 4.75 ppb | 54.24 ppb | +- | 4.44 ppb | 116.07 ppm | +- | 22.18 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1610.1 | +- | 717.7 | <42.98 |  |  | <15.71 |  |  | 200.9 | +- | 26.82
[6 empty] | (ppb or ppm) | 130.42 ppb | +- | 58.13 ppb | [6 empty] | 49.42 ppb | +- | 6.6 ppb
QBITS-CUTE CW04 |  | 10.7 g | 6.866 | 230217 | 2023-02-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
New Sample Holder Circuit Boards |  | 5 Units |  |  | 2023-02-24 | (mBq/kg) | 241.6 | +- | 16.41 | 459.1 | +- | 83.62 | 11.7 | +- | 2.477 | 208.7 | +- | 16.62 | 1802.2 | +- | 517.1 | <40.46 |  |  | <25.22
[6 empty] | (ppb or ppm) | 19.57 ppb | +- | 1.33 ppb | 37.19 ppb | +- | 6.77 ppb | 20.59 ppb | +- | 4.36 ppb | 51.34 ppb | +- | 4.09 ppb | 58.21 ppm | +- | 16.7 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <984.90 |  |  | 47.63 | +- | 40.28 | <13.69 |  |  | 174.6 | +- | 24.3
[6 empty] | (ppb or ppm) | <79.78 ppb | [8 empty] | 42.95 ppb | +- | 5.98 ppb
QBITS-CUTE CW05 | Kurt-J Lesker (99.999% Pure) | 4.246 g | 6.759 | 230224 230228 | 2023-02-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Aluminum Pellets (New) |  |  |  |  | 2023-03-03 | (mBq/kg) | <13.58 |  |  | <133.00 |  |  | <6.66 |  |  | 16.2 | +- | 13.42 | <754.90 |  |  | <90.26 |  |  | <12.60
[6 empty] | (ppb or ppm) | <1.1 ppb |  |  | <10.77 ppb |  |  | <11.72 ppb |  |  | 3.99 ppb | +- | 3.3 ppb | <24.38 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co |  |  | 24Na
[6 empty] | (mBq/kg) | <967.10 |  |  | <128.60 |  |  | <15.93 |  |  | 41.26 | +- | 18.96 | 9.158 | +- | 4.551 | <16.20 |  |  | <38.64
[6 empty] | (ppb or ppm) | <78.34 ppb | [8 empty] | 10.15 ppb | +- | 4.66 ppb
QBITS-CUTE CW06 | Alfa Aeser Puratonics (99.999% Pure) | 1.97 g | 13.031 | 230303 230314 | 2023-03-03 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Aluminum Slugs (New) |  |  |  |  | 2023-03-17 | (mBq/kg) | <17.31 |  |  | <131.90 |  |  | <7.83 |  |  | <18.91 |  |  | 1780.1 | +- | 962 | <209.00 |  |  | <38.08
[6 empty] | (ppb or ppm) | <1.4 ppb |  |  | <10.68 ppb |  |  | <13.78 ppb |  |  | <4.65 ppb |  |  | 57.5 ppm | +- | 31.07 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co |  |  | 24Na
[6 empty] | (mBq/kg) | <1762.00 |  |  | <160.60 |  |  | <37.99 |  |  | 53.61 | +- | 43.65 | <15.02 |  |  | <27.79 |  |  | <47.76
[6 empty] | (ppb or ppm) | <142.72 ppb | [8 empty] | 13.19 ppb | +- | 10.74 ppb
QBITS-CUTE CW07 | Kurt-J Lesker (99.995% Pure) | 2.66 g | 10.704 | 230317 | 2023-03-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Titanium Pellets (New) |  |  |  |  | 2023-03-28 | (mBq/kg) | 21.97 | +- | 20.06 | <154.40 |  |  | 4.998 | +- | 4.338 | <23.08 |  |  | <1538.00 |  |  | <129.80 |  |  | <72.47
[6 empty] | (ppb or ppm) | 1.78 ppb | +- | 1.62 ppb | <12.51 ppb |  |  | 8.8 ppb | +- | 7.63 ppb | <5.68 ppb |  |  | <49.68 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <2802.00 |  |  | <137.70 |  |  | <45.93 |  |  | <47.48
[6 empty] | (ppb or ppm) | <226.96 ppb | [8 empty] | <11.68 ppb
QBITS-CUTE CW08 | Chalmers University of Technology | 4.806 g | 6.866 | 230328 | 2023-03-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tin Plated Printed Circuit Boards |  |  |  |  | 2023-04-04 | (mBq/kg) | 5386 | +- | 191.3 | 9544 | +- | 688.6 | 304.2 | +- | 16.54 | 6481 | +- | 250.1 | 7503 | +- | 1695 | <70.62 |  |  | <69.07
| 13 units |  |  |  |  | (ppb or ppm) | 436.27 ppb | +- | 15.5 ppb | 773.06 ppb | +- | 55.78 ppb | 535.39 ppb | +- | 29.11 ppb | 1594.33 ppb | +- | 61.52 ppb | 242.35 ppm | +- | 54.75 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 3163.1 | +- | 1665 | <376.80 |  |  | 199.85 | +- | 78.96 | 6254 | +- | 253.1
[6 empty] | (ppb or ppm) | 256.21 ppb | +- | 134.87 ppb | [6 empty] | 1538.48 ppb | +- | 62.26 ppb
QBITS-CUTE CW09 | Chalmers University of Technology | 12.9568 g | 6.812 | 231129 | 2023-11-29 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
IR Filter from Bluefors |  |  |  |  | 2023-12-06 | (mBq/kg) | <8.66 |  |  | 1459 | +- | 216.5 | 28.91 | +- | 3.364 | 20.58 | +- | 9.113 | 505.53 | +- | 297.6 | <40.70 |  |  | <20.96
| Sample #2 |  |  |  |  | (ppb or ppm) | <0.7 ppb |  |  | 118.18 ppb | +- | 17.54 ppb | 50.88 ppb | +- | 5.92 ppb | 5.06 ppb | +- | 2.24 ppb | 16.33 ppm | +- | 9.61 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 77515 | +- | 5479 | <55.55 |  |  | 8.8599 | +- | 9.627 | 23.51 | +- | 13.18
[6 empty] | (ppb or ppm) | 6.28 ppm | +- | 0.44 ppm | [6 empty] | 5.78 ppb | +- | 3.24 ppb
QBITS-CUTE CW10 | Chalmers University of Technology | 5.092 g | 5.843 | 231206 231209 | 2023-12-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Screws M2 | Brass screws small 16 units Skruvcenter |  |  |  | 2023-12-12 | (mBq/kg) | <14.96 |  |  | <62.19 |  |  | <6.64 |  |  | <11.83 |  |  | 391.51 | +- | 525.7 | <79.64 |  |  | 27.71 | +- | 28.12
| Sample #5 16 units |  |  |  |  | (ppb or ppm) | <1.21 ppb |  |  | <5.04 ppb |  |  | <11.69 ppb |  |  | <2.91 ppb |  |  | 12.65 ppm | +- | 16.98 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <603.70 |  |  | <146.40 |  |  | <15.44 |  |  | <30.20
[6 empty] | (ppb or ppm) | <48.9 ppb | [8 empty] | <7.43 ppb
QBITS-CUTE CW11 | Chalmers University of Technology | 0.584 g | 6.86 | 231212 | 2023-12-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
PCB Circuits | AlOx Tin plated PCB from CERcuits |  |  |  | 2023-12-19 | (mBq/kg) | 3337 | +- | 276.7 | 8104 | +- | 1197 | 189.5 | +- | 41.39 | 6149 | +- | 407.2 | 14077 | +- | 7897 | <679.10 |  |  | <469.20
| Sample #6 |  |  |  |  | (ppb or ppm) | 270.3 ppb | +- | 22.41 ppb | 656.42 ppb | +- | 96.96 ppb | 333.52 ppb | +- | 72.85 ppb | 1512.65 ppb | +- | 100.17 ppb | 454.69 ppm | +- | 255.07 ppm
| 2 Units |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <4672.00 |  |  | 984.92 | +- | 946.92 | <228.90 |  |  | 4253 | +- | 573.2
[6 empty] | (ppb or ppm) | <378.43 ppb | [8 empty] | 1046.24 ppb | +- | 141.01 ppb
QBITS-CUTE CW12 | Chalmers University of Technology | 1.506 g | 9.007 | 231219 | 2023-12-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Circuit Boards | PCB from CIRcuits |  |  |  | 2023-12-28 | (mBq/kg) | 3144 | +- | 161.8 | 6719 | +- | 604.5 | 215.9 | +- | 21.15 | 5079 | +- | 256.1 | 2901.8 | +- | 2820 | <323.40 |  |  | <153.20
| Sample #8-1 |  |  |  |  | (ppb or ppm) | 254.66 ppb | +- | 13.11 ppb | 544.24 ppb | +- | 48.96 ppb | 379.98 ppb | +- | 37.22 ppb | 1249.43 ppb | +- | 63 ppb | 93.73 ppm | +- | 91.09 ppm
| 5 Units |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1439.3 | +- | 1707 | <335.30 |  |  | 132.08 | +- | 109 | 4354 | +- | 307.1
[6 empty] | (ppb or ppm) | 116.58 ppb | +- | 138.27 ppb | [6 empty] | 1071.08 ppb | +- | 75.55 ppb
HALO Measurements:
HALO CW01 | HALO Cf-252 Source | 34.012 g |  | 221104 | 2022-11-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
SRS-10-003 Cf-252 Source |  |  |  |  | 2022-11-04 | (mBq/kg)
| Special measurement to determine strength of the HALO source |  |  |  |  | (ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)
HALO CW02 | HALO Cf-252 “Dummy” Source | 51.03 g | 11.412 | 221107 22110701 | 2022-12-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
SRS-10-003 Cf-252 “Dummy” Source |  |  |  |  | 2022-12-19 | (mBq/kg) | 2.081 | +- | 1.699 | <42.61 |  |  | <0.61 |  |  | 9.586 | +- | 2.219 | <40.49 |  |  | <8.46 |  |  | 2.684 | +- | 1.471
[6 empty] | (ppb or ppm) | 0.17 ppb | +- | 0.14 ppb | <3.45 ppb |  |  | <1.07 ppb |  |  | 2.36 ppb | +- | 0.55 ppb | <1.31 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | <537.50 |  |  | <9.77 |  |  | <1.82 |  |  | 10.03 | +- | 3.517 | <1.07 |  |  | <2.86
[6 empty] | (ppb or ppm) | <43.54 ppb | [8 empty] | 2.47 ppb | +- | 0.87 ppb
CNL Measurements:
CNL CW01 | [5 empty] | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Place Holder | [5 empty] | (mBq/kg)
(ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)
Super CDMS Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
11 samples have been counted for SuperCDMS, results are only available to SuperCDMS collaboration members.
DARKSIDE Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable) |  |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 228Ac |  |  | 137Cs |  |  | 210Pb |  | Comments
6 samples have been counted for DarkSide, results are only available to DarkSide collaboration members.
In Progress and To Be Measured:
Sample Description |  | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Background CW20 |  | N/A | 35.57 | 230916 23091801 230927 230929 23092901 | 2023-09-16 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  | 230930 23093001 231001 231003 23100401 | 2023-10-26 | (mBq) | 0.1658 | +- | 0.02688 | 0.1943 | +- | 0.1314 | 0.008162 | +- | 0.006008 | 0.1948 | +- | 0.02804 | 6.0279 | +- | 1.347 | 0.069115 | +- | 0.03101 | 0.09021 | +- | 0.05005
231008 231012 231014 231018 23101801
|  |  |  | 231023 231024 23102401 231025 |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (Mbq) | <1.19 |  |  | 0.10168 | +- | 0.1381 | 0.022679 | +- | 0.03624 | 0.2475 | +- | 0.07385
PICO CW51 | Part Number: C0805X105K8RAC7210 | 0.140 g |  |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
KEMET Capacitors | Type: 1uF 10V X7R 0805 |  |  |  |  | (mBq/kg)
| 13 units |  |  |  |  | (ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)
LEGEND CW01 | Taiwan Semiconductor Manufacturing Company Limited | 0.57 g | 7.093 | 250425 | 2024-04-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
ASIC Chips | J/N A16-07-374-4 |  |  |  |  | (mBq/kg) | <132.50 |  |  | <359.80 |  |  | 29.65 | +- | 26.03 | <163.00 |  |  | 7703.2 | +- | 6561 | <643.10 |  |  | <263.30
Cold ATA V1 | Die Size: 7.85 mm x 7.85 mm x 0.3 mm |  |  |  |  | (ppb or ppm) | <10.73 ppb |  |  | <29.14 ppb |  |  | 52.18 ppb | +- | 45.81 ppb | <40.1 ppb |  |  | 248.81 ppm | +- | 211.92 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
| Number of Units: 13 |  |  |  |  | (mBq/kg) | <5312.00 |  |  | <623.50 |  |  | <210.40 |  |  | <424.30
[6 empty] | (ppb or ppm) | <430.27 ppb | [8 empty] | <104.38 ppb
NEXO 5 | Part Name: MMC-S5B8 S/N: 935131424610-T3N | 0.08 g |  |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
MuRata Silicon Capacitors | Type: 3.3 uF |  |  |  |  | (mBq/kg)
| www.murata.com 100 units |  |  |  |  | (ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)
Next Sample | [5 empty] | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
(mBq/kg)
(ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)