## Collected Ge Detector Sample Re
Some Useful Information Concerning the results:
The Conversion factors for the primordial nuclides are given by: | [9 empty] | 1 Bq 238U/kg = 81 ppb U (81 x 10-9 gU/g) | [5 empty] | The 238U decay chain gammas used are: | [5 empty] | 226Ra: 186.1 keV
214Pb: 295.21 and 351.92 keV
214Bi: 609.31, 1120.29, 1764.49 and 2204.21 keV
The relationships are valid for any daughters in the 238U, 235U or 232Th chain only if the chain is in equilibrium. | [9 empty] | 1 Bq 232Th/kg = 246 ppb Th (246 x 10-9 gTh/g) | [5 empty] | The 232Th decay chain gammas used are: | [5 empty] | 212Pb: 238.63 and 300.09 keV
208Tl: 583.19 and 2614.53 keV,
228Ac: 911.21 keV
[10 empty] | 1 Bq 40K/kg = 32300 ppb K (32300 x 10-6 gK/g) | [5 empty] | The 40K decay chain gamma used is: | [5 empty] | 40K: 1460.83 keV
[10 empty] | 1 Bq 235U/kg = 1.76 ppm U (1.76 x 10-6 gU/g) | [5 empty] | The 235U decay chain gammas used are: | [5 empty] | 235U: 143.76, 163.33 and 205.31 keV
Background Measurements of the Vue des Alpes Detector
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Background 1 | Efficiency assumes no samples |  | 36.187 | 171213 171218 | 2017-12-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-01-19 | (mBq) | 1.873 | +- | 0.1674 | 5.039 | +- | 1.612 | 0.0612 | +- | 0.05173 | 1.354 | +- | 0.1679 | 15.262 | +- | 1.927 | 0.2008 | +- | 0.096 | 0.06328 | +- | 0.05422
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 76.664 | +- | 14.04 | <0.67 |  |  | <0.081 |  |  | 0.8836 | +- | 0.2577
Background 2 | Efficiency assumes no samples |  | 17.497 | 180212 | 2018-02-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-03-02 | (mBq) | 1.322 | +- | 0.2181 | 3.725 | +- | 2.318 | 0.2334 | +- | 0.07634 | 0.7936 | +- | 0.2178 | 18.386 | +- | 2.838 | 0.39164 | +- | 0.1591 | 0.02449 | +- | 0.08185
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.72 | +- | 20.73 | <0.98 |  |  | <0.12 |  |  | 1.033 | +- | 0.3662
Background 3 | Efficiency assumes no samples |  | 51.75 | 180319 180417 | 2018-03-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-05-12 | (mBq) | 2.225 | +- | 0.1474 | 4.95 | +- | 1.385 | 0.14 | +- | 0.0455 | 1.274 | +- | 0.1334 | 17.109 | +- | 1.737 | 0.43336 | +- | 0.09 | <0.068
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.15 | +- | 12.63 | 0.5036 | +- | 0.55 | 0.075 | +- | 0.065 | 1.212 | +- | 0.2177
Background 4 | Efficiency assumes no samples |  | 82.542 | 180522 18053001 180801 | 2018-05-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-08-12 | (mBq) | 2.053 | +- | 0.1186 | 6.239 | +- | 1.172 | 0.1237 | +- | 0.03392 | 1.328 | +- | 0.1141 | 15.52 | +- | 1.393 | 0.31949 | +- | 0.0674 | 0.1219 | +- | 0.03733
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 79.586 | +- | 9.876 | 0.731 | +- | 0.4578 | 0.0398 | +- | 0.04739 | 1.13 | +- | 0.18
Background 5 | Efficiency assumes no samples |  | 76.532 | 190910 190914 190915 | 2019-09-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2019-11-28 | (mBq) | 2.04 | +- | 0.122 | 6.373 | +- | 1.193 | 0.133 | +- | 0.03459 | 1.164 | +- | 0.1159 | 16.218 | +- | 1.452 | 0.38528 | +- | 0.0757 | 0.097 | +- | 0.0385
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 82.073 | +- | 10.15 | 0.48395 | +- | 0.443 | <0.052 |  |  | 1.477 | +- | 0.1925
Background 6 | Efficiency assumes no samples |  | 55.304 | 200324 200430 | 2020-03-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2020-05-19 | (mBq) | 3.397 | +- | 0.1763 | 7.801 | +- | 1.57 | 0.1535 | +- | 0.04259 | 2.404 | +- | 0.1694 | 29.088 | +- | 2.378 | 0.24865 | +- | 0.0871 | 0.03165 | +- | 0.05048
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 191.42 | +- | 21.95 | 0.48686 | +- | 0.5507 | 0.043088 | +- | 0.06009 | 2.092 | +- | 0.2595
Background 7 | Efficiency assumes no samples |  | 7.257 | 210111 210119 | 2021-01-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2021-01-20 | (mBq) | 2.063 | +- | 0.3767 | 8.573 | +- | 3.769 | 0.2188 | +- | 0.1077 | 1.139 | +- | 0.3709 | 29.316 | +- | 5.234 | 0.36096 | +- | 0.2337 | <0.12
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 172.52 | +- | 44.42 | <1.55 |  |  | <0.21 |  |  | 1.727 | +- | 0.6023
Combined Background | Combined Backgrounds of Runs 1+2+3+4+5 |  | 238.322 |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector | Efficiency assumes no samples |  |  |  |  | (mBq) | 2.414 | +- | 0.094 | 6.61 | +- | 0.85 | 0.333 | +- | 0.06 | 1.534 | +- | 0.082 | 18.6526 | +- | 1.2214 | 0.4053 | +- | 0.0522 | 0.151 | +- | 0.026
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 99.2049 | +- | 7.7138 | 0.3194 | +- | 0.3263 | 0.0063 | +- | 0.0342 | 1.394 | +- | 0.041
Completed Sample Measurements Counted on the Vue des Alpes Detector
The measurements of the samples below take into account the background measurement shown above. If a measurement is below the background then the upper bound shown is the 90% confidence limit.
CNL Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
CNL V01 |  | 49.2919 g | 4.443 | 230202 | 2023-02-02 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Pembroke Site 3 10-15 cm |  |  |  |  | 2023-02-07 | (mBq/kg) | 10880 | +- | 279.7 | 39520 | +- | 2966 | 501.2 | +- | 33.79 | 13540 | +- | 396 | 669460 | +- | 34440 | 4596.9 | +- | 254.7 | <80.57
[6 empty] | (ppb or ppm) | 881.28 ppb | +- | 22.66 ppb | 3201.12 ppb | +- | 240.25 ppb | 882.11 ppb | +- | 59.47 ppb | 3330.84 ppb | +- | 97.42 ppb | 21623.56 ppm | +- | 1112.41 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 57367 | +- | 108600 | <724.90 |  |  | 202.97 | +- | 60.98 | 13100 | +- | 432.4
[6 empty] | (ppb or ppm) | 4.65 ppm | +- | 8.8 ppm | [6 empty] | 3222.6 ppb | +- | 106.37 ppb
CNL V02 |  | 41.7266 g | 3.002 | 230207 | 2023-02-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Pembroke Site 2 10-15 cm |  |  |  |  | 2023-02-10 | (mBq/kg) | 10760 | +- | 302.5 | 26190 | +- | 2540 | 598.7 | +- | 41.24 | 10610 | +- | 348.3 | 705960 | +- | 36490 | 6556 | +- | 366.6 | <127.10
[6 empty] | (ppb or ppm) | 871.56 ppb | +- | 24.5 ppb | 2121.39 ppb | +- | 205.74 ppb | 1053.71 ppb | +- | 72.58 ppb | 2610.06 ppb | +- | 85.68 ppb | 22802.51 ppm | +- | 1178.63 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 256440 | +- | 123600 | <999.10 |  |  | 93.624 | +- | 66.66 | 10820 | +- | 418.8
[6 empty] | (ppb or ppm) | 20.77 ppm | +- | 10.01 ppm | [6 empty] | 2661.72 ppb | +- | 103.02 ppb
CNL V03 |  | 48.4052 g | 2.841 | 230210 | 2023-02-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Pembroke Site 1 10-15 cm |  |  |  |  | 2023-02-13 | (mBq/kg) | 9014 | +- | 367.5 | 29300 | +- | 2834 | 450.9 | +- | 40.66 | 12620 | +- | 403.2 | 807110 | +- | 41600 | 4355.5 | +- | 257 | <51.11
[6 empty] | (ppb or ppm) | 730.13 ppb | +- | 29.77 ppb | 2373.3 ppb | +- | 229.55 ppb | 793.58 ppb | +- | 71.56 ppb | 3104.52 ppb | +- | 99.19 ppb | 26069.65 ppm | +- | 1343.68 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 128320 | +- | 137500 | <655.20 |  |  | 139.87 | +- | 76.46 | 12320 | +- | 461
[6 empty] | (ppb or ppm) | 10.39 ppm | +- | 11.14 ppm | [6 empty] | 3030.72 ppb | +- | 113.41 ppb
CNL V04 |  | 51.0620 g | 3.899 | 230213 | 2023-02-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Lemieux Island 10-15 cm |  |  |  |  | 2023-02-17 | (mBq/kg) | 12050 | +- | 309.5 | 45830 | +- | 3314 | 746.3 | +- | 40.25 | 19810 | +- | 552.8 | 641560 | +- | 30300 | 2681.6 | +- | 166.6 | <71.39
[6 empty] | (ppb or ppm) | 976.05 ppb | +- | 25.07 ppb | 3712.23 ppb | +- | 268.43 ppb | 1313.49 ppb | +- | 70.84 ppb | 4873.26 ppb | +- | 135.99 ppb | 20722.39 ppm | +- | 978.69 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 641560 | +- | 33030 | 715.55 | +- | 458 | 361.53 | +- | 166.6 | 18830 | +- | 576.3
[6 empty] | (ppb or ppm) | 51.97 ppm | +- | 2.68 ppm | [6 empty] | 4632.18 ppb | +- | 141.77 ppb
CNL V05 |  | 39.3869 g | 3.871 | 230217 | 2023-02-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-004 10-15 cm |  |  |  |  | 2023-02-21 | (mBq/kg) | 26070 | +- | 598.9 | 77410 | +- | 5281 | 1596 | +- | 64.14 | 34040 | +- | 912.8 | 590410 | +- | 30510 | 972.83 | +- | 110.3 | 43.7 | +- | 77.43
[6 empty] | (ppb or ppm) | 2111.67 ppb | +- | 48.51 ppb | 6270.21 ppb | +- | 427.76 ppb | 2808.96 ppb | +- | 112.89 ppb | 8373.84 ppb | +- | 224.55 ppb | 19070.24 ppm | +- | 985.47 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <291400.00 |  |  | 160.19 | +- | 600 | 705.75 | +- | 94.7 | 32860 | +- | 933.7
[6 empty] | (ppb or ppm) | <23.6 ppm | [8 empty] | 8083.56 ppb | +- | 229.69 ppb
CNL V06 |  | 43.9680 g | 2.957 | 230221 | 2023-02-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-003 10-15 cm |  |  |  |  | 2023-02-24 | (mBq/kg) | 19080 | +- | 487.5 | 56920 | +- | 4889 | 1207 | +- | 62.66 | 41210 | +- | 1090 | 622680 | +- | 32210 | 4544.2 | +- | 274.3 | <85.05
[6 empty] | (ppb or ppm) | 1545.48 ppb | +- | 39.49 ppb | 4610.52 ppb | +- | 396.01 ppb | 2124.32 ppb | +- | 110.28 ppb | 10137.66 ppb | +- | 268.14 ppb | 20112.56 ppm | +- | 1040.38 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 286110 | +- | 189600 | <1059.00 |  |  | 715.21 | +- | 105.4 | 40500 | +- | 1125
[6 empty] | (ppb or ppm) | 23.17 ppm | +- | 15.36 ppm | [6 empty] | 9963 ppb | +- | 276.75 ppb
CNL V07 |  | 39.8567 g | 3.766 | 230224 230228 | 2023-02-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-002 10-15 cm |  |  |  |  | 2023-02-28 | (mBq/kg) | 11660 | +- | 310 | 32310 | +- | 2711 | 611 | +- | 38.36 | 12530 | +- | 383.4 | 371710 | +- | 19330 | 5223.6 | +- | 293.2 | 64.58 | +- | 58.44
[6 empty] | (ppb or ppm) | 944.46 ppb | +- | 25.11 ppb | 2617.11 ppb | +- | 219.59 ppb | 1075.36 ppb | +- | 67.51 ppb | 3082.38 ppb | +- | 94.32 ppb | 12006.23 ppm | +- | 624.36 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 78726 | +- | 115200 | 231.36 | +- | 444.7 | 206.38 | +- | 62.46 | 12400 | +- | 430.4
[6 empty] | (ppb or ppm) | 6.38 ppm | +- | 9.33 ppm | [6 empty] | 3050.4 ppb | +- | 105.88 ppb
CNL V08 |  | 41.9112 g | 2.915 | 29022801 | 2023-02-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-005 10-15 cm |  |  |  |  | 2023-03-03 | (mBq/kg) | 23300 | +- | 564.1 | 64860 | +- | 4750 | 1431 | +- | 65.68 | 29120 | +- | 814.2 | 333920 | +- | 27700 | 7712.2 | +- | 427.6 | <167.10
[6 empty] | (ppb or ppm) | 1887.3 ppb | +- | 45.69 ppb | 5253.66 ppb | +- | 384.75 ppb | 2518.56 ppb | +- | 115.6 ppb | 7163.52 ppb | +- | 200.29 ppb | 10785.62 ppm | +- | 894.71 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 119830 | +- | 180100 | 1389.2 | +- | 687.1 | 294.54 | +- | 91.79 | 28480 | +- | 856.4
[6 empty] | (ppb or ppm) | 9706.23 ppb | +- | 14588.1 ppb | [6 empty] | 7006.08 ppb | +- | 210.67 ppb
CNL V09 |  | 41.9332 g | 3.853 | 230303 | 2023-03-03 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-006 10-15 cm |  |  |  |  | 2023-03-07 | (mBq/kg) | 25500 | +- | 581.2 | 68240 | +- | 4574 | 1435 | +- | 58.89 | 27010 | +- | 739.8 | 539400 | +- | 27910 | 1702.6 | +- | 131.9 | <93.04
[6 empty] | (ppb or ppm) | 2065.5 ppb | +- | 47.08 ppb | 5527.44 ppb | +- | 370.49 ppb | 2525.6 ppb | +- | 103.65 ppb | 6644.46 ppb | +- | 181.99 ppb | 17422.62 ppm | +- | 901.49 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 19616 | +- | 155700 | <1074.00 |  |  | 403.79 | +- | 82.89 | 26630 | +- | 780.6
[6 empty] | (ppb or ppm) | 1.59 ppm | +- | 12.61 ppm | [6 empty] | 6550.98 ppb | +- | 192.03 ppb
CNL V10 |  | 49.9195 g | 2.724 | 230307 | 2023-03-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Princess Sodalite Mine 10-15 cm |  |  |  |  | 2023-03-10 | (mBq/kg) | 22930 | +- | 552.9 | 63690 | +- | 4739 | 1049 | +- | 57.99 | 26130 | +- | 746.8 | 828250 | +- | 42690 | 62.474 | +- | 93.55 | <150.30
[6 empty] | (ppb or ppm) | 1857.33 ppb | +- | 44.78 ppb | 5158.89 ppb | +- | 383.86 ppb | 1846.24 ppb | +- | 102.06 ppb | 6427.98 ppb | +- | 183.71 ppb | 26752.48 ppm | +- | 1378.89 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <252500.00 |  |  | 1041 | +- | 655.6 | 377.08 | +- | 98.05 | 27410 | +- | 831
[6 empty] | (ppb or ppm) | <20.45 ppm | [8 empty] | 6742.86 ppb | +- | 204.43 ppb
CNL V11 |  | 43.4597 g | 6.3 | 230310 230314 | -619747 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Silver Crater Mine 10-15 cm |  |  |  |  | 2023-03-17 | (mBq/kg) | 99480 | +- | 1861 | 341400 | +- | 15640 | 8150 | +- | 186.3 | 45590 | +- | 1156 | 599190 | +- | 30810 | 5185.4 | +- | 287.6 | <169.50
[6 empty] | (ppb or ppm) | 8057.88 ppb | +- | 150.74 ppb | 27653.4 ppb | +- | 1266.84 ppb | 14344 ppb | +- | 327.89 ppb | 11215.14 ppb | +- | 284.38 ppb | 19353.84 ppm | +- | 995.16 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 349250 | +- | 196200 | 592.52 | +- | 637.6 | 730.47 | +- | 95.68 | 39990 | +- | 1103
[6 empty] | (ppb or ppm) | 28.29 ppm | +- | 15.89 ppm | [6 empty] | 9837.54 ppb | +- | 271.34 ppb
CNL V12 |  | 29.3726 g | 3.897 | 230317 | 2023-03-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Silver Crate Mine 5-10 cm |  |  |  |  | 2023-03-21 | (mBq/kg) | 164900 | +- | 3116 | 495300 | +- | 23130 | 11780 | +- | 279.8 | 52230 | +- | 1402 | 592680 | +- | 30740 | 9499.8 | +- | 527.6 | <237.60
[6 empty] | (ppb or ppm) | 13356.9 ppb | +- | 252.4 ppb | 40119.3 ppb | +- | 1873.53 ppb | 20732.8 ppb | +- | 492.45 ppb | 12848.58 ppb | +- | 344.89 ppb | 19143.56 ppm | +- | 992.9 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 117790 | +- | 306300 | <2320.00 |  |  | 849.02 | +- | 151.5 | 47310 | +- | 1409
[6 empty] | (ppb or ppm) | 9.54 ppm | +- | 24.81 ppm | [6 empty] | 11638.26 ppb | +- | 346.61 ppb
CNL V13 |  | 35.6399 g | 2.878 | 230321 | 2023-03-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-2 5-10 cm |  |  |  |  | 2023-03-24 | (mBq/kg) | 16170 | +- | 434.5 | 42940 | +- | 3753 | 733.1 | +- | 53.65 | 14460 | +- | 467.7 | 719350 | +- | 37270 | 5218.8 | +- | 313.4 | <188.70
[6 empty] | (ppb or ppm) | 1309.77 ppb | +- | 35.19 ppb | 3478.14 ppb | +- | 303.99 ppb | 1290.26 ppb | +- | 94.42 ppb | 3557.16 ppb | +- | 115.05 ppb | 23235.01 ppm | +- | 1203.82 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 287490 | +- | 175600 | <962.90 |  |  | 260.74 | +- | 95.67 | 13920 | +- | 543.2
[6 empty] | (ppb or ppm) | 23.29 ppm | +- | 14.22 ppm | [6 empty] | 3424.32 ppb | +- | 133.63 ppb
CNL V14 |  | 17.3276 g | 3.869 | 230324 | 2023-03-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-8 5-10 cm |  |  |  |  | 2023-03-28 | (mBq/kg) | 7913 | +- | 272.5 | 25650 | +- | 2906 | 485.3 | +- | 48.01 | 8285 | +- | 325.4 | 450450 | +- | 23810 | 16529 | +- | 889.6 | <136.30
[6 empty] | (ppb or ppm) | 640.95 ppb | +- | 22.07 ppb | 2077.65 ppb | +- | 235.39 ppb | 854.13 ppb | +- | 84.5 ppb | 2038.11 ppb | +- | 80.05 ppb | 14549.54 ppm | +- | 769.06 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 286810 | +- | 154000 | 1499.8 | +- | 717.5 | 96.874 | +- | 85.97 | 7417 | +- | 410.4
[6 empty] | (ppb or ppm) | 23.23 ppm | +- | 12.47 ppm | [6 empty] | 1824.58 ppb | +- | 100.96 ppb
CNL V15 |  | 40.0602 g | 2.916 | 230328 | 2023-03-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-5 5-10 cm |  |  |  |  | 2023-03-31 | (mBq/kg) | 16440 | +- | 425.1 | 35480 | +- | 3479 | 764.6 | +- | 50.13 | 10810 | +- | 366.7 | 704010 | +- | 36420 | 4833.4 | +- | 288.7 | <138.00
[6 empty] | (ppb or ppm) | 1331.64 ppb | +- | 34.43 ppb | 2873.88 ppb | +- | 281.8 ppb | 1345.7 ppb | +- | 88.23 ppb | 2659.26 ppb | +- | 90.21 ppb | 22739.52 ppm | +- | 1176.37 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 175180 | +- | 155600 | 716.91 | +- | 608.1 | 130.04 | +- | 84.14 | 10500 | +- | 441.8
[6 empty] | (ppb or ppm) | 14.19 ppm | +- | 12.6 ppm | [6 empty] | 2583 ppb | +- | 108.68 ppb
CNL V16 |  | 45.9848 g | 3.923 | 230331 | 2023-03-31 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-9 10-15 cm |  |  |  |  | 2023-04-04 | (mBq/kg) | 12370 | +- | 319.5 | 32200 | +- | 2829 | 582.3 | +- | 37.65 | 11900 | +- | 367.9 | 663670 | +- | 34190 | 3926.3 | +- | 227.9 | <65.56
[6 empty] | (ppb or ppm) | 1001.97 ppb | +- | 25.88 ppb | 2608.2 ppb | +- | 229.15 ppb | 1024.85 ppb | +- | 66.26 ppb | 2927.4 ppb | +- | 90.5 ppb | 21436.54 ppm | +- | 1104.34 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 41530 | +- | 119000 | <709.80 |  |  | 165.29 | +- | 65 | 11400 | +- | 412
[6 empty] | (ppb or ppm) | 3363.93 ppb | +- | 9639 ppb | [6 empty] | 2804.4 ppb | +- | 101.35 ppb
CNL V17 |  | 5.6586 g | 6.755 | 230404 | 2023-04-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-7 0-5 cm |  |  |  |  | 2023-04-11 | (mBq/kg) | 3382 | +- | 206.7 | 17150 |  | 2924 | 243.2 | +- | 52.88 | 6710 | +- | 322.6 | 170070 | +- | 9840 | 21146 | +- | 1156 | <73.91
[6 empty] | (ppb or ppm) | 273.94 ppb | +- | 16.74 ppb | 1389.15 ppb | +- | 236.84 ppb | 428.03 ppb | +- | 93.07 ppb | 1650.66 ppb | +- | 79.36 ppb | 5493.26 ppm | +- | 317.83 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 3012500 | +- | 393900 | <2021.00 |  |  | 250.25 | +- | 85.25 | 6325 | +- | 433.4
[6 empty] | (ppb or ppm) | 244.01 ppm | +- | 31.91 ppm | [6 empty] | 1555.95 ppb | +- | 106.62 ppb
CNL V18 |  | 42.6733 g | 2.64 | 230411 | 2023-04-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-9 10-15 cm |  |  |  |  | 2023-04-14 | (mBq/kg) | 8319 | +- | 256.6 | 26760 | +- | 2724 | 417.3 | +- | 41.61 | 8318 | +- | 298.6 | 618890 | +- | 32080 | 8730.7 | +- | 479.1 | <177.70
[6 empty] | (ppb or ppm) | 673.84 ppb | +- | 20.78 ppb | 2167.56 ppb | +- | 220.64 ppb | 734.45 ppb | +- | 73.23 ppb | 2046.23 ppb | +- | 73.46 ppb | 19990.15 ppm | +- | 1036.18 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 108290 | +- | 137400 | 311.44 | +- | 465.1 | 184.44 | +- | 77.41 | 7448 | +- | 367.6
[6 empty] | (ppb or ppm) | 8.77 ppm | +- | 11.13 ppm | [6 empty] | 1832.21 ppb | +- | 90.43 ppb
CNL V19 |  | 41.4027 g | 4.92 | 230414 | 2023-04-14 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-6 10-15 cm |  |  |  |  | 2023-04-19 | (mBq/kg) | 9160 | +- | 247.7 | 22370 | +- | 2222 | 459.4 | +- | 32.23 | 9486 | +- | 298.2 | 647360 | +- | 33310 | 4011.6 | +- | 227.8 | <56.49
[6 empty] | (ppb or ppm) | 741.96 ppb | +- | 20.06 ppb | 1811.97 ppb | +- | 179.98 ppb | 808.54 ppb | +- | 56.72 ppb | 2333.56 ppb | +- | 73.36 ppb | 20909.73 ppm | +- | 1075.91 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 124390 | +- | 105800 | 540.29 | +- | 412.5 | 173.6 | +- | 58.73 | 9330 | +- | 347.1
[6 empty] | (ppb or ppm) | 10.08 ppm | +- | 8.57 ppm | [6 empty] | 2295.18 ppb | +- | 85.39 ppb
CNL V20 |  | 29.3445 g | 5.894 | 230419 | 2023-04-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-8 5-10 cm |  |  |  |  | 2023-04-25 | (mBq/kg) | 12290 | +- | 313.1 | 27040 | +- | 2374 | 532 | +- | 35.45 | 10870 | +- | 335.9 | 564900 | +- | 29100 | 13061 | +- | 677.9 | 66.18 | +- | 63.58
[6 empty] | (ppb or ppm) | 995.49 ppb | +- | 25.36 ppb | 2190.24 ppb | +- | 192.29 ppb | 936.32 ppb | +- | 62.39 ppb | 2674.02 ppb | +- | 82.63 ppb | 18246.27 ppm | +- | 939.93 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 145200 | +- | 107100 | <601.50 |  |  | 203.01 | +- | 60.09 | 9552 | +- | 371.8
[6 empty] | (ppb or ppm) | 11.76 ppm | +- | 8.68 ppm | [6 empty] | 2349.79 ppb | +- | 91.46 ppb
CNL V21 |  | 27.2959 g | 6.926 | 230425 | 2023-04-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-1 0-5 cm |  |  |  |  | 2023-05-02 | (mBq/kg) | 13010 | +- | 325 | 35750 | +- | 2724 | 602.8 | +- | 36.84 | 12570 | +- | 376.4 | 663360 | +- | 34110 | 18312 | +- | 938.5 | 63.93 | +- | 67.23
[6 empty] | (ppb or ppm) | 1053.81 ppb | +- | 26.33 ppb | 2895.75 ppb | +- | 220.64 ppb | 1060.93 ppb | +- | 64.84 ppb | 3092.22 ppb | +- | 92.59 ppb | 21426.53 ppm | +- | 1101.75 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 343740 | +- | 119000 | 1077 | +- | 490.8 | 71.448 | +- | 60.62 | 11610 | +- | 417.3
[6 empty] | (ppb or ppm) | 27.84 ppm | +- | 9.64 ppm | [6 empty] | 2856.06 ppb | +- | 102.66 ppb
In Progress Sample Measurements Counted on the Vue des Alpes Detector
In Progress and To Be Measured:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Next Sample | [5 empty] | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
(mBq/kg)
(ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)