## Collected Ge Detector Sample Re
Some Useful Information Concerning the results:
The Conversion factors for the primordial nuclides are given by: | [9 empty] | 1 Bq 238U/kg = 81 ppb U (81 x 10-9 gU/g) | [5 empty] | The 238U decay chain gammas used are: | [5 empty] | 226Ra: 186.1 keV
214Pb: 295.21 and 351.92 keV
214Bi: 609.31, 1120.29, 1764.49 and 2204.21 keV
The relationships are valid for any daughters in the 238U, 235U or 232Th chain only if the chain is in equilibrium. | [9 empty] | 1 Bq 232Th/kg = 246 ppb Th (246 x 10-9 gTh/g) | [5 empty] | The 232Th decay chain gammas used are: | [5 empty] | 212Pb: 238.63 and 300.09 keV
208Tl: 583.19 and 2614.53 keV,
228Ac: 911.21 keV
[10 empty] | 1 Bq 40K/kg = 32300 ppb K (32300 x 10-6 gK/g) | [5 empty] | The 40K decay chain gamma used is: | [5 empty] | 40K: 1460.83 keV
[10 empty] | 1 Bq 235U/kg = 1.76 ppm U (1.76 x 10-6 gU/g) | [5 empty] | The 235U decay chain gammas used are: | [5 empty] | 235U: 143.76, 163.33 and 205.31 keV
Background Measurements of the Vue des Alpes Detector
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Background 1 | Efficiency assumes no samples |  | 36.187 | 171213 171218 | 2017-12-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-01-19 | (mBq) | 1.873 | +- | 0.1674 | 5.039 | +- | 1.612 | 0.0612 | +- | 0.05173 | 1.354 | +- | 0.1679 | 15.262 | +- | 1.927 | 0.2008 | +- | 0.096 | 0.06328 | +- | 0.05422
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 76.664 | +- | 14.04 | <0.67 |  |  | <0.081 |  |  | 0.8836 | +- | 0.2577
Background 2 | Efficiency assumes no samples |  | 17.497 | 180212 | 2018-02-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-03-02 | (mBq) | 1.322 | +- | 0.2181 | 3.725 | +- | 2.318 | 0.2334 | +- | 0.07634 | 0.7936 | +- | 0.2178 | 18.386 | +- | 2.838 | 0.39164 | +- | 0.1591 | 0.02449 | +- | 0.08185
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.72 | +- | 20.73 | <0.98 |  |  | <0.12 |  |  | 1.033 | +- | 0.3662
Background 3 | Efficiency assumes no samples |  | 51.75 | 180319 180417 | 2018-03-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-05-12 | (mBq) | 2.225 | +- | 0.1474 | 4.95 | +- | 1.385 | 0.14 | +- | 0.0455 | 1.274 | +- | 0.1334 | 17.109 | +- | 1.737 | 0.43336 | +- | 0.09 | <0.068
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.15 | +- | 12.63 | 0.5036 | +- | 0.55 | 0.075 | +- | 0.065 | 1.212 | +- | 0.2177
Background 4 | Efficiency assumes no samples |  | 82.542 | 180522 18053001 180801 | 2018-05-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-08-12 | (mBq) | 2.053 | +- | 0.1186 | 6.239 | +- | 1.172 | 0.1237 | +- | 0.03392 | 1.328 | +- | 0.1141 | 15.52 | +- | 1.393 | 0.31949 | +- | 0.0674 | 0.1219 | +- | 0.03733
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 79.586 | +- | 9.876 | 0.731 | +- | 0.4578 | 0.0398 | +- | 0.04739 | 1.13 | +- | 0.18
Background 5 | Efficiency assumes no samples |  | 76.532 | 190910 190914 190915 | 2019-09-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2019-11-28 | (mBq) | 2.04 | +- | 0.122 | 6.373 | +- | 1.193 | 0.133 | +- | 0.03459 | 1.164 | +- | 0.1159 | 16.218 | +- | 1.452 | 0.38528 | +- | 0.0757 | 0.097 | +- | 0.0385
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 82.073 | +- | 10.15 | 0.48395 | +- | 0.443 | <0.052 |  |  | 1.477 | +- | 0.1925
Background 6 | Efficiency assumes no samples |  | 55.304 | 200324 200430 | 2020-03-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2020-05-19 | (mBq) | 3.397 | +- | 0.1763 | 7.801 | +- | 1.57 | 0.1535 | +- | 0.04259 | 2.404 | +- | 0.1694 | 29.088 | +- | 2.378 | 0.24865 | +- | 0.0871 | 0.03165 | +- | 0.05048
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 191.42 | +- | 21.95 | 0.48686 | +- | 0.5507 | 0.043088 | +- | 0.06009 | 2.092 | +- | 0.2595
Background 7 | Efficiency assumes no samples |  | 7.257 | 210111 210119 | 2021-01-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2021-01-20 | (mBq) | 2.063 | +- | 0.3767 | 8.573 | +- | 3.769 | 0.2188 | +- | 0.1077 | 1.139 | +- | 0.3709 | 29.316 | +- | 5.234 | 0.36096 | +- | 0.2337 | <0.12
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 172.52 | +- | 44.42 | <1.55 |  |  | <0.21 |  |  | 1.727 | +- | 0.6023
Combined Background | Combined Backgrounds of Runs 1+2+3+4+5 |  | 238.322 |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector | Efficiency assumes no samples |  |  |  |  | (mBq) | 2.414 | +- | 0.094 | 6.61 | +- | 0.85 | 0.333 | +- | 0.06 | 1.534 | +- | 0.082 | 18.6526 | +- | 1.2214 | 0.4053 | +- | 0.0522 | 0.151 | +- | 0.026
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 99.2049 | +- | 7.7138 | 0.3194 | +- | 0.3263 | 0.0063 | +- | 0.0342 | 1.394 | +- | 0.041
Completed Sample Measurements Counted on the Vue des Alpes Detector
The measurements of the samples below take into account the background measurement shown above. If a measurement is below the background then the upper bound shown is the 90% confidence limit.
CUTE Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
CUTE V01 | CUTE Source SRS-19-003 | 194.0 g | 6.791 | 210120 | 2021-01-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
CUTE Fe-55 Calibration Sample | 55-Fe source inside a copper holder |  |  |  | 2021-01-27 | (mBq/kg) | <1.44 |  |  | <247.60 |  |  | 6.192 | +- | 3.943 | 7.232 | +- | 2.439 | 7.669 | +- | 20.01 | <1.56 |  |  | 175.3 | +- | 8.401
| The source strength is 222 kBq |  |  |  |  | (ppb or ppm) | <0.12 ppb |  |  | <20.06 ppb |  |  | 10.9 ppb | +- | 6.94 ppb | 1.78 ppb | +- | 0.6 ppb | 247.71 ppb | +- | 646.32 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | <639400.00 | High limit due to large number of 55-Fe X-rays |  | 5.8584 | +- | 9.093 | <2.18 |  |  | 4.19 | +- | 4.472 | <53.66 |  |  | <1.79 |  |  | + large number of X-Rays from 55-Fe
[6 empty] | (ppb or ppm) | <51.79 ppm | [8 empty] | 1.03 ppb | +- | 1.1 ppb
CUTE V02 | Loctite Epoxy by Henkel | 52.3 g | 5.861 | 210415 | 2021-04-15 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Loctite EA-0151 Epoxy | PN: 398454 |  |  |  | 2021-04-21 | (mBq/kg) | <8.65 |  |  | <64.10 |  |  | <3.01 |  |  | <9.01 |  |  | <78.76 |  |  | <4.56 |  |  | <6.37
| Batch No: LB0KAA3587 21/043 |  |  |  |  | (ppb or ppm) | <0.7 ppb |  |  | <5.19 ppb |  |  | <5.3 ppb |  |  | <2.22 ppb |  |  | <2.54 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 897.72 | +- | 981.5 | <74.36 |  |  | <6.39 |  |  | <27.55
[6 empty] | (ppb or ppm) | 72.72 ppb | +- | 79.5 ppb | [6 empty] | <6.78 ppb
CUTE V03 |  | 51.6 g | 7.883 | 240411 | 2024-04-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
G10 Plastic Tubes & Inserts |  |  |  |  | 2024-04-19 | (mBq/kg) | 8100 | +- | 184.2 | 25000 | +- | 1808 | 397.2 | +- | 19.02 | 20790 | +- | 512.1 | 16364 | +- | 950 | <26.87 |  |  | <14.42
[6 empty] | (ppb or ppm) | 656.1 ppb | +- | 14.92 ppb | 2.03 ppm | +- | 0.15 ppm | 699.07 ppb | +- | 33.48 ppb | 5.11 ppm | +- | 0.13 ppm | 528.56 ppm | +- | 30.69 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 30602 | +- | 83100 | 354.71 | +- | 202 | 236.28 | +- | 30.01 | 18730 | +- | 476.9
[6 empty] | (ppb or ppm) | 2.48 ppm | +- | 6.73 ppm | [6 empty] | 4.61 ppm | +- | 0.12 ppm
CUTE V04 |  | 16.9 g | 5.778 | 240419 240425 | 2024-04-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
G10 Plastic Tubes |  |  |  |  | 2024-04-25 | (mBq/kg) | 10710 | +- | 323.5 | 39930 | +- | 3861 | 527.3 | +- | 51.78 | 27360 | +- | 793.2 | 22372 | +- | 1778 | <86.57 |  |  | <78.20
[6 empty] | (ppb or ppm) | 867.51 ppb | +- | 26.2 ppb | 3.23 ppm | +- | 0.31 ppm | 928.05 ppb | +- | 91.13 ppb | 6.73 ppm | +- | 0.2 ppm | 722.62 ppm | +- | 57.43 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 100830 | +- | 253700 | <1163.00 |  |  | 415.95 | +- | 80.19 | 25420 | +- | 812.2
[6 empty] | (ppb or ppm) | 8.17 ppm | +- | 20.55 ppm | [6 empty] | 6.25 ppm | +- | 0.2 ppm
In Progress Sample Measurements Counted on the Vue des Alpes Detector
In Progress and To Be Measured:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Next Sample | [5 empty] | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
(mBq/kg)
(ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)