## Collected Ge Detector Sample Re
Some Useful Information Concerning the results:
The Conversion factors for the primordial nuclides are given by: | [9 empty] | 1 Bq 238U/kg = 81 ppb U (81 x 10-9 gU/g) | [5 empty] | The 238U decay chain gammas used are: | [5 empty] | 226Ra: 186.1 keV
214Pb: 295.21 and 351.92 keV
214Bi: 609.31, 1120.29, 1764.49 and 2204.21 keV
The relationships are valid for any daughters in the 238U, 235U or 232Th chain only if the chain is in equilibrium. | [9 empty] | 1 Bq 232Th/kg = 246 ppb Th (246 x 10-9 gTh/g) | [5 empty] | The 232Th decay chain gammas used are: | [5 empty] | 212Pb: 238.63 and 300.09 keV
208Tl: 583.19 and 2614.53 keV,
228Ac: 911.21 keV
[10 empty] | 1 Bq 40K/kg = 32300 ppb K (32300 x 10-6 gK/g) | [5 empty] | The 40K decay chain gamma used is: | [5 empty] | 40K: 1460.83 keV
[10 empty] | 1 Bq 235U/kg = 1.76 ppm U (1.76 x 10-6 gU/g) | [5 empty] | The 235U decay chain gammas used are: | [5 empty] | 235U: 143.76, 163.33 and 205.31 keV
Background Measurements of the Vue des Alpes Detector
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Background 1 | Efficiency assumes no samples |  | 36.187 | 171213 171218 | 2017-12-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-01-19 | (mBq) | 1.873 | +- | 0.1674 | 5.039 | +- | 1.612 | 0.0612 | +- | 0.05173 | 1.354 | +- | 0.1679 | 15.262 | +- | 1.927 | 0.2008 | +- | 0.096 | 0.06328 | +- | 0.05422
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 76.664 | +- | 14.04 | <0.67 |  |  | <0.081 |  |  | 0.8836 | +- | 0.2577
Background 2 | Efficiency assumes no samples |  | 17.497 | 180212 | 2018-02-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-03-02 | (mBq) | 1.322 | +- | 0.2181 | 3.725 | +- | 2.318 | 0.2334 | +- | 0.07634 | 0.7936 | +- | 0.2178 | 18.386 | +- | 2.838 | 0.39164 | +- | 0.1591 | 0.02449 | +- | 0.08185
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
Background 3 | Efficiency assumes no samples |  | 51.75 | 180319 180417 | 2018-03-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-05-12 | (mBq) | 2.225 | +- | 0.1474 | 4.95 | +- | 1.385 | 0.14 | +- | 0.0455 | 1.274 | +- | 0.1334 | 17.109 | +- | 1.737 | 0.43336 | +- | 0.09 | <0.068
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.15 | +- | 12.63 | 0.5036 | +- | 0.55 | 0.075 | +- | 0.065 | 1.212 | +- | 0.2177
Background 4 | Efficiency assumes no samples |  | 82.542 | 180522 18053001 180801 | 2018-05-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-08-12 | (mBq) | 2.053 | +- | 0.1186 | 6.239 | +- | 1.172 | 0.1237 | +- | 0.03392 | 1.328 | +- | 0.1141 | 15.52 | +- | 1.393 | 0.31949 | +- | 0.0674 | 0.1219 | +- | 0.03733
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 79.586 | +- | 9.876 | 0.731 | +- | 0.4578 | 0.0398 | +- | 0.04739 | 1.13 | +- | 0.18
Background 5 | Efficiency assumes no samples |  | 76.532 | 190910 190914 190915 | 2019-09-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2019-11-28 | (mBq) | 2.04 | +- | 0.122 | 6.373 | +- | 1.193 | 0.133 | +- | 0.03459 | 1.164 | +- | 0.1159 | 16.218 | +- | 1.452 | 0.38528 | +- | 0.0757 | 0.097 | +- | 0.0385
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 82.073 | +- | 10.15 | 0.48395 | +- | 0.443 | <0.052 |  |  | 1.477 | +- | 0.1925
Background 6 | Efficiency assumes no samples |  | 55.304 | 200324 200430 | 2020-03-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2020-05-19 | (mBq) | 3.397 | +- | 0.1763 | 7.801 | +- | 1.57 | 0.1535 | +- | 0.04259 | 2.404 | +- | 0.1694 | 29.088 | +- | 2.378 | 0.24865 | +- | 0.0871 | 0.03165 | +- | 0.05048
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 191.42 | +- | 21.95 | 0.48686 | +- | 0.5507 | 0.043088 | +- | 0.06009 | 2.092 | +- | 0.2595
Background 7 | Efficiency assumes no samples |  | 7.257 | 210111 210119 | 2021-01-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2021-01-20 | (mBq) | 2.063 | +- | 0.3767 | 8.573 | +- | 3.769 | 0.2188 | +- | 0.1077 | 1.139 | +- | 0.3709 | 29.316 | +- | 5.234 | 0.36096 | +- | 0.2337 | <0.12
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 172.52 | +- | 44.42 | <1.55 |  |  | <0.21 |  |  | 1.727 | +- | 0.6023
Background 8 |  | N/A | 67.574 | 210517 210524 210627 210630 | 2021-05-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2021-07-25 | (mBq) | 1.906 | +- | 0.1261 | 4.782 | +- | 1.185 | 0.179 | +- | 0.03537 | 1.278 | +- | 0.1224 | 19.493 | +- | 1.701 | 0.38762 | +- | 0.08147 | 0.03217 | +- | 0.04114
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 141.39 | +- | 16.91 | <1.25 |  |  | <0.061 |  |  | 1.183 | +- | 0.2005
Combined Background | Combined Backgrounds of Runs 1+2+3+4+5 |  | 238.322 |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector | Efficiency assumes no samples |  |  |  |  | (mBq) | 2.414 | +- | 0.094 | 6.61 | +- | 0.85 | 0.333 | +- | 0.06 | 1.534 | +- | 0.082 | 18.6526 | +- | 1.2214 | 0.4053 | +- | 0.0522 | 0.151 | +- | 0.026
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 99.2049 | +- | 7.7138 | 0.3194 | +- | 0.3263 | 0.0063 | +- | 0.0342 | 1.394 | +- | 0.041
Completed Sample Measurements Counted on the Vue des Alpes Detector
The measurements of the samples below take into account the background measurement shown above. If a measurement is below the background then the upper bound shown is the 90% confidence limit.
SNOLAB Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
SNOLAB V01 | Background is NOT subtracted |  | 20.583 | 170901 | 2017-09-01 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Acrylic Sample Beaker H | Efficiency assumes no samples |  |  |  | 2017-09-22 | (mBq) | 1.439 | +- | 0.145 | 0.182 | +- | 0.729 | 0.354 | +- | 0.104 | 0.798 | +- | 0.114 | 7.886 | +- | 1.318 | 0.197 | +- | 0.082 | 0.04 | +- | 0.048
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 32.615 | +- | 5.256 | 0.38 | +- | 0.563 | <0.052 |  |  | 0.455 | +- | 0.199
SNOLAB V01A | Background Subtracted |  | 20.853 | 170901 | 2017-09-01 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Acrylic Sample Beaker H | Efficiency assumes no samples |  |  |  | 2017-09-22 | (mBq) | 0.294 | +- | 0.142 | <0.095 |  |  | 0.149 | +- | 0.105 | 0.007 | +- | 0.115 | <0.26 |  |  | <0.096 |  |  | 0.012 | +- | 0.05
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 12.805 | +- | 5.104 | 0.38 | +- | 0.563 | <0.052 |  |  | <0.082
SNOLAB V02 | Lot No. A4B8 02/20 | 1871.5 g | 6.979 | 180511 | 2018-05-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Alconox | Alconox Inc. |  |  |  | 2018-05-18 | (mBq/kg) | 184.6 | +- | 4.642 | 104.5 | +- | 18.43 | 5.728 | +- | 0.54 | 139.5 | +- | 4.386 | 8664 | +- | 446.7 | <0.65 |  |  | 0.8054 | +- | 0.7874
[6 empty] | (ppb or ppm) | 14.95 ppb | +- | 0.38 ppb | 8.46 ppb | +- | 1.49 ppb | 10.08 ppb | +- | 0.95 ppb | 34.32 ppb | +- | 1.08 ppb | 279.85 ppm | +- | 14.43 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <79.93 |  |  | 5.6059 | +- | 5.455 | <1.21 |  |  | 61.58 | +- | 3.484
[6 empty] | (ppb or ppm) | <6.47 ppb | [8 empty] | 15.15 ppb | +- | 0.86 ppb
SNOLAB V03 | Sterling AP 308/3-8L, AWS E308-16, E308L-16 | 2294.9 g | 3.91 | 180518 | 2018-05-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Premium Welding Electrodes --- Calibration | McKay Welding Products, Troy, Ohio, Lot No 24L056, Heat No. 738228 |  |  |  | 2018-05-22 | (mBq/kg) | 32330 | +- | 564.6 | 15260 | +- | 854.5 | 1110 | +- | 24.75 | 4810 | +- | 116.1 | 261860 | +- | 13380 | <1.68 |  |  | <6.29
[6 empty] | (ppb or ppm) | 2.62 ppm | +- | 0.05 ppm | 1.24 ppm | +- | 0.07 ppm | 1.95 ppm | +- | 0.04 ppm | 1.18 ppm | +- | 0.03 ppm | 8458.08 ppm | +- | 432.17 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 414.69 | +- | 1085 |  | NM |  |  | NM |  | 4376 | +- | 120.2
[6 empty] | (ppb or ppm) | 33.59 ppb | +- | 87.89 ppb | [6 empty] | 1.08 ppm | +- | 0.03 ppm
SNOLAB V04 | Sources prepared on Nov 21, 2019 | 4.4 g | 12.638 | 191129 | 2019-11-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Calibration Source Wipes | The Kimwipes are from sources SRS-19-004 |  | Standard Background Subtraction Using Combined Backgrounds |  | 2019-12-11 | (mBq/kg) | 455.6 | +- | 79.18 | 1335 | +- | 733.5 | 12.51 | +- | 22.94 | 300.5 | +- | 76.03 | 5534.6 | +- | 1118 | <48.00 |  |  | 11.39 | +- | 31.69
[6 empty] | (ppb or ppm) | 36.9 ppb | +- | 6.41 ppb | 108.14 ppb | +- | 59.41 ppb | 22.02 ppb | +- | 40.37 ppb | 73.92 ppb | +- | 18.7 ppb | 178.77 ppm | +- | 36.11 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 986.42 | +- | 6241 | <229.50 |  |  | 28.188 | +- | 26.26 | 99.8 | +- | 116.5
[6 empty] | (ppb or ppm) | 79.9 ppb | +- | 505.52 ppb | [6 empty] | 24.55 ppb | +- | 28.66 ppb
SNOLAB V04A | Sources prepared on Nov 21, 2019 | 4.4 g | 12.638 | 191129 | 2019-11-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Calibration Source Wipes | The Kimwipes are from sources SRS-19-004 |  | Background Subtraction Using The Results from SNOLAB 05 |  | 2019-12-11 | (mBq/kg) | 8.961 | +- | 114.2 | 264.9 | +- | 989.3 | <23.57 |  |  | 18.4 | +- | 111.7 | 0.00306 | +- | 1567 | <137.40 |  |  | 0.0228 | +- | 38.3
[6 empty] | (ppb or ppm) | 0.73 ppb | +- | 9.25 ppb | 21.46 ppb | +- | 80.13 ppb | <41.48 ppb |  |  | 4.53 ppb | +- | 27.48 ppb | 0 ppm | +- | 50.61 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 0.0168 | +- | 24810 | <944.30 |  |  | 0.025 | +- | 47.83 | 6.347 | +- | 165.8
[6 empty] | (ppb or ppm) | 0 ppb | +- | 2009.61 ppb | [6 empty] | 1.56 ppb | +- | 40.79 ppb
SNOLAB V05 | Sample prepared on Dec 05, 2019 | 4.4 g | 7.882 | 191211 | 2019-12-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Kimwipes Blank | These are clean Kimwipes |  | Standard Background Subtraction Using Combined Backgrounds |  | 2019-12-19 | (mBq/kg) | 40.49 | +- | 84.56 | 801.4 | +- | 823.4 | <14.59 |  |  | 55.51 | +- | 88.79 | 635.3 | +- | 990.9 | <83.35 |  |  | <43.88
[6 empty] | (ppb or ppm) | 3.28 ppb | +- | 6.85 ppb | 64.91 ppb | +- | 66.7 ppb | <25.68 ppb |  |  | 13.66 ppb | +- | 21.84 ppb | 20.52 ppm | +- | 32.01 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 20521 | +- | 11650 | 212.32 | +- | 352.6 | 17.376 | +- | 29.66 | 115.1 | +- | 142.6
[6 empty] | (ppb or ppm) | 1662.2 ppb | +- | 943.65 ppb | [6 empty] | 28.31 ppb | +- | 35.08 ppb
SNOLAB V06 | Vial cap is downwards | 0.832 g | 17.701 | 191219 191220 | 2019-12-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-K2 nEXO Calibration Sample | Specific Activity: 14.0 +- 0.09 Bq of 40K / [g RGK-1] |  |  |  | 2020-01-06 | (mBq/kg) | 4315 | +- | 513.7 | 8816 | +- | 15340 | 146.6 | +- | 180.2 | 297.1 | +- | 420.2 | 15233000 | +- | 801100 | 615.97 | +- | 484.9 | <981.60
| [11.57 – 11.72] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 349.52 ppb | +- | 41.61 ppb | 714.1 ppb | +- | 1242.54 ppb | 258.02 ppb | +- | 317.15 ppb | 73.09 ppb | +- | 103.37 ppb | 492025.9 ppm | +- | 25875.53 ppm
[6 empty] | (Bq) | [12 empty] | 12.6739 | +- | 0.50149
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 795110 | +- | 2797000 | 1644 | +- | 3001 | 594.49 | +- | 511.2 | <1864.00
[6 empty] | (ppb or ppm) | 64.4 ppm | +- | 226.56 ppm | [6 empty] | <458.54 ppb
SNOLAB V07 | Vial cap is upwards | 2.046 g | 1.937 | 200106 | 2020-01-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-U1 nEXO Calibration Sample | Specific Activity: 4.94 +- 0.03 Bq of 238U / [g RGU-1] |  |  |  | 2020-01-08 | (mBq/kg) | 5114000 | +- | 93980 | 6168000 | +- | 377900 | 231300 | +- | 5404 | 2079 | +- | 1486 | 4783.5 |  | 28500 | <637.00 |  |  | <3318.00
| [10.04 – 10.17] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 414.23 ppm | +- | 7.61 ppm | 499.61 ppm | +- | 30.61 ppm | 407.09 ppm | +- | 9.51 ppm | 511.43 ppb | +- | 365.56 ppb | 154.51 ppm | +- | 920.55 ppm
[6 empty] | (Bq) | 10.4632 | +- | 0.19228 | 12.6197 | +- | 0.77318 |  | Average: |  | 10.5888 | +- | 0.1866
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 66658000 | +- | 9102000 | <10990.00 |  |  | <960.20 |  |  | <8204.00
[6 empty] | (ppb or ppm) | 5399.3 ppm | +- | 737.26 ppm | [6 empty] | <2.02 ppm
[6 empty] | (Bq) | 136.3823 | +- | 18.6227
SNOLAB V08 | Vial cap is upwards | 3.327 g | 5 | 200108 | 2020-01-08 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-T1 nEXO Calibration Sample | Specific Activity: 3.25 +- 0.07 Bq of 232Th / [g RGTh-1] |  |  |  | 2020-01-13 | (mBq/kg) | 50150 | +- | 2512 | 511500 | +- | 73360 | 3829 | +- | 540.7 | 2961000 | +- | 68880 | 233380 | +- | 15970 | <1441.00 |  |  | <729.70
| [10.58 – 11.05] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 4.06 ppm | +- | 0.2 ppm | 41.43 ppm | +- | 5.94 ppm | 6.74 ppm | +- | 0.95 ppm | 728.41 ppm | +- | 16.94 ppm | 7538.17 ppm | +- | 515.83 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 9.8512 | +- | 0.22916
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <1159000.00 |  |  | 23356 | +- | 7448 | 26100 | +- | 1789 | 2804000 | +- | 64000
[6 empty] | (ppb or ppm) | <93.88 ppm | [8 empty] | 689.78 ppm | +- | 15.74 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 9.3289 | +- | 0.21293 |  | Average: |  | 9.5709 | +- | 0.15599
SNOLAB V09 | Vial cap is upwards | 0.832 g | 6.896 | 200113 | 2020-01-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-K2 nEXO Calibration Sample | Specific Activity: 14.0 +- 0.09 Bq of 40K / [g RGK-1] |  |  |  | 2020-01-20 | (mBq/kg) | 494.3 | +- | 686.5 | 1890 | +- | 24090 | <437.50 |  |  | 918.8 | +- | 767.2 | 15495000 | +- | 818300 | 763.17 | +- | 785.3 | <1728.00
| [11.57 – 11.72 Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 40.04 ppb | +- | 55.61 ppb | 153.09 ppb | +- | 1951.29 ppb | <0.77 ppm |  |  | 226.02 ppb | +- | 188.73 ppb | 500488.5 ppm | +- | 26431.09 ppm
[6 empty] | (Bq) | [12 empty] | 12.8918 | +- | 0.68083
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 9895400 | +- | 4772000 | 2661.3 | +- | 4678 | 180.42 | +- | 825.5 | <1627.00
[6 empty] | (ppb or ppm) | 801.53 ppm | +- | 386.53 ppm | [6 empty] | <400.24 ppb
SNOLAB V10 | G.F. Thompson Co Ltd. | 78.7 g | 2.104 | 200120 | 2020-01-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Masters Pro-Dope with Teflon | Newmarket ON |  |  |  | 2020-01-22 | (mBq/kg) | 10880 | +- | 279.9 | 23870 | +- | 2191 | 555 | +- | 34.31 | 12220 | +- | 361.1 | 53843 | +- | 3100 | <64.63 |  |  | 20.91 | +- | 27.51
Pipe Sealant | [5 empty] | (ppb or ppm) | 881.28 ppb | +- | 22.67 ppb | 1933.47 ppb | +- | 177.47 ppb | 976.8 ppb | +- | 60.39 ppb | 3006.12 ppb | +- | 88.83 ppb | 1739.13 ppm | +- | 100.13 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 113630 | +- | 135500 | 340.15 | +- | 308.3 | 82.984 | +- | 43.57 | 13140 | +- | 422.8
[6 empty] | (ppb or ppm) | 9204.03 ppb | +- | 10975.5 ppb | [6 empty] | 3232.44 ppb | +- | 104.01 ppb
SNOLAB V11 | Vial cap is upwards | 2.102 g | 4.791 | 200122 | 2020-01-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-U2 nEXO Calibration Sample | Specific Activity: 4.94 +- 0.03 Bq of 238U / [g RGU-1] |  |  |  | 2020-01-27 | (mBq/kg) | 4911000 | +- | 88240 | 5713000 | +- | 311500 | 221400 | +- | 4801 | 2384 | +- | 993.7 | <15330.00 |  |  | <336.50 |  |  | <1583.00
| [10.32 – 10.45] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 397.79 ppm | +- | 7.15 ppm | 462.75 ppm | +- | 25.23 ppm | 389.66 ppm | +- | 8.45 ppm | 586.46 ppb | +- | 244.45 ppb | <495.16 ppm
[6 empty] | (Bq) | 10.3229 | +- | 0.18548 | 12.0087 | +- | 0.62923 |  | Average: |  | 10.4577 | +- | 0.17791
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 71438000 | +- | 6386000 | <472.00 |  |  | <461.00 |  |  | 7337 | +- | 4053
[6 empty] | (ppb or ppm) | 5786.48 ppm | +- | 517.27 ppm | [6 empty] | 1804.9 ppb | +- | 997.04 ppb
[6 empty] | (Bq) | 146.1621 | +- | 13.0658
SNOLAB V12 | Vial cap is upwards | 3.180 g | 2.958 | 200127 | 2020-01-27 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-T2 nEXO Calibration Sample | Specific Activity: 3.25 +- 0.07 Bq of 232Th / [g RGTh-1] |  |  |  | 2020-01-30 | (mBq/kg) | 50670 | +- | 2986 | 787600 | +- | 99220 | 3309 | +- | 723.5 | 3087000 | +- | 72390 | 224340 | +- | 18380 | <1361.00 |  |  | <610.70
| [10.11 – 10.56] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 4.1 ppm | +- | 0.24 ppm | 63.8 ppm | +- | 8.04 ppm | 5.82 ppm | +- | 1.27 ppm | 759.4 ppm | +- | 17.81 ppm | 7246.18 ppm | +- | 593.67 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 10.2704 | +- | 0.24084
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <453000.00 |  |  | 52273 | +- | 10520 | 27111 | +- | 2143 | 2936000 | +- | 67660
[6 empty] | (ppb or ppm) | <36.69 ppm | [8 empty] | 722.26 ppm | +- | 16.64 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 9.7681 | +- | 0.2251 |  | Average: |  | 10.0023 | +- | 0.16446
SNOLAB V13 | Vial cap is upwards | 0.626 g | 4.91 | 200130 | 2020-01-30 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-K1 nEXO Calibration Sample | Specific Activity: 14.0 +- 0.09 Bq of 40K / [g RGK-1] |  |  |  | 2020-02-04 | (mBq/kg) | 2371 | +- | 1068 | 68810 | +- | 33160 | <311.40 |  |  | <1200.00 |  |  | 15444000 | +- | 836630 | <1641.00 |  |  | 104.3 | +- | 1429
| [ 8.70 – 8.82] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 192.05 ppb | +- | 86.51 ppb | 5.57 ppm | +- | 2.69 ppm | <0.55 ppm |  |  | <295.2 ppb |  |  | 498841.2 ppm | +- | 27023.15 ppm
[6 empty] | (Bq) | [12 empty] | 9.6679 | +- | 0.52373
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <562400.00 |  |  | <928.20 |  |  | <933.30 |  |  | <4932.00
[6 empty] | (ppb or ppm) | <45.55 ppm | [8 empty] | <1213.27 ppb
SNOLAB V14 | Part No: FG-p | 63.9 g | 6.021 | 200204 | 2020-02-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Plastic Process Equipment Inc. Food Grade Anti-Seize Thread Compound | Macedonia, OH, 44056 |  |  |  | 2020-02-10 | (mBq/kg) | 1519 | +- | 48.29 | 1630 | +- | 268.8 | 75.37 | +- | 7.183 | 601.7 | +- | 30.42 | 189.63 | +- | 118.3 | 7.4677 | +- | 8.364 | <4.27
Pipe Sealant Fortified with PTFE | [5 empty] | (ppb or ppm) | 123.04 ppb | +- | 3.91 ppb | 132.03 ppb | +- | 21.77 ppb | 132.65 ppb | +- | 12.64 ppb | 148.02 ppb | +- | 7.48 ppb | 6.13 ppm | +- | 3.82 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 46194 | +- | 25930 | <70.64 |  |  | <9.71 |  |  | 430.3 | +- | 38.22
[6 empty] | (ppb or ppm) | 3741.71 ppb | +- | 2100.33 ppb | [6 empty] | 105.85 ppb | +- | 9.4 ppb
SNOLAB V15 | Type: 801-16-BLU | 357.4 g | 2.819 | 200210 | 2020-02-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Push Lok Plus Muli-Purpose Hose | Parker Hannifin Corp |  |  |  | 2020-02-13 | (mBq/kg) | 1803 | +- | 51.43 | 5141 | +- | 534.4 | 106 | +- | 8.728 | 3102 | +- | 91.69 | 9333.1 | +- | 568.8 | 5.009 | +- | 9.986 | <5.85
[6 empty] | (ppb or ppm) | 146.04 ppb | +- | 4.17 ppb | 416.42 ppb | +- | 43.29 ppb | 186.56 ppb | +- | 15.36 ppb | 763.09 ppb | +- | 22.56 ppb | 301.46 ppm | +- | 18.37 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 5826 | +- | 7988 | <124.20 |  |  | 22.169 | +- | 10.2 | 3004 | +- | 101.9
[6 empty] | (ppb or ppm) | 471.91 ppb | +- | 647.03 ppb | [6 empty] | 738.98 ppb | +- | 25.07 ppb
SNOLAB V16 | ¾” Natural polypropylene | 623.5 g | 7.646 | 200213 200214 | 2020-02-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
ENPURE Polypropylene Tubing | Type: SCH-80 Natural |  |  |  | 2020-02-21 | (mBq/kg) | 6.106 | +- | 1.706 | 17.12 | +- | 23.48 | <1.34 |  |  | 12.55 | +- | 1.969 | 4.02 | +- | 15.06 | <1.20 |  |  | 1.021 | +- | 0.5066
| Lot: JJ05-B-12021321352 |  |  |  |  | (ppb or ppm) | 0.49 ppb | +- | 0.14 ppb | 1.39 ppb | +- | 1.9 ppb | <2.36 ppb |  |  | 3.09 ppb | +- | 0.48 ppb | 129.85 ppb | +- | 486.44 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1062.5 | +- | 569.2 | <6.47 |  |  | 0.87879 | +- | 0.6374 | <3.34
[6 empty] | (ppb or ppm) | 86.06 ppb | +- | 46.11 ppb | [6 empty] | <0.82 ppb
SNOLAB V17 | Clear Non-Toxic PVC with Polyester Braid and Rigid PVC Helix | 240.5 g | 3.847 | 200221 | 2020-02-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Reinforced FDA PVC Suction Hose |  |  |  |  | 2020-02-25 | (mBq/kg) | 494.3 | +- | 19.11 | 739.3 | +- | 127.2 | 26.56 | +- | 3.282 | 60.84 | +- | 8.144 | 678.07 | +- | 100.8 | <3.93 |  |  | <3.08
[6 empty] | (ppb or ppm) | 40.04 ppb | +- | 1.55 ppb | 59.88 ppb | +- | 10.3 ppb | 46.75 ppb | +- | 5.78 ppb | 14.97 ppb | +- | 2 ppb | 21.9 ppm | +- | 3.26 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2178.2 | +- | 1735 | 60.512 | +- | 28.93 | <2.38 |  |  | 48.45 | +- | 12.81
[6 empty] | (ppb or ppm) | 176.43 ppb | +- | 140.54 ppb | [6 empty] | 11.92 ppb | +- | 3.15 ppb
SNOLAB V18 | NIST Sample | 28.825 g | 12.736 | 200225 200302 | 2020-02-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Powdered Ocean Sediment Calibration Sample | Cross Comparison Sample A |  |  |  | 2020-03-09 | (mBq/kg) | 11640 | +- | 255.7 | 21440 | +- | 2071 | 588.4 | +- | 23 | 12650 | +- | 331.7 | 255240 | +- | 13160 | 7109 | +- | 366.4 | <29.11
| Sample placed directly on top of the detector |  |  |  |  | (ppb or ppm) | 942.84 ppb | +- | 20.71 ppb | 1736.64 ppb | +- | 167.75 ppb | 1035.58 ppb | +- | 40.48 ppb | 3.11 ppm | +- | 0.08 ppm | 8244.25 ppm | +- | 425.07 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 241Am
[6 empty] | (mBq/kg) | <15060.00 |  |  | <452.60 |  |  | 91.497 | +- | 29.67 | 12730 | +- | 352 | 20580 | +- | 1569
[6 empty] | (ppb or ppm) | <1219.86 ppb | [8 empty] | 3.13 ppm | +- | 0.09 ppm
SNOLAB V19 | NIST Sample | 33.454 g | 21.617 | 200519 | 2020-05-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Powdered Ocean Sediment Calibration Sample | Cross Comparison Sample B |  |  |  | 2020-06-10 | (mBq/kg) | 10980 | +- | 223 | 17630 | +- | 1702 | 571.6 | +- | 18.5 | 12060 | +- | 299.1 | 247900 | +- | 12720 | 6397.8 | +- | 352.7 | <25.82
| Sample placed directly on top of the detector |  |  |  |  | (ppb or ppm) | 889.38 ppb | +- | 18.06 ppb | 1428.03 ppb | +- | 137.86 ppb | 1006.02 ppb | +- | 32.56 ppb | 2.97 ppm | +- | 0.07 ppm | 8007.17 ppm | +- | 410.86 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 241Am
[6 empty] | (mBq/kg) | <8362.00 |  |  | <272.30 |  |  | 74.544 | +- | 20.54 | 12000 | +- | 307.1 | 28410 | +- | 1509
[6 empty] | (ppb or ppm) | <677.32 ppb | [8 empty] | 2.95 ppm | +- | 0.08 ppm
SNOLAB V20 | Lake in City of Greater Sudbury, Ontario | 993.9 g | 14.666 | 200309 | 2020-03-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Long Lake Water | Sample collected on 2019-10-14 at 16:54 |  |  |  | 2020-03-24 | (mBq/kg) | 2.151 | +- | 0.5584 | <12.40 |  |  | <0.27 |  |  | <0.61 |  |  | 52.142 | +- | 8.263 | 1.003 | +- | 0.383 | <0.15
| Sample placed in beaker “∆” |  |  |  |  | (ppb or ppm) | 174.23 ppt | +- | 45.23 ppt | <1 ppb |  |  | <0.48 ppb |  |  | <0.15 ppb |  |  | 1684.19 ppb | +- | 266.89 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1151.5 | +- | 456.7 | 6.8076 | +- | 2.466 | <0.38 |  |  | 1.243 | +- | 0.8064
[6 empty] | (ppb or ppm) | 93.27 ppb | +- | 36.99 ppb | [6 empty] | 0.31 ppb | +- | 0.2 ppb
SNOLAB V21 | Lake in City of Greater Sudbury, Ontario | 981.4 g | 17.117 | 200610 200624 | 2020-06-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Simon Lake Water | Sample collected on 2019-11-01 at 18:15 |  |  |  | 2020-06-30 | (mBq/kg) | 1.599 | +- | 0.506 | 12.4 | +- | 13.22 | <0.44 |  |  | <0.82 |  |  | 402 | +- | 25.3 | <0.46 |  |  | 0.1133 | +- | 0.2314
| Sample placed in beaker “H” |  |  |  |  | (ppb or ppm) | 129.52 ppt | +- | 40.99 ppt | 1004.4 ppt | +- | 1070.82 ppt | <0.77 ppb |  |  | <0.2 ppb |  |  | 12984.6 ppb | +- | 817.19 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1144.6 | +- | 434.2 | <2.80 |  |  | <0.27 |  |  | 4.993 | +- | 1.066
[6 empty] | (ppb or ppm) | 92.71 ppb | +- | 35.17 ppb | [6 empty] | 1.23 ppb | +- | 0.26 ppb
SNOLAB V22 | Lake in City of Greater Sudbury, Ontario | 935.2 g | 21.708 | 200630 | 2020-06-30 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Robinson Lake Water | Sample collected on 2019-10-22 at 12:43 |  |  |  | 2020-07-22 | (mBq/kg) | 5.117 | +- | 0.5798 | 21.36 | +- | 12.7 | <0.33 |  |  | 0.3155 | +- | 0.4927 | 95.614 | +- | 9.55 | 0.23 | +- | 0.2964 | <0.14
| Sample placed in beaker “□” |  |  |  |  | (ppb or ppm) | 414.48 ppt | +- | 46.96 ppt | 1.73 ppb | +- | 1.03 ppb | <0.58 ppb |  |  | 0.08 ppb | +- | 0.12 ppb | 3088.33 ppb | +- | 308.47 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 536.35 | +- | 394.1 | 0.5047 | +- | 2.041 | 0.23 | +- | 0.2964 | 1.943 | +- | 0.7858
[6 empty] | (ppb or ppm) | 43.44 ppb | +- | 31.92 ppb | [6 empty] | 0.48 ppb | +- | 0.19 ppb
SNOLAB V23 | Lake in City of Greater Sudbury, Ontario | 986.0 g | 16.659 | 200722 | 2020-07-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Meatbird Lake Water | Sample collected on 2019-10-08 at 10:50 |  |  |  | 2020-08-11 | (mBq/kg) | 0.19 | +- | 0.4465 | <13.83 |  |  | 0.3666 | +- | 0.2156 | <0.32 |  |  | 124.59 | +- | 11.09 | <0.39 |  |  | <0.21
| Sample placed in beaker “∆” |  |  |  |  | (ppb or ppm) | 15.39 ppt | +- | 36.17 ppt | <1.12 ppb |  |  | 0.65 ppb | +- | 0.38 ppb | <0.08 ppb |  |  | 4024.26 ppb | +- | 358.21 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <567.10 |  |  | 0.975 | +- | 1.909 | 0.02689 | +- | 0.1934 | 0.792 | +- | 0.7894
[6 empty] | (ppb or ppm) | <45.94 ppb | [8 empty] | 0.19 ppb | +- | 0.19 ppb
SNOLAB V24 | Lake in City of Greater Sudbury, Ontario | 842.7 g | 13.5 | 200811 20081101 | 2020-08-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Kelly Lake Water | Sample collected on 2019-10-7 at 10:30 |  |  |  | 2020-08-25 | (mBq/kg) | 3.308 | +- | 0.7805 | 11.09 | +- | 30.51 | <0.60 |  |  | 0.3551 | +- | 0.7696 | 501.76 | +- | 33.74 | 0.5963 | +- | 0.4861 | <0.24
[6 empty] | (ppb or ppm) | 267.95 ppt | +- | 63.22 ppt | 898.29 ppt | +- | 2471.31 ppt | <1.06 ppb |  |  | 0.09 ppb | +- | 0.19 ppb | 16.21 ppm | +- | 1.09 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2300.5 | +- | 1355 | 2.4308 | +- | 3.233 | <0.43 |  |  | 6.605 | +- | 1.395
[6 empty] | (ppb or ppm) | 186.34 ppb | +- | 109.76 ppb | [6 empty] | 1.62 ppb | +- | 0.34 ppb
SNOLAB V25 | Lake in City of Greater Sudbury, Ontario | 1032.2 g | 13.395 | 201006 | 2020-10-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Robinson Lake Water | Sample collected on 2020-09-19 at 14:30 |  |  |  | 2020-10-20 | (mBq/kg) | 1.353 | +- | 0.5589 | <29.58 |  |  | 0.417 | +- | 0.2379 | <1.04 |  |  | 100.12 | +- | 11.32 | 0.4959 | +- | 0.3941 | <0.16
[6 empty] | (ppb or ppm) | 109.59 ppt | +- | 45.27 ppt | <2.4 ppb |  |  | 0.73 ppb | +- | 0.42 ppb | <0.26 ppb |  |  | 3233.88 ppb | +- | 365.64 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 881.22 | +- | 989.9 | 10.405 | +- | 3.225 | 0.33892 | +- | 0.2337 | 1.877 | +- | 0.9873
[6 empty] | (ppb or ppm) | 71.38 ppb | +- | 80.18 ppb | [6 empty] | 0.46 ppb | +- | 0.24 ppb
SNOLAB V26 | Lake in City of Greater Sudbury, Ontario | 985.3 g | 10.743 | 201023 | 2020-10-23 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
St. Charles Lake Water | Sample collected on 2020-09-19 at 14:53 |  |  |  | 2020-11-03 | (mBq/kg) | 1.5 | +- | 0.62 | 27.39 | +- | 23.77 | 0.279 | +- | 0.2828 | 0.3211 | +- | 0.6553 | 56.582 | +- | 9.839 | 1.2519 | +- | 0.4744 | 0.1718 | +- | 0.2352
[6 empty] | (ppb or ppm) | 0.12 ppb | +- | 0.05 ppb | 2.22 ppb | +- | 1.93 ppb | 0.49 ppb | +- | 0.5 ppb | 0.08 ppb | +- | 0.16 ppb | 1.83 ppm | +- | 0.32 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 23258 | +- | 29210 | 13.734 | +- | 3.761 | <0.24 |  |  | 0.159 | +- | 0.9299
[6 empty] | (ppb or ppm) | 1883.9 ppb | +- | 2366.01 ppb | [6 empty] | 0.04 ppb | +- | 0.23 ppb
SNOLAB V27 | Lake in City of Greater Sudbury, Ontario | 956.4 g | 11.146 | 201106 | 2020-11-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Long Lake Water | Sample collected on 2020-09-27 at 15:00 |  |  |  | 2020-11-17 | (mBq/kg) | <0.98 |  |  | <16.85 |  |  | 0.4023 | +- | 0.2632 | <0.63 |  |  | 35.529 | +- | 8.902 | 0.61398 | +- | 0.4425 | <0.31
[6 empty] | (ppb or ppm) | <79.38 ppt |  |  | <1.36 ppb |  |  | 0.71 ppb | +- | 0.46 ppb | <0.15 ppb |  |  | 1147.59 ppb | +- | 287.53 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1542.4 | +- | 1182 | 14.09 | +- | 3.542 | <0.39 |  |  | <0.84
[6 empty] | (ppb or ppm) | 124.93 ppb | +- | 95.74 ppb | [6 empty] | <0.21 ppb
SNOLAB V28 | Lake Bed Sand from Lohi Lake in City of Greater Sudbury, Ontario | 600.4 g | 2.93 | 201020 | 2020-10-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Lohi Lake Sand | Sample collected on 2020-09-19 at 17:10 |  |  |  | 2020-10-23 | (mBq/kg) | 7917 | +- | 172.9 | 17980 | +- | 1656 | 422.1 | +- | 18.84 | 6592 | +- | 179.8 | 381970 | +- | 19550 | N/D – No observable signal due to U & Th Interference | +- | 508.4 | <27.73
[6 empty] | (ppb or ppm) | 641.28 ppb | +- | 14 ppb | 1456.38 ppb | +- | 134.14 ppb | 742.9 ppb | +- | 33.16 ppb | 1621.63 ppb | +- | 44.23 ppb | 12337.63 ppm | +- | 631.47 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 455740 | +- | 198800 | N/D – No observable signal due to U & Th Interference | +- | 710.5 | N/D – No observable signal due to U & Th Interference | +- | 18.84 | 6194 | +- | 190.9
[6 empty] | (ppb or ppm) | 36914.94 ppb | +- | 16102.8 ppb | [6 empty] | 1523.72 ppb | +- | 46.96 ppb
SNOLAB V29 | Lake Bed Sand from Ramsey Lake in City of Greater Sudbury, Ontario | 605.1 g | 2.93 | 201103 | 2020-11-03 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Ramsey Lake Sand | Sample collected on 2020-09-15 at 18:50 |  |  |  | 2020-11-06 | (mBq/kg) | 14950 | +- | 305.7 | 28000 | +- | 2515 | 765.4 | +- | 27.2 | 16230 | +- | 407.2 | 513760 | +- | 26280 | N/D – No observable signal due to U & Th Interference |  |  | <34.32
[6 empty] | (ppb or ppm) | 1210.95 ppb | +- | 24.76 ppb | 2268 ppb | +- | 203.72 ppb | 1347.1 ppb | +- | 47.87 ppb | 3992.58 ppb | +- | 100.17 ppb | 16594.45 ppm | +- | 848.84 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 156490 | +- | 184600 | N/D – No observable signal due to U & Th Interference |  |  | N/D – No observable signal due to U & Th Interference | +- | 24.88 | 14800 | +- | 392.3
[6 empty] | (ppb or ppm) | 12675.69 ppb | +- | 14952.6 ppb | [6 empty] | 3640.8 ppb | +- | 96.51 ppb
SNOLAB V30 | Lake Bed Sand from St Charles Lake in City of Greater Sudbury, Ontario | 591.6 g | 2.534 | 201117 | 2020-11-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
St Charles Lake Sand | Sample collected on 2020-09-19 at 14:55 |  |  |  | 2020-11-20 | (mBq/kg) | 12370 | +- | 263.6 | 29950 | +- | 2642 | 606.5 | +- | 25.82 | 16700 | +- | 422.7 | 572120 | +- | 29270 | N/D – No observable signal due to U & Th Interference |  |  | 17.15 | +- | 25.66
[6 empty] | (ppb or ppm) | 1001.97 ppb | +- | 21.35 ppb | 2425.95 ppb | +- | 214 ppb | 1067.44 ppb | +- | 45.44 ppb | 4108.2 ppb | +- | 103.98 ppb | 18479.48 ppm | +- | 945.42 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <266100.00 |  |  | N/D – No observable signal due to U & Th Interference |  |  | N/D – No observable signal due to U & Th Interference | +- | 30.21 | 16120 | +- | 429.6
[6 empty] | (ppb or ppm) | <21.55 ppm | [8 empty] | 3965.52 ppb | +- | 105.68 ppb
SNOLAB V31 | Lake in City of Greater Sudbury, Ontario | 1023.10 g | 13.687 | 201120 | 2020-11-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Lake Water | Sample collected on 2020-10-23 at 13:00 |  |  |  | 2020-12-04 | (mBq/kg) | 2.203 | +- | 0.5796 | 15.37 | +- | 20.24 | <0.42 |  |  | 0.9665 | +- | 0.6031 | 63.512 | +- | 9.205 | <0.38 |  |  | <0.22
[6 empty] | (ppb or ppm) | 178.44 ppt | +- | 46.95 ppt | 1.24 ppb | +- | 1.64 ppb | <0.74 ppb |  |  | 0.24 ppb | +- | 0.15 ppb | 2051.44 ppb | +- | 297.32 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <1381.00 |  |  | 57.033 | +- | 5.721 | <0.24 |  |  | 3.169 | +- | 0.9504
[6 empty] | (ppb or ppm) | <111.86 ppb | [8 empty] | 0.78 ppb | +- | 0.23 ppb
SNOLAB V32 | Lake in City of Greater Sudbury, Ontario | 961.0 g | 13.777 | 201204 | 2020-12-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Meatbird Lake Water | Sample collected on October 05, 2020 @ 14:00 |  |  |  | 2020-12-18 | (mBq/kg) | 0.8306 | +- | 0.5845 | <24.71 |  |  | <0.30 |  |  | 0.064 | +- | 0.5975 | 167.44 | +- | 15.2 | 0.18371 | +- | 0.3529 | 0.2729 | +- | 0.2459
[6 empty] | (ppb or ppm) | 67.28 ppt | +- | 47.34 ppt | <2 ppb |  |  | <0.53 ppb |  |  | 0.02 ppb | +- | 0.15 ppb | 5.41 ppm | +- | 0.49 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 977.51 | +- | 1067 | 5.128 | +- | 2.804 | 0.29759 | +- | 0.2745 | 4.56 | +- | 1.18
[6 empty] | (ppb or ppm) | 79.18 ppb | +- | 86.43 ppb | [6 empty] | 1.12 ppb | +- | 0.29 ppb
SNOLAB V33 | Lucas Cell Coating | 68.3 g | 23.652 | 201218 | 2020-12-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
ZnS(Ag) Powder | Saint-Gobain |  |  | The sample is also counted in the Well detector as SNOLAB CW10 | 2021-01-11 | (mBq/kg) | <3.30 |  |  | <98.38 |  |  | 0.8626 | +- | 0.9022 | <4.37 |  |  | 21.744 | +- | 34.92 | <3.09 |  |  | <0.89
[6 empty] | (ppb or ppm) | <0.27 ppb |  |  | <7.97 ppb |  |  | 1.52 ppb | +- | 1.59 ppb | <1.08 ppb |  |  | 702.33 ppb | +- | 1127.92 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 53891 | +- | 34430 | 5.0645 | +- | 11.94 | <1.33 |  |  | <5.33
[6 empty] | (ppb or ppm) | 4.37 ppm | +- | 2.79 ppm | [6 empty] | <1.31 ppb
SNOLAB V34 | 72% Cocoa | 202.9 g | 4.98 | 240425 | 2024-04-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Merci (Stork) Dark Chocolate | Made in Germany |  |  |  | 2024-04-30 | (mBq/kg) | 453.3 | +- | 23.5 | 952.5 | +- | 379.5 | 5.613 | +- | 6.602 | 285.7 | +- | 21.85 | 20507 | +- | 10540 | 62.385 | +- | 14.87 | <19.40
[6 empty] | (ppb or ppm) | 36.72 ppb | +- | 1.9 ppb | 77.15 ppb | +- | 30.74 ppb | 9.88 ppb | +- | 11.62 ppb | 70.28 ppb | +- | 5.38 ppb | 662.38 ppm | +- | 340.44 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <29620.00 |  |  | 227.13 | +- | 95.34 | 22.85 | +- | 14.15 | 630.1 | +- | 51.65
[6 empty] | (ppb or ppm) | <2.4 ppm | [8 empty] | 155 ppb | +- | 12.71 ppb
PICO Measurements:
Sample Description |  | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
PICO V01 | Model No. S1328PD3T72 Omega | 60.1 g | 15.659 | 200825 | 2020-08-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
White RTD Wire Leads |  |  |  |  | 2020-09-10 | (mBq/kg) | 5.375 | +- | 7.215 | <127.80 |  |  | 2.542 | +- | 2.395 | <9.16 |  |  | 69.281 | +- | 72.04 | <5.89 |  |  | <1.78
[6 empty] | (ppb or ppm) | 0.44 ppb | +- | 0.58 ppb | <10.35 ppb |  |  | 4.47 ppb | +- | 4.22 ppb | <2.25 ppb |  |  | 2237.78 ppb | +- | 2326.89 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 140700 | +- | 38520 | <53.17 |  |  | <1.90 |  |  | <14.98
[6 empty] | (ppb or ppm) | 11.4 ppm | +- | 3.12 ppm | [6 empty] | <3.69 ppb
PICO V02 | Model No. S1328PD3T72 Omega | 23.9 g | 11.639 | 200910 | 2020-09-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Red RTD Wire Leads |  |  |  |  | 2020-09-22 | (mBq/kg) | 4.669 | +- | 20.77 | <422.90 |  |  | 6.185 | +- | 6.586 | <27.98 |  |  | 74.988 | +- | 232.8 | <17.53 |  |  | <3.64
[6 empty] | (ppb or ppm) | 0.38 ppb | +- | 1.68 ppb | <34.25 ppb |  |  | 10.89 ppb | +- | 11.59 ppb | <6.88 ppb |  |  | 2.42 ppm | +- | 7.52 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 148160 | +- | 69770 | 132.14 | +- | 84.27 | <14.75 |  |  | <42.97
[6 empty] | (ppb or ppm) | 12 ppm | +- | 5.65 ppm | [6 empty] | <10.57 ppb
PICO V03 | Hard Copper Wire | 45.2 g | 13.687 | 200922 | 2020-09-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
PICO Heater Cables |  |  |  |  | 2020-10-06 | (mBq/kg) | <13.25 |  |  | <127.70 |  |  | <3.89 |  |  | 6.695 | +- | 9.819 | <148.60 |  |  | 9.1666 | +- | 6.88 | <5.33
[6 empty] | (ppb or ppm) | <1.07 ppb |  |  | <10.34 ppb |  |  | <6.85 ppb |  |  | 1.65 ppb | +- | 2.42 ppb | <4.8 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 58Co |  |  | 57Co
[6 empty] | (mBq/kg) | 143110 | +- | 62600 | 3.3293 | +- | 41.87 | 2.5467 | +- | 4.666 | 21.23 | +- | 16.07 | 0.54315 | +- | 3.654 | 7.049 | +- | 6.27
[6 empty] | (ppb or ppm) | 11.59 ppm | +- | 5.07 ppm | [6 empty] | 5.22 ppb | +- | 3.95 ppb
PICO V04 | Preamps for Piezos for the PICO and SBC experiments | 21.2 g | 25.61 | 210421 210429 | 2021-04-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Kapton Preamp Boards |  | Quantity: 28 |  |  | 2021-05-17 | (mBq/kg) | 1504 | +- | 45.52 | 4544 | +- | 459.1 | 89.11 | +- | 7.122 | 2226 | +- | 71.56 | 4239.8 | +- | 321.9 | <9.34 |  |  | <5.98
[6 empty] | (ppb or ppm) | 121.82 ppb | +- | 3.69 ppb | 368.06 ppb | +- | 37.19 ppb | 156.83 ppb | +- | 12.53 ppb | 547.6 ppb | +- | 17.6 ppb | 136.95 ppm | +- | 10.4 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 259030 | +- | 88590 | <108.00 |  |  | 12.595 | +- | 9.45 | 2254 | +- | 80.58
[6 empty] | (ppb or ppm) | 20.98 ppm | +- | 7.18 ppm | [6 empty] | 554.48 ppb | +- | 19.82 ppb
PICO V05 | BLU-25751706 Lot 230328 3/32 “ x 7” | 276.2 g | 5.088 | 231026 23102601 231028 231029 | 2023-10-26 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  | 231030 | 2023-10-31 | (mBq/kg) | <186.40 |  |  | 9577 | +- | 4406 | <193.20 |  |  | 570400 | +- | 13220 | 14541 | +- | 902.7 | 61.79 | +- | 71.9 | <50.83
| Sample contains 20 electrodes |  |  |  |  | (ppb or ppm) | <15.1 ppb |  |  | 775.74 ppb | +- | 356.89 ppb | <340.03 ppb |  |  | 140.32 ppm | +- | 3.25 ppm | 469.67 ppm | +- | 29.16 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <2.62E6 |  |  | 2903 | +- | 633.1 | 2985.6 | +- | 182.1 | 207300 | +- | 4975
[6 empty] | (ppb or ppm) | <212.22 ppm | [8 empty] | 51 ppm | +- | 1.22 ppm
PICO V06 | BLU-25751706 Lot 220713 3/32 “ x 7” | 422.2 g | 4.687 | 231031 23103101 231101 231102 231103 | 2023-10-31 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  |  | 2023-11-05 | (mBq/kg) | <68.47 |  |  | 927.8 | +- | 1908 | 64.68 | +- | 53.28 | 93870 | +- | 2196 | 7938.6 | +- | 475.8 | <66.87 |  |  | <10.82
| Sample contains 30 electrodes |  |  |  |  | (ppb or ppm) | <5.55 ppb |  |  | 75.15 ppb | +- | 154.55 ppb | 113.84 ppb | +- | 93.77 ppb | 23.09 ppm | +- | 0.54 ppm | 256.42 ppm | +- | 15.37 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2324000 | +- | 2368000 | 1866.6 | +- | 317.7 | 1527 | +- | 84.04 | 105300 | +- | 2557
[6 empty] | (ppb or ppm) | 188.24 ppm | +- | 191.81 ppm | [6 empty] | 25.9 ppm | +- | 0.63 ppm
PICO V07 | BLU-25751708 Lot 230204 1/8 “ x 7” | 251.5 g | 3.126 | 231105 | 2023-11-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  |  | 2023-11-08 | (mBq/kg) | <115.70 |  |  | <7092.00 |  |  | <235.50 |  |  | 577200 | +- | 13410 | 16792 | +- | 1114 | <111.70 |  |  | <59.50
| Sample contains 10 electrodes |  |  |  |  | (ppb or ppm) | <9.37 ppb |  |  | <574.45 ppb |  |  | <414.48 ppb |  |  | 141.99 ppm | +- | 3.3 ppm | 542.38 ppm | +- | 35.98 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <7.92E6 |  |  | 3190.9 | +- | 949.6 | 3172.3 | +- | 181.6 | 222400 | +- | 5483
[6 empty] | (ppb or ppm) | <641.52 ppm | [8 empty] | 54.71 ppm | +- | 1.35 ppm
PICO V08 | BLU-25751708 Lot 230515 1/8 “ x 7” | 1004.4 g | 7.786 | 231108 | 2023-11-08 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  |  | 2023-11-16 | (mBq/kg) | <33.37 |  |  | 263.4 | +- | 2273 | 299.5 | +- | 113.6 | 599300 | +- | 14060 | 17403 | +- | 929.6 | <59.25 |  |  | <14.68
| Sample contains 40 electrodes |  |  |  |  | (ppb or ppm) | <2.7 ppb |  |  | 21.34 ppb | +- | 184.11 ppb | 527.12 ppb | +- | 199.94 ppb | 147.43 ppm | +- | 3.46 ppm | 562.12 ppm | +- | 30.03 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <1.79E7 |  |  | 3007.2 | +- | 395 | 2659.2 | +- | 146.1 | 213800 | +- | 5309
[6 empty] | (ppb or ppm) | <1449.9 ppm | [8 empty] | 52.59 ppm | +- | 1.31 ppm
PICO V09 | Molex wire ring connectors | 100.6 g | 8.757 | 240305 240313 | 2024-03-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Wire Ring connectors | Types: 16-14 MX 22-18 MX |  |  |  | 2024-03-15 | (mBq/kg) | <4.41 |  |  | 99.35 | +- | 249 | <4.27 |  |  | <5.03 |  |  | <64.56 |  |  | <3.73 |  |  | <3.75
[6 empty] | (ppb or ppm) | <357.21 ppt |  |  | 8.05 ppb | +- | 20.17 ppb | <7.52 ppb |  |  | <1.24 ppb |  |  | <2.09 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 371410 | +- | 337700 | 46.919 | +- | 30.78 | 7.5284 | +- | 3.965 | <7.22
[6 empty] | (ppb or ppm) | 30.08 ppm | +- | 27.35 ppm | [6 empty] | <1.78 ppb
NEXO Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
NEXO V01 | TIVAR UMHW | 1065.6 g | 19.495 | 231122 231209 | 2023-11-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
TIVAR UHMW EC J-202 | Cuboid Pieces |  |  |  | 2023-12-12 | (mBq/kg) | <0.62 |  |  | 21.02 | +- | 12.66 | <0.41 |  |  | <1.12 |  |  | 49.561 | +- | 8.095 | <0.43 |  |  | <0.14
[6 empty] | (ppb or ppm) | <50.22 ppt |  |  | 1.7 ppb | +- | 1.03 ppb | <0.72 ppb |  |  | <0.28 ppb |  |  | 1.6 ppm | +- | 0.26 ppm
| 4 Units |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 210Po
[6 empty] | (mBq/kg) | 3108 | +- | 1928 | 3.3424 | +- | 2.339 | 0.28084 | +- | 0.2337 | <1.71
[6 empty] | (ppb or ppm) | 251.75 ppb | +- | 156.17 ppb | [6 empty] | <0.42 ppb
NEXO V02 | TIVAR ESD | 2478.77 g | 15.853 | 231212 231219 | 2023-12-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
TIVAR ESD J-203 | Cylinders |  |  |  | 2023-12-28 | (mBq/kg) | 0.5197 | +- | 0.1785 | <8.46 |  |  | 0.08824 | +- | 0.07382 | 0.2078 | +- | 0.1823 | 5.9626 | +- | 2.111 | <0.065 |  |  | <0.12
[6 empty] | (ppb or ppm) | 42.1 ppt | +- | 14.46 ppt | <0.69 ppb |  |  | 0.16 ppb | +- | 0.13 ppb | 51.12 ppt | +- | 44.85 ppt | 192.59 ppb | +- | 68.19 ppb
| 27 units counted |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1835.7 | +- | 816.5 | 1.2325 | +- | 0.7827 | <0.061 |  |  | <0.25
[6 empty] | (ppb or ppm) | 148.69 ppb | +- | 66.14 ppb | [6 empty] | <61.5 ppt
NEXO V03 | Semi-Gloss – White Oil Based | 73.3 g | 9.869 | 240216 | 2024-02-16 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tremclad Rust Paint | Part No: 254933 |  |  |  | 2024-02-26 | (mBq/kg) | 737.5 | +- | 26.44 | 2487 | +- | 276.2 | 35.97 | +- | 4.716 | 619.4 | +- | 27.16 | 3033.1 | +- | 238.9 | 14.649 | +- | 7.056 | <4.29
| Rust-Oleum, Concord, ON |  |  |  |  | (ppb or ppm) | 59.74 ppb | +- | 2.14 ppb | 201.45 ppb | +- | 22.37 ppb | 63.31 ppb | +- | 8.3 ppb | 152.37 ppb | +- | 6.68 ppb | 97.97 ppm | +- | 7.72 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 210Po
[6 empty] | (mBq/kg) | <19550.00 |  |  | <75.74 |  |  | 18.179 | +- | 6.3 | 592.4 | +- | 35.18
[6 empty] | (ppb or ppm) | <1.58 ppm | [8 empty] | 145.73 ppb | +- | 8.65 ppb
CUTE Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
CUTE V01 | CUTE Source SRS-19-003 | 194.0 g | 6.791 | 210120 | 2021-01-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
CUTE Fe-55 Calibration Sample | 55-Fe source inside a copper holder |  |  |  | 2021-01-27 | (mBq/kg) | <1.44 |  |  | <247.60 |  |  | 6.192 | +- | 3.943 | 7.232 | +- | 2.439 | 7.669 | +- | 20.01 | <1.56 |  |  | 175.3 | +- | 8.401
| The source strength is 222 kBq |  |  |  |  | (ppb or ppm) | <0.12 ppb |  |  | <20.06 ppb |  |  | 10.9 ppb | +- | 6.94 ppb | 1.78 ppb | +- | 0.6 ppb | 247.71 ppb | +- | 646.32 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | <639400.00 | High limit due to large number of 55-Fe X-rays |  | 5.8584 | +- | 9.093 | <2.18 |  |  | 4.19 | +- | 4.472 | <53.66 |  |  | <1.79 |  |  | + large number of X-Rays from 55-Fe
[6 empty] | (ppb or ppm) | <51.79 ppm | [8 empty] | 1.03 ppb | +- | 1.1 ppb
CUTE V02 | Loctite Epoxy by Henkel | 52.3 g | 5.861 | 210415 | 2021-04-15 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Loctite EA-0151 Epoxy | PN: 398454 |  |  |  | 2021-04-21 | (mBq/kg) | <8.65 |  |  | <64.10 |  |  | <3.01 |  |  | <9.01 |  |  | <78.76 |  |  | <4.56 |  |  | <6.37
| Batch No: LB0KAA3587 21/043 |  |  |  |  | (ppb or ppm) | <0.7 ppb |  |  | <5.19 ppb |  |  | <5.3 ppb |  |  | <2.22 ppb |  |  | <2.54 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 897.72 | +- | 981.5 | <74.36 |  |  | <6.39 |  |  | <27.55
[6 empty] | (ppb or ppm) | 72.72 ppb | +- | 79.5 ppb | [6 empty] | <6.78 ppb
CUTE V03 |  | 51.6 g | 7.883 | 240411 | 2024-04-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
G10 Plastic Tubes & Inserts |  |  |  |  | 2024-04-19 | (mBq/kg) | 8100 | +- | 184.2 | 25000 | +- | 1808 | 397.2 | +- | 19.02 | 20790 | +- | 512.1 | 16364 | +- | 950 | <26.87 |  |  | <14.42
[6 empty] | (ppb or ppm) | 656.1 ppb | +- | 14.92 ppb | 2.03 ppm | +- | 0.15 ppm | 699.07 ppb | +- | 33.48 ppb | 5.11 ppm | +- | 0.13 ppm | 528.56 ppm | +- | 30.69 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 30602 | +- | 83100 | 354.71 | +- | 202 | 236.28 | +- | 30.01 | 18730 | +- | 476.9
[6 empty] | (ppb or ppm) | 2.48 ppm | +- | 6.73 ppm | [6 empty] | 4.61 ppm | +- | 0.12 ppm
CUTE V04 |  | 16.9 g | 5.778 | 240419 240425 | 2024-04-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
G10 Plastic Tubes |  |  |  |  | 2024-04-25 | (mBq/kg) | 10710 | +- | 323.5 | 39930 | +- | 3861 | 527.3 | +- | 51.78 | 27360 | +- | 793.2 | 22372 | +- | 1778 | <86.57 |  |  | <78.20
[6 empty] | (ppb or ppm) | 867.51 ppb | +- | 26.2 ppb | 3.23 ppm | +- | 0.31 ppm | 928.05 ppb | +- | 91.13 ppb | 6.73 ppm | +- | 0.2 ppm | 722.62 ppm | +- | 57.43 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 100830 | +- | 253700 | <1163.00 |  |  | 415.95 | +- | 80.19 | 25420 | +- | 812.2
[6 empty] | (ppb or ppm) | 8.17 ppm | +- | 20.55 ppm | [6 empty] | 6.25 ppm | +- | 0.2 ppm
SBC Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
SBC V01 | Stainless Steel | 364.72 g | 8.777 | 210127 210202 | 2021-01-27 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Vessel Body |  |  |  | 2021-02-05 | (mBq/kg) | 1.909 | +- | 1.127 | <61.70 |  |  | 0.4688 | +- | 0.4291 | <2.64 |  |  | <23.78 |  |  | 0.70548 | +- | 0.6779 | 5.059 | +- | 0.8153
| PV Drawing Item #3 |  |  |  |  | (ppb or ppm) | 154.63 ppt | +- | 91.29 ppt | <5 ppb |  |  | 0.83 ppb | +- | 0.76 ppb | <0.65 ppb |  |  | <0.77 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 148260 | +- | 118900 | <7.53 |  |  | 2.8376 | +- | 0.7004 | <1.76 |  |  | 4.727 | +- | 1.904 | <0.86
[6 empty] | (ppb or ppm) | 12.01 ppm | +- | 9.63 ppm | [6 empty] | <0.43 ppb
SBC V02 | Stainless Steel | 254.35 g | 6.791 | 210205 200207 | 2021-02-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Feedthrough Tube for the Pressure Vessel |  |  |  | 2021-02-12 | (mBq/kg) | 3.86 | +- | 2.028 | <67.98 |  |  | 0.3404 | +- | 0.7203 | 4.659 | +- | 2.153 | <34.73 |  |  | 1.0881 | +- | 1.447 | 1.016 | +- | 0.8182
| PV Drawing Item #8 |  |  |  |  | (ppb or ppm) | 312.66 ppt | +- | 164.27 ppt | <5.51 ppb |  |  | 0.6 ppb | +- | 1.27 ppb | 1.15 ppb | +- | 0.53 ppb | <1.12 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 141470 | +- | 129400 | 7.4193 | +- | 8.756 | 2.3993 | +- | 1.235 | <7.34 |  |  | <2.44 |  |  | <1.17
[6 empty] | (ppb or ppm) | 11.46 ppm | +- | 10.48 ppm | [6 empty] | <1.81 ppb
SBC V03 | Stainless Steel | 189.45 g | 11.541 | 210212 210216 | 2021-02-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Vessel Spool |  |  |  | 2021-02-24 | (mBq/kg) | 0.9399 | +- | 1.763 | 126.8 | +- | 71.45 | 2.376 | +- | 0.7457 | <3.89 |  |  | <32.52 |  |  | 1.5692 | +- | 1.16 | 75.03 | +- | 4.169
| PV Drawing Item #6 |  |  |  |  | (ppb or ppm) | 0.08 ppb | +- | 0.14 ppb | 10.27 ppb | +- | 5.79 ppb | 4.18 ppb | +- | 1.31 ppb | <0.96 ppb |  |  | <1.05 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 122400 | +- | 106900 | <12.46 |  |  | 6.4552 | +- | 1.469 | 2.611 | +- | 3.374 | <3.72 |  |  | <2.38
[6 empty] | (ppb or ppm) | 9.91 ppm | +- | 8.66 ppm | [6 empty] | 0.64 ppb | +- | 0.83 ppb
SBC V04 | Stainless Steel | 134.65 g | 8.639 | 210224 210301 | 2021-02-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Vessel Dome |  |  |  | 2021-03-05 | (mBq/kg) | <3.93 |  |  | <108.50 |  |  | <0.74 |  |  | 4.873 | +- | 2.867 | <42.76 |  |  | <1.27 |  |  | <1.33
| PV Drawing Item #1 |  |  |  |  | (ppb or ppm) | <0.32 ppb |  |  | <8.79 ppb |  |  | <1.3 ppb |  |  | 1.2 ppb | +- | 0.71 ppb | <1.38 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 74615 | +- | 120200 | 3.12 | +- | 8.982 | 4.1339 | +- | 1.37 | 12.3 | +- | 4.232 | 3.749 | +- | 2.117 | <1.59
[6 empty] | (ppb or ppm) | 6.04 ppm | +- | 9.74 ppm | [6 empty] | 3.03 ppb | +- | 1.04 ppb
SBC V05 | Stainless Steel | 372.40 g | 14.722 | 210305 | 2021-03-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Relief Pipe |  |  |  | 2021-03-22 | (mBq/kg) | 3.971 | +- | 1.466 | 82.29 | +- | 53.49 | 0.8206 | +- | 0.6175 | 6.322 | +- | 1.587 | <5.66 |  |  | <0.80 |  |  | 3.23 | +- | 0.8053
[6 empty] | (ppb or ppm) | 0.32 ppb | +- | 0.12 ppb | 6.67 ppb | +- | 4.33 ppb | 1.44 ppb | +- | 1.09 ppb | 1.56 ppb | +- | 0.39 ppb | <0.18 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 130530 | +- | 136300 | <6.77 |  |  | 1.6401 | +- | 0.7656 | <2.38 |  |  | <2.28 |  |  | 0.54838 | +- | 0.6217
[6 empty] | (ppb or ppm) | 10.57 ppm | +- | 11.04 ppm | [6 empty] | <0.59 ppb
SBC V06 | Silver Plated Bolts | 744.9 g | 7.182 | 231228 | 2023-12-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Silter Plated Bolts | Label: S30400 THE |  |  |  | 2024-01-05 | (mBq/kg) | <1.40 |  |  | 26.97 | +- | 36.14 | 0.4047 | +- | 0.3929 | 3.417 | +- | 1.103 | <9.62 |  |  | <0.50 |  |  | 23.25 | +- | 1.558
| Units: 6 |  |  |  |  | (ppb or ppm) | <0.11 ppb |  |  | 2.18 ppb | +- | 2.93 ppb | 0.71 ppb | +- | 0.69 ppb | 0.84 ppb | +- | 0.27 ppb | <0.31 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 18039 | +- | 20200 | 10.836 | +- | 5.093 | 0.8271 | +- | 0.6252 | 2.577 | +- | 1.833
[6 empty] | (ppb or ppm) | 1.46 ppm | +- | 1.64 ppm | [6 empty] | 0.63 ppb | +- | 0.45 ppb
SBC V07 | PTFE | 48.5 g | 11.708 | 240105 | 2024-01-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
IA PTFE inner Tower |  |  |  |  | 2024-01-17 | (mBq/kg) | <4.70 |  |  | 181.1 | +- | 111.8 | <4.56 |  |  | <5.05 |  |  | <63.58 |  |  | <11.23 |  |  | <2.16
[6 empty] | (ppb or ppm) | <0.38 ppb |  |  | 14.67 ppb | +- | 9.06 ppb | <8.03 ppb |  |  | <1.24 ppb |  |  | <2.05 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 23168 | +- | 13680 | <40.04 |  |  | <8.61 |  |  | <9.43
[6 empty] | (ppb or ppm) | 1.88 ppm | +- | 1.11 ppm | [6 empty] | <2.32 ppb
SBC V08 |  | 1695.2 g | 7.854 | 240117 | 2024-01-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Bellows Flange | 8 Pieces |  |  |  | 2024-01-25 | (mBq/kg) | 4.943 | +- | 0.8595 | 15.12 | +- | 27.78 | 0.1853 | +- | 0.3833 | 12.42 | +- | 1.2 | <7.91 |  |  | <0.59 |  |  | 10.29 | +- | 0.7999
[6 empty] | (ppb or ppm) | 0.4 ppb | +- | 0.07 ppb | 1.22 ppb | +- | 2.25 ppb | 0.33 ppb | +- | 0.67 ppb | 3.06 ppb | +- | 0.3 ppb | <0.26 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 14818 | +- | 21610 | 5.0672 | +- | 3.814 | 0.83976 | +- | 0.4803 | 15.84 | +- | 1.965 | <1.30 |  |  | 0.97213 | +- | 0.4483
[6 empty] | (ppb or ppm) | 1.2 ppm | +- | 1.75 ppm | [6 empty] | 3.9 ppb | +- | 0.48 ppb
SBC V09 | Standard Stainles Steel Tubes which are welded together | 150.8 g | 6.73 | 240209 240210 24021001 | 2024-02-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Welded Tubes |  |  |  |  | 2024-02-16 | (mBq/kg) | <1.61 |  |  | <73.78 |  |  | <0.93 |  |  | 4.495 | +- | 2.51 | <21.85 |  |  | <1.57 |  |  | 6.462 | +- | 1.508
[6 empty] | (ppb or ppm) | <0.13 ppb |  |  | <5.98 ppb |  |  | <1.64 ppb |  |  | 1.11 ppb | +- | 0.62 ppb | <705.76 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | <14390.00 |  |  | <12.37 |  |  | 2.8532 | +- | 1.386 | 2.074 | +- | 3.812 | <1.65 |  |  | <2.51
[6 empty] | (ppb or ppm) | <1.17 ppm | [8 empty] | 0.51 ppb | +- | 0.94 ppb
DEAP Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
DEAP V01 | Carlisle Brushes | 151.7 g | 8.819 | 210322 | 2021-03-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Carlisle Brushes | Part No: 4039002 | Quantity: 5 |  |  | 2021-03-31 | (mBq/kg) | 90.62 | +- | 6.712 | 237.9 | +- | 99.48 | 5.367 | +- | 1.655 | 16.09 | +- | 4.774 | 648.34 | +- | 82.23 | <1.48 |  |  | <2.27
[6 empty] | (ppb or ppm) | 7.34 ppb | +- | 0.54 ppb | 19.27 ppb | +- | 8.06 ppb | 9.45 ppb | +- | 2.91 ppb | 3.96 ppb | +- | 1.17 ppb | 20.94 ppm | +- | 2.66 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 31077 | +- | 23210 | <20.17 |  |  | 2.5718 | +- | 2.074 | 31.42 | +- | 8.244
[6 empty] | (ppb or ppm) | 2.52 ppm | +- | 1.88 ppm | [6 empty] | 7.73 ppb | +- | 2.03 ppb
DEAP V02 | Gordon Brushes | 207.0 g | 14.715 | 210331 210405 | 2021-03-31 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Gordon Brushes | Part No: 906389 | Quantity: 5 |  |  | 2021-04-15 | (mBq/kg) | 0.4271 | +- | 1.656 | <67.28 |  |  | <1.63 |  |  | 10.54 | +- | 2.125 | <24.83 |  |  | <1.70 |  |  | 1.944 | +- | 0.8329
[6 empty] | (ppb or ppm) | 0.03 ppb | +- | 0.13 ppb | <5.45 ppb |  |  | <2.87 ppb |  |  | 2.59 ppb | +- | 0.52 ppb | <0.8 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 53660 | +- | 41610 | <13.76 |  |  | 1.8481 | +- | 0.8673 | 10.81 | +- | 3.288
[6 empty] | (ppb or ppm) | 4.35 ppm | +- | 3.37 ppm | [6 empty] | 2.66 ppb | +- | 0.81 ppb
CNL Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
CNL V01 |  | 49.2919 g | 4.443 | 230202 | 2023-02-02 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Pembroke Site 3 10-15 cm |  |  |  |  | 2023-02-07 | (mBq/kg) | 10880 | +- | 279.7 | 39520 | +- | 2966 | 501.2 | +- | 33.79 | 13540 | +- | 396 | 669460 | +- | 34440 | 4596.9 | +- | 254.7 | <80.57
[6 empty] | (ppb or ppm) | 881.28 ppb | +- | 22.66 ppb | 3201.12 ppb | +- | 240.25 ppb | 882.11 ppb | +- | 59.47 ppb | 3330.84 ppb | +- | 97.42 ppb | 21623.56 ppm | +- | 1112.41 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 57367 | +- | 108600 | <724.90 |  |  | 202.97 | +- | 60.98 | 13100 | +- | 432.4
[6 empty] | (ppb or ppm) | 4.65 ppm | +- | 8.8 ppm | [6 empty] | 3222.6 ppb | +- | 106.37 ppb
CNL V02 |  | 41.7266 g | 3.002 | 230207 | 2023-02-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Pembroke Site 2 10-15 cm |  |  |  |  | 2023-02-10 | (mBq/kg) | 10760 | +- | 302.5 | 26190 | +- | 2540 | 598.7 | +- | 41.24 | 10610 | +- | 348.3 | 705960 | +- | 36490 | 6556 | +- | 366.6 | <127.10
[6 empty] | (ppb or ppm) | 871.56 ppb | +- | 24.5 ppb | 2121.39 ppb | +- | 205.74 ppb | 1053.71 ppb | +- | 72.58 ppb | 2610.06 ppb | +- | 85.68 ppb | 22802.51 ppm | +- | 1178.63 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 256440 | +- | 123600 | <999.10 |  |  | 93.624 | +- | 66.66 | 10820 | +- | 418.8
[6 empty] | (ppb or ppm) | 20.77 ppm | +- | 10.01 ppm | [6 empty] | 2661.72 ppb | +- | 103.02 ppb
CNL V03 |  | 48.4052 g | 2.841 | 230210 | 2023-02-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Pembroke Site 1 10-15 cm |  |  |  |  | 2023-02-13 | (mBq/kg) | 9014 | +- | 367.5 | 29300 | +- | 2834 | 450.9 | +- | 40.66 | 12620 | +- | 403.2 | 807110 | +- | 41600 | 4355.5 | +- | 257 | <51.11
[6 empty] | (ppb or ppm) | 730.13 ppb | +- | 29.77 ppb | 2373.3 ppb | +- | 229.55 ppb | 793.58 ppb | +- | 71.56 ppb | 3104.52 ppb | +- | 99.19 ppb | 26069.65 ppm | +- | 1343.68 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 128320 | +- | 137500 | <655.20 |  |  | 139.87 | +- | 76.46 | 12320 | +- | 461
[6 empty] | (ppb or ppm) | 10.39 ppm | +- | 11.14 ppm | [6 empty] | 3030.72 ppb | +- | 113.41 ppb
CNL V04 |  | 51.0620 g | 3.899 | 230213 | 2023-02-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Lemieux Island 10-15 cm |  |  |  |  | 2023-02-17 | (mBq/kg) | 12050 | +- | 309.5 | 45830 | +- | 3314 | 746.3 | +- | 40.25 | 19810 | +- | 552.8 | 641560 | +- | 30300 | 2681.6 | +- | 166.6 | <71.39
[6 empty] | (ppb or ppm) | 976.05 ppb | +- | 25.07 ppb | 3712.23 ppb | +- | 268.43 ppb | 1313.49 ppb | +- | 70.84 ppb | 4873.26 ppb | +- | 135.99 ppb | 20722.39 ppm | +- | 978.69 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 641560 | +- | 33030 | 715.55 | +- | 458 | 361.53 | +- | 166.6 | 18830 | +- | 576.3
[6 empty] | (ppb or ppm) | 51.97 ppm | +- | 2.68 ppm | [6 empty] | 4632.18 ppb | +- | 141.77 ppb
CNL V05 |  | 39.3869 g | 3.871 | 230217 | 2023-02-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-004 10-15 cm |  |  |  |  | 2023-02-21 | (mBq/kg) | 26070 | +- | 598.9 | 77410 | +- | 5281 | 1596 | +- | 64.14 | 34040 | +- | 912.8 | 590410 | +- | 30510 | 972.83 | +- | 110.3 | 43.7 | +- | 77.43
[6 empty] | (ppb or ppm) | 2111.67 ppb | +- | 48.51 ppb | 6270.21 ppb | +- | 427.76 ppb | 2808.96 ppb | +- | 112.89 ppb | 8373.84 ppb | +- | 224.55 ppb | 19070.24 ppm | +- | 985.47 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <291400.00 |  |  | 160.19 | +- | 600 | 705.75 | +- | 94.7 | 32860 | +- | 933.7
[6 empty] | (ppb or ppm) | <23.6 ppm | [8 empty] | 8083.56 ppb | +- | 229.69 ppb
CNL V06 |  | 43.9680 g | 2.957 | 230221 | 2023-02-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-003 10-15 cm |  |  |  |  | 2023-02-24 | (mBq/kg) | 19080 | +- | 487.5 | 56920 | +- | 4889 | 1207 | +- | 62.66 | 41210 | +- | 1090 | 622680 | +- | 32210 | 4544.2 | +- | 274.3 | <85.05
[6 empty] | (ppb or ppm) | 1545.48 ppb | +- | 39.49 ppb | 4610.52 ppb | +- | 396.01 ppb | 2124.32 ppb | +- | 110.28 ppb | 10137.66 ppb | +- | 268.14 ppb | 20112.56 ppm | +- | 1040.38 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 286110 | +- | 189600 | <1059.00 |  |  | 715.21 | +- | 105.4 | 40500 | +- | 1125
[6 empty] | (ppb or ppm) | 23.17 ppm | +- | 15.36 ppm | [6 empty] | 9963 ppb | +- | 276.75 ppb
CNL V07 |  | 39.8567 g | 3.766 | 230224 230228 | 2023-02-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-002 10-15 cm |  |  |  |  | 2023-02-28 | (mBq/kg) | 11660 | +- | 310 | 32310 | +- | 2711 | 611 | +- | 38.36 | 12530 | +- | 383.4 | 371710 | +- | 19330 | 5223.6 | +- | 293.2 | 64.58 | +- | 58.44
[6 empty] | (ppb or ppm) | 944.46 ppb | +- | 25.11 ppb | 2617.11 ppb | +- | 219.59 ppb | 1075.36 ppb | +- | 67.51 ppb | 3082.38 ppb | +- | 94.32 ppb | 12006.23 ppm | +- | 624.36 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 78726 | +- | 115200 | 231.36 | +- | 444.7 | 206.38 | +- | 62.46 | 12400 | +- | 430.4
[6 empty] | (ppb or ppm) | 6.38 ppm | +- | 9.33 ppm | [6 empty] | 3050.4 ppb | +- | 105.88 ppb
CNL V08 |  | 41.9112 g | 2.915 | 29022801 | 2023-02-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-005 10-15 cm |  |  |  |  | 2023-03-03 | (mBq/kg) | 23300 | +- | 564.1 | 64860 | +- | 4750 | 1431 | +- | 65.68 | 29120 | +- | 814.2 | 333920 | +- | 27700 | 7712.2 | +- | 427.6 | <167.10
[6 empty] | (ppb or ppm) | 1887.3 ppb | +- | 45.69 ppb | 5253.66 ppb | +- | 384.75 ppb | 2518.56 ppb | +- | 115.6 ppb | 7163.52 ppb | +- | 200.29 ppb | 10785.62 ppm | +- | 894.71 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 119830 | +- | 180100 | 1389.2 | +- | 687.1 | 294.54 | +- | 91.79 | 28480 | +- | 856.4
[6 empty] | (ppb or ppm) | 9706.23 ppb | +- | 14588.1 ppb | [6 empty] | 7006.08 ppb | +- | 210.67 ppb
CNL V09 |  | 41.9332 g | 3.853 | 230303 | 2023-03-03 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-006 10-15 cm |  |  |  |  | 2023-03-07 | (mBq/kg) | 25500 | +- | 581.2 | 68240 | +- | 4574 | 1435 | +- | 58.89 | 27010 | +- | 739.8 | 539400 | +- | 27910 | 1702.6 | +- | 131.9 | <93.04
[6 empty] | (ppb or ppm) | 2065.5 ppb | +- | 47.08 ppb | 5527.44 ppb | +- | 370.49 ppb | 2525.6 ppb | +- | 103.65 ppb | 6644.46 ppb | +- | 181.99 ppb | 17422.62 ppm | +- | 901.49 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 19616 | +- | 155700 | <1074.00 |  |  | 403.79 | +- | 82.89 | 26630 | +- | 780.6
[6 empty] | (ppb or ppm) | 1.59 ppm | +- | 12.61 ppm | [6 empty] | 6550.98 ppb | +- | 192.03 ppb
CNL V10 |  | 49.9195 g | 2.724 | 230307 | 2023-03-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Princess Sodalite Mine 10-15 cm |  |  |  |  | 2023-03-10 | (mBq/kg) | 22930 | +- | 552.9 | 63690 | +- | 4739 | 1049 | +- | 57.99 | 26130 | +- | 746.8 | 828250 | +- | 42690 | 62.474 | +- | 93.55 | <150.30
[6 empty] | (ppb or ppm) | 1857.33 ppb | +- | 44.78 ppb | 5158.89 ppb | +- | 383.86 ppb | 1846.24 ppb | +- | 102.06 ppb | 6427.98 ppb | +- | 183.71 ppb | 26752.48 ppm | +- | 1378.89 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <252500.00 |  |  | 1041 | +- | 655.6 | 377.08 | +- | 98.05 | 27410 | +- | 831
[6 empty] | (ppb or ppm) | <20.45 ppm | [8 empty] | 6742.86 ppb | +- | 204.43 ppb
CNL V11 |  | 43.4597 g | 6.3 | 230310 230314 | -619747 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Silver Crater Mine 10-15 cm |  |  |  |  | 2023-03-17 | (mBq/kg) | 99480 | +- | 1861 | 341400 | +- | 15640 | 8150 | +- | 186.3 | 45590 | +- | 1156 | 599190 | +- | 30810 | 5185.4 | +- | 287.6 | <169.50
[6 empty] | (ppb or ppm) | 8057.88 ppb | +- | 150.74 ppb | 27653.4 ppb | +- | 1266.84 ppb | 14344 ppb | +- | 327.89 ppb | 11215.14 ppb | +- | 284.38 ppb | 19353.84 ppm | +- | 995.16 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 349250 | +- | 196200 | 592.52 | +- | 637.6 | 730.47 | +- | 95.68 | 39990 | +- | 1103
[6 empty] | (ppb or ppm) | 28.29 ppm | +- | 15.89 ppm | [6 empty] | 9837.54 ppb | +- | 271.34 ppb
CNL V12 |  | 29.3726 g | 3.897 | 230317 | 2023-03-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Silver Crate Mine 5-10 cm |  |  |  |  | 2023-03-21 | (mBq/kg) | 164900 | +- | 3116 | 495300 | +- | 23130 | 11780 | +- | 279.8 | 52230 | +- | 1402 | 592680 | +- | 30740 | 9499.8 | +- | 527.6 | <237.60
[6 empty] | (ppb or ppm) | 13356.9 ppb | +- | 252.4 ppb | 40119.3 ppb | +- | 1873.53 ppb | 20732.8 ppb | +- | 492.45 ppb | 12848.58 ppb | +- | 344.89 ppb | 19143.56 ppm | +- | 992.9 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 117790 | +- | 306300 | <2320.00 |  |  | 849.02 | +- | 151.5 | 47310 | +- | 1409
[6 empty] | (ppb or ppm) | 9.54 ppm | +- | 24.81 ppm | [6 empty] | 11638.26 ppb | +- | 346.61 ppb
CNL V13 |  | 35.6399 g | 2.878 | 230321 | 2023-03-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-2 5-10 cm |  |  |  |  | 2023-03-24 | (mBq/kg) | 16170 | +- | 434.5 | 42940 | +- | 3753 | 733.1 | +- | 53.65 | 14460 | +- | 467.7 | 719350 | +- | 37270 | 5218.8 | +- | 313.4 | <188.70
[6 empty] | (ppb or ppm) | 1309.77 ppb | +- | 35.19 ppb | 3478.14 ppb | +- | 303.99 ppb | 1290.26 ppb | +- | 94.42 ppb | 3557.16 ppb | +- | 115.05 ppb | 23235.01 ppm | +- | 1203.82 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 287490 | +- | 175600 | <962.90 |  |  | 260.74 | +- | 95.67 | 13920 | +- | 543.2
[6 empty] | (ppb or ppm) | 23.29 ppm | +- | 14.22 ppm | [6 empty] | 3424.32 ppb | +- | 133.63 ppb
CNL V14 |  | 17.3276 g | 3.869 | 230324 | 2023-03-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-8 5-10 cm |  |  |  |  | 2023-03-28 | (mBq/kg) | 7913 | +- | 272.5 | 25650 | +- | 2906 | 485.3 | +- | 48.01 | 8285 | +- | 325.4 | 450450 | +- | 23810 | 16529 | +- | 889.6 | <136.30
[6 empty] | (ppb or ppm) | 640.95 ppb | +- | 22.07 ppb | 2077.65 ppb | +- | 235.39 ppb | 854.13 ppb | +- | 84.5 ppb | 2038.11 ppb | +- | 80.05 ppb | 14549.54 ppm | +- | 769.06 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 286810 | +- | 154000 | 1499.8 | +- | 717.5 | 96.874 | +- | 85.97 | 7417 | +- | 410.4
[6 empty] | (ppb or ppm) | 23.23 ppm | +- | 12.47 ppm | [6 empty] | 1824.58 ppb | +- | 100.96 ppb
CNL V15 |  | 40.0602 g | 2.916 | 230328 | 2023-03-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-5 5-10 cm |  |  |  |  | 2023-03-31 | (mBq/kg) | 16440 | +- | 425.1 | 35480 | +- | 3479 | 764.6 | +- | 50.13 | 10810 | +- | 366.7 | 704010 | +- | 36420 | 4833.4 | +- | 288.7 | <138.00
[6 empty] | (ppb or ppm) | 1331.64 ppb | +- | 34.43 ppb | 2873.88 ppb | +- | 281.8 ppb | 1345.7 ppb | +- | 88.23 ppb | 2659.26 ppb | +- | 90.21 ppb | 22739.52 ppm | +- | 1176.37 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 175180 | +- | 155600 | 716.91 | +- | 608.1 | 130.04 | +- | 84.14 | 10500 | +- | 441.8
[6 empty] | (ppb or ppm) | 14.19 ppm | +- | 12.6 ppm | [6 empty] | 2583 ppb | +- | 108.68 ppb
CNL V16 |  | 45.9848 g | 3.923 | 230331 | 2023-03-31 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-9 10-15 cm |  |  |  |  | 2023-04-04 | (mBq/kg) | 12370 | +- | 319.5 | 32200 | +- | 2829 | 582.3 | +- | 37.65 | 11900 | +- | 367.9 | 663670 | +- | 34190 | 3926.3 | +- | 227.9 | <65.56
[6 empty] | (ppb or ppm) | 1001.97 ppb | +- | 25.88 ppb | 2608.2 ppb | +- | 229.15 ppb | 1024.85 ppb | +- | 66.26 ppb | 2927.4 ppb | +- | 90.5 ppb | 21436.54 ppm | +- | 1104.34 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 41530 | +- | 119000 | <709.80 |  |  | 165.29 | +- | 65 | 11400 | +- | 412
[6 empty] | (ppb or ppm) | 3363.93 ppb | +- | 9639 ppb | [6 empty] | 2804.4 ppb | +- | 101.35 ppb
CNL V17 |  | 5.6586 g | 6.755 | 230404 | 2023-04-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-7 0-5 cm |  |  |  |  | 2023-04-11 | (mBq/kg) | 3382 | +- | 206.7 | 17150 |  | 2924 | 243.2 | +- | 52.88 | 6710 | +- | 322.6 | 170070 | +- | 9840 | 21146 | +- | 1156 | <73.91
[6 empty] | (ppb or ppm) | 273.94 ppb | +- | 16.74 ppb | 1389.15 ppb | +- | 236.84 ppb | 428.03 ppb | +- | 93.07 ppb | 1650.66 ppb | +- | 79.36 ppb | 5493.26 ppm | +- | 317.83 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 3012500 | +- | 393900 | <2021.00 |  |  | 250.25 | +- | 85.25 | 6325 | +- | 433.4
[6 empty] | (ppb or ppm) | 244.01 ppm | +- | 31.91 ppm | [6 empty] | 1555.95 ppb | +- | 106.62 ppb
CNL V18 |  | 42.6733 g | 2.64 | 230411 | 2023-04-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-9 10-15 cm |  |  |  |  | 2023-04-14 | (mBq/kg) | 8319 | +- | 256.6 | 26760 | +- | 2724 | 417.3 | +- | 41.61 | 8318 | +- | 298.6 | 618890 | +- | 32080 | 8730.7 | +- | 479.1 | <177.70
[6 empty] | (ppb or ppm) | 673.84 ppb | +- | 20.78 ppb | 2167.56 ppb | +- | 220.64 ppb | 734.45 ppb | +- | 73.23 ppb | 2046.23 ppb | +- | 73.46 ppb | 19990.15 ppm | +- | 1036.18 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 108290 | +- | 137400 | 311.44 | +- | 465.1 | 184.44 | +- | 77.41 | 7448 | +- | 367.6
[6 empty] | (ppb or ppm) | 8.77 ppm | +- | 11.13 ppm | [6 empty] | 1832.21 ppb | +- | 90.43 ppb
CNL V19 |  | 41.4027 g | 4.92 | 230414 | 2023-04-14 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-6 10-15 cm |  |  |  |  | 2023-04-19 | (mBq/kg) | 9160 | +- | 247.7 | 22370 | +- | 2222 | 459.4 | +- | 32.23 | 9486 | +- | 298.2 | 647360 | +- | 33310 | 4011.6 | +- | 227.8 | <56.49
[6 empty] | (ppb or ppm) | 741.96 ppb | +- | 20.06 ppb | 1811.97 ppb | +- | 179.98 ppb | 808.54 ppb | +- | 56.72 ppb | 2333.56 ppb | +- | 73.36 ppb | 20909.73 ppm | +- | 1075.91 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 124390 | +- | 105800 | 540.29 | +- | 412.5 | 173.6 | +- | 58.73 | 9330 | +- | 347.1
[6 empty] | (ppb or ppm) | 10.08 ppm | +- | 8.57 ppm | [6 empty] | 2295.18 ppb | +- | 85.39 ppb
CNL V20 |  | 29.3445 g | 5.894 | 230419 | 2023-04-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-8 5-10 cm |  |  |  |  | 2023-04-25 | (mBq/kg) | 12290 | +- | 313.1 | 27040 | +- | 2374 | 532 | +- | 35.45 | 10870 | +- | 335.9 | 564900 | +- | 29100 | 13061 | +- | 677.9 | 66.18 | +- | 63.58
[6 empty] | (ppb or ppm) | 995.49 ppb | +- | 25.36 ppb | 2190.24 ppb | +- | 192.29 ppb | 936.32 ppb | +- | 62.39 ppb | 2674.02 ppb | +- | 82.63 ppb | 18246.27 ppm | +- | 939.93 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 145200 | +- | 107100 | <601.50 |  |  | 203.01 | +- | 60.09 | 9552 | +- | 371.8
[6 empty] | (ppb or ppm) | 11.76 ppm | +- | 8.68 ppm | [6 empty] | 2349.79 ppb | +- | 91.46 ppb
CNL V21 |  | 27.2959 g | 6.926 | 230425 | 2023-04-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-1 0-5 cm |  |  |  |  | 2023-05-02 | (mBq/kg) | 13010 | +- | 325 | 35750 | +- | 2724 | 602.8 | +- | 36.84 | 12570 | +- | 376.4 | 663360 | +- | 34110 | 18312 | +- | 938.5 | 63.93 | +- | 67.23
[6 empty] | (ppb or ppm) | 1053.81 ppb | +- | 26.33 ppb | 2895.75 ppb | +- | 220.64 ppb | 1060.93 ppb | +- | 64.84 ppb | 3092.22 ppb | +- | 92.59 ppb | 21426.53 ppm | +- | 1101.75 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 343740 | +- | 119000 | 1077 | +- | 490.8 | 71.448 | +- | 60.62 | 11610 | +- | 417.3
[6 empty] | (ppb or ppm) | 27.84 ppm | +- | 9.64 ppm | [6 empty] | 2856.06 ppb | +- | 102.66 ppb
LEGEND Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
LEGEND V01 |  | Copper: 58.01 g | 22.462 | 240125 | 2024-01-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Majorana Copper Embedded into Teflon Holder |  | Teflon: 275.132 g |  |  | 2024-02-09 | (mBq/kg) | <0.75 |  |  | 8.16 | +- | 27.66 | <0.88 |  |  | <1.58 |  |  | <9.16 |  |  | <0.97 |  |  | <0.52
[5 empty] | 2024-02-26 | (ppb or ppm) | <60.75 ppt |  |  | 0.66 ppb | +- | 2.24 ppb | <1.55 ppb |  |  | <0.39 ppb |  |  | <295.87 ppb
[5 empty] | March 5,2024 | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 5230.7 | +- | 2585 | <9.68 |  |  | 0.562 | +- | 0.5315 | <2.17 |  |  | <1.55 |  |  | <1.06
[6 empty] | (ppb or ppm) | 0.42 ppm | +- | 0.21 ppm | [6 empty] | <0.53 ppb
DarT Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
DArT V01 | McMaster-Carr | 823.8 g | 26.57 | 240315 24031801 | 2024-03-15 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Acrylic Sheet | PN; 8560K222 |  |  |  | 2024-04-11 | (mBq/kg) | 0.5951 | +- | 0.534 | <19.40 |  |  | <0.31 |  |  | <0.51 |  |  | <6.36 |  |  | <0.35 |  |  | <0.26
| Units: 50 |  |  |  |  | (ppb or ppm) | 48.2 ppt | +- | 43.25 ppt | <1.57 ppb |  |  | <0.55 ppb |  |  | <125.46 ppt |  |  | <205.43 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 3874.5 | +- | 1856 | <3.71 |  |  | 0.21949 | +- | 0.2633 | <0.75
[6 empty] | (ppb or ppm) | 313.83 ppb | +- | 150.34 ppb | [6 empty] | <184.5 ppt
In Progress Sample Measurements Counted on the Vue des Alpes Detector
In Progress and To Be Measured:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Background 11 |  | N/A | 35.762 | 23092101 230928 230929 23092901 230930 23093001 231001 231003 231008 231012 231014 231018 23101801 231020 231023 231024 231025 231116 | Sep 21 – Oct 26, 2023 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | Nov 16- 22, 2023 | (mBq) | 1.784 | +- | 0.1798 | 4.947 | +- | 2.399 | 0.066 | +- | 0.054 | 1.522 | +- | 0.185 | 13.043 | +- | 1.742 | 0.35176 | +- | 0.1182 | <0.16
[6 empty] | Results: |  |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 153.23 | +- | 30.36 | 1.1502 | +- | 0.7831 | <0.16 |  |  | 1.045 | +- | 0.3198
SNOLAB V35 | Medium Grey | 107.4 g | 1.276 | 240430 | 2024-04-30 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Sikafloor 22NA PurCem | Type: SS-0721 |  |  |  |  | (mBq/kg) | 6096 | +- | 188.3 | 16150 | +- | 1853 | 290.4 | +- | 29.44 | 5830 | +- | 210.1 | 351500 | +- | 18310 | 40.688 | +- | 51.89 | <62.20
[6 empty] | (ppb or ppm) | 493.78 ppb | +- | 15.25 ppb | 1308.15 ppb | +- | 150.09 ppb | 511.1 ppb | +- | 51.81 ppb | 1.43 ppm | +- | 0.05 ppm | 11353.45 ppm | +- | 591.41 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 11187 | +- | 13800 | 165.85 | +- | 361.2 | 68.979 | +- | 52.77 | 6692 | +- | 295.4
[6 empty] | (ppb or ppm) | 906.15 ppb | +- | 1117.8 ppb | [6 empty] | 1.65 ppm | +- | 0.07 ppm
SNOLAB V36 |  | 200.5 g |  |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Cadbury Dark Chocolate | Made in Canada |  |  |  |  | (mBq/kg)
[6 empty] | (ppb or ppm) | 0 ppb | +- | 0 ppb | 0 ppb | +- | 0 ppb | 0 ppb | +- | 0 ppb | 0 ppb | +- | 0 ppb | 0 ppm | +- | 0 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
[6 empty] | (ppb or ppm) | #VALUE! | [8 empty] | 0 ppb | +- | 0 ppb
Next Sample | [5 empty] | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
(mBq/kg)
(ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)