## Collected Ge Detector Sample Re
Some Useful Information Concerning the results:
The Conversion factors for the primordial nuclides are given by: | [9 empty] | 1 Bq 238U/kg = 81 ppb U (81 x 10-9 gU/g) | [5 empty] | The 238U decay chain gammas used are: | [5 empty] | 226Ra: 186.1 keV
214Pb: 295.21 and 351.92 keV
214Bi: 609.31, 1120.29, 1764.49 and 2204.21 keV
The relationships are valid for any daughters in the 238U, 235U or 232Th chain only if the chain is in equilibrium. | [9 empty] | 1 Bq 232Th/kg = 246 ppb Th (246 x 10-9 gTh/g) | [5 empty] | The 232Th decay chain gammas used are: | [5 empty] | 212Pb: 238.63 and 300.09 keV
208Tl: 583.19 and 2614.53 keV,
228Ac: 911.21 keV
[10 empty] | 1 Bq 40K/kg = 32300 ppb K (32300 x 10-6 gK/g) | [5 empty] | The 40K decay chain gamma used is: | [5 empty] | 40K: 1460.83 keV
[10 empty] | 1 Bq 235U/kg = 1.76 ppm U (1.76 x 10-6 gU/g) | [5 empty] | The 235U decay chain gammas used are: | [5 empty] | 235U: 143.76, 163.33 and 205.31 keV
Background Measurements of the Vue des Alpes Detector
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Background 1 | Efficiency assumes no samples |  | 36.187 | 171213 171218 | 2017-12-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-01-19 | (mBq) | 1.873 | +- | 0.1674 | 5.039 | +- | 1.612 | 0.0612 | +- | 0.05173 | 1.354 | +- | 0.1679 | 15.262 | +- | 1.927 | 0.2008 | +- | 0.096 | 0.06328 | +- | 0.05422
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 76.664 | +- | 14.04 | <0.67 |  |  | <0.081 |  |  | 0.8836 | +- | 0.2577
Background 2 | Efficiency assumes no samples |  | 17.497 | 180212 | 2018-02-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-03-02 | (mBq) | 1.322 | +- | 0.2181 | 3.725 | +- | 2.318 | 0.2334 | +- | 0.07634 | 0.7936 | +- | 0.2178 | 18.386 | +- | 2.838 | 0.39164 | +- | 0.1591 | 0.02449 | +- | 0.08185
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.72 | +- | 20.73 | <0.98 |  |  | <0.12 |  |  | 1.033 | +- | 0.3662
Background 3 | Efficiency assumes no samples |  | 51.75 | 180319 180417 | 2018-03-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-05-12 | (mBq) | 2.225 | +- | 0.1474 | 4.95 | +- | 1.385 | 0.14 | +- | 0.0455 | 1.274 | +- | 0.1334 | 17.109 | +- | 1.737 | 0.43336 | +- | 0.09 | <0.068
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.15 | +- | 12.63 | 0.5036 | +- | 0.55 | 0.075 | +- | 0.065 | 1.212 | +- | 0.2177
Background 4 | Efficiency assumes no samples |  | 82.542 | 180522 18053001 180801 | 2018-05-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-08-12 | (mBq) | 2.053 | +- | 0.1186 | 6.239 | +- | 1.172 | 0.1237 | +- | 0.03392 | 1.328 | +- | 0.1141 | 15.52 | +- | 1.393 | 0.31949 | +- | 0.0674 | 0.1219 | +- | 0.03733
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 79.586 | +- | 9.876 | 0.731 | +- | 0.4578 | 0.0398 | +- | 0.04739 | 1.13 | +- | 0.18
Background 5 | Efficiency assumes no samples |  | 76.532 | 190910 190914 190915 | 2019-09-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2019-11-28 | (mBq) | 2.04 | +- | 0.122 | 6.373 | +- | 1.193 | 0.133 | +- | 0.03459 | 1.164 | +- | 0.1159 | 16.218 | +- | 1.452 | 0.38528 | +- | 0.0757 | 0.097 | +- | 0.0385
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 82.073 | +- | 10.15 | 0.48395 | +- | 0.443 | <0.052 |  |  | 1.477 | +- | 0.1925
Background 6 | Efficiency assumes no samples |  | 55.304 | 200324 200430 | 2020-03-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2020-05-19 | (mBq) | 3.397 | +- | 0.1763 | 7.801 | +- | 1.57 | 0.1535 | +- | 0.04259 | 2.404 | +- | 0.1694 | 29.088 | +- | 2.378 | 0.24865 | +- | 0.0871 | 0.03165 | +- | 0.05048
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 191.42 | +- | 21.95 | 0.48686 | +- | 0.5507 | 0.043088 | +- | 0.06009 | 2.092 | +- | 0.2595
Background 7 | Efficiency assumes no samples |  | 7.257 | 210111 210119 | 2021-01-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2021-01-20 | (mBq) | 2.063 | +- | 0.3767 | 8.573 | +- | 3.769 | 0.2188 | +- | 0.1077 | 1.139 | +- | 0.3709 | 29.316 | +- | 5.234 | 0.36096 | +- | 0.2337 | <0.12
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 172.52 | +- | 44.42 | <1.55 |  |  | <0.21 |  |  | 1.727 | +- | 0.6023
Combined Background | Combined Backgrounds of Runs 1+2+3+4+5 |  | 238.322 |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector | Efficiency assumes no samples |  |  |  |  | (mBq) | 2.414 | +- | 0.094 | 6.61 | +- | 0.85 | 0.333 | +- | 0.06 | 1.534 | +- | 0.082 | 18.6526 | +- | 1.2214 | 0.4053 | +- | 0.0522 | 0.151 | +- | 0.026
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 99.2049 | +- | 7.7138 | 0.3194 | +- | 0.3263 | 0.0063 | +- | 0.0342 | 1.394 | +- | 0.041
Completed Sample Measurements Counted on the Vue des Alpes Detector
The measurements of the samples below take into account the background measurement shown above. If a measurement is below the background then the upper bound shown is the 90% confidence limit.
PICO Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
PICO V01 | Model No. S1328PD3T72 Omega | 60.1 g | 15.659 | 200825 | 2020-08-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
White RTD Wire Leads |  |  |  |  | 2020-09-10 | (mBq/kg) | 5.375 | +- | 7.215 | <127.80 |  |  | 2.542 | +- | 2.395 | <9.16 |  |  | 69.281 | +- | 72.04 | <5.89 |  |  | <1.78
[6 empty] | (ppb or ppm) | 0.44 ppb | +- | 0.58 ppb | <10.35 ppb |  |  | 4.47 ppb | +- | 4.22 ppb | <2.25 ppb |  |  | 2237.78 ppb | +- | 2326.89 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 140700 | +- | 38520 | <53.17 |  |  | <1.90 |  |  | <14.98
[6 empty] | (ppb or ppm) | 11.4 ppm | +- | 3.12 ppm | [6 empty] | <3.69 ppb
PICO V02 | Model No. S1328PD3T72 Omega | 23.9 g | 11.639 | 200910 | 2020-09-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Red RTD Wire Leads |  |  |  |  | 2020-09-22 | (mBq/kg) | 4.669 | +- | 20.77 | <422.90 |  |  | 6.185 | +- | 6.586 | <27.98 |  |  | 74.988 | +- | 232.8 | <17.53 |  |  | <3.64
[6 empty] | (ppb or ppm) | 0.38 ppb | +- | 1.68 ppb | <34.25 ppb |  |  | 10.89 ppb | +- | 11.59 ppb | <6.88 ppb |  |  | 2.42 ppm | +- | 7.52 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 148160 | +- | 69770 | 132.14 | +- | 84.27 | <14.75 |  |  | <42.97
[6 empty] | (ppb or ppm) | 12 ppm | +- | 5.65 ppm | [6 empty] | <10.57 ppb
PICO V03 | Hard Copper Wire | 45.2 g | 13.687 | 200922 | 2020-09-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
PICO Heater Cables |  |  |  |  | 2020-10-06 | (mBq/kg) | <13.25 |  |  | <127.70 |  |  | <3.89 |  |  | 6.695 | +- | 9.819 | <148.60 |  |  | 9.1666 | +- | 6.88 | <5.33
[6 empty] | (ppb or ppm) | <1.07 ppb |  |  | <10.34 ppb |  |  | <6.85 ppb |  |  | 1.65 ppb | +- | 2.42 ppb | <4.8 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 58Co |  |  | 57Co
[6 empty] | (mBq/kg) | 143110 | +- | 62600 | 3.3293 | +- | 41.87 | 2.5467 | +- | 4.666 | 21.23 | +- | 16.07 | 0.54315 | +- | 3.654 | 7.049 | +- | 6.27
[6 empty] | (ppb or ppm) | 11.59 ppm | +- | 5.07 ppm | [6 empty] | 5.22 ppb | +- | 3.95 ppb
PICO V04 | Preamps for Piezos for the PICO and SBC experiments | 21.2 g | 25.61 | 210421 210429 | 2021-04-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Kapton Preamp Boards |  | Quantity: 28 |  |  | 2021-05-17 | (mBq/kg) | 1504 | +- | 45.52 | 4544 | +- | 459.1 | 89.11 | +- | 7.122 | 2226 | +- | 71.56 | 4239.8 | +- | 321.9 | <9.34 |  |  | <5.98
[6 empty] | (ppb or ppm) | 121.82 ppb | +- | 3.69 ppb | 368.06 ppb | +- | 37.19 ppb | 156.83 ppb | +- | 12.53 ppb | 547.6 ppb | +- | 17.6 ppb | 136.95 ppm | +- | 10.4 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 259030 | +- | 88590 | <108.00 |  |  | 12.595 | +- | 9.45 | 2254 | +- | 80.58
[6 empty] | (ppb or ppm) | 20.98 ppm | +- | 7.18 ppm | [6 empty] | 554.48 ppb | +- | 19.82 ppb
PICO V05 | BLU-25751706 Lot 230328 3/32 “ x 7” | 276.2 g | 5.088 | 231026 23102601 231028 231029 | 2023-10-26 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  | 231030 | 2023-10-31 | (mBq/kg) | <186.40 |  |  | 9577 | +- | 4406 | <193.20 |  |  | 570400 | +- | 13220 | 14541 | +- | 902.7 | 61.79 | +- | 71.9 | <50.83
| Sample contains 20 electrodes |  |  |  |  | (ppb or ppm) | <15.1 ppb |  |  | 775.74 ppb | +- | 356.89 ppb | <340.03 ppb |  |  | 140.32 ppm | +- | 3.25 ppm | 469.67 ppm | +- | 29.16 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <2.62E6 |  |  | 2903 | +- | 633.1 | 2985.6 | +- | 182.1 | 207300 | +- | 4975
[6 empty] | (ppb or ppm) | <212.22 ppm | [8 empty] | 51 ppm | +- | 1.22 ppm
PICO V06 | BLU-25751706 Lot 220713 3/32 “ x 7” | 422.2 g | 4.687 | 231031 23103101 231101 231102 231103 | 2023-10-31 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  |  | 2023-11-05 | (mBq/kg) | <68.47 |  |  | 927.8 | +- | 1908 | 64.68 | +- | 53.28 | 93870 | +- | 2196 | 7938.6 | +- | 475.8 | <66.87 |  |  | <10.82
| Sample contains 30 electrodes |  |  |  |  | (ppb or ppm) | <5.55 ppb |  |  | 75.15 ppb | +- | 154.55 ppb | 113.84 ppb | +- | 93.77 ppb | 23.09 ppm | +- | 0.54 ppm | 256.42 ppm | +- | 15.37 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2324000 | +- | 2368000 | 1866.6 | +- | 317.7 | 1527 | +- | 84.04 | 105300 | +- | 2557
[6 empty] | (ppb or ppm) | 188.24 ppm | +- | 191.81 ppm | [6 empty] | 25.9 ppm | +- | 0.63 ppm
PICO V07 | BLU-25751708 Lot 230204 1/8 “ x 7” | 251.5 g | 3.126 | 231105 | 2023-11-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  |  | 2023-11-08 | (mBq/kg) | <115.70 |  |  | <7092.00 |  |  | <235.50 |  |  | 577200 | +- | 13410 | 16792 | +- | 1114 | <111.70 |  |  | <59.50
| Sample contains 10 electrodes |  |  |  |  | (ppb or ppm) | <9.37 ppb |  |  | <574.45 ppb |  |  | <414.48 ppb |  |  | 141.99 ppm | +- | 3.3 ppm | 542.38 ppm | +- | 35.98 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <7.92E6 |  |  | 3190.9 | +- | 949.6 | 3172.3 | +- | 181.6 | 222400 | +- | 5483
[6 empty] | (ppb or ppm) | <641.52 ppm | [8 empty] | 54.71 ppm | +- | 1.35 ppm
PICO V08 | BLU-25751708 Lot 230515 1/8 “ x 7” | 1004.4 g | 7.786 | 231108 | 2023-11-08 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  |  | 2023-11-16 | (mBq/kg) | <33.37 |  |  | 263.4 | +- | 2273 | 299.5 | +- | 113.6 | 599300 | +- | 14060 | 17403 | +- | 929.6 | <59.25 |  |  | <14.68
| Sample contains 40 electrodes |  |  |  |  | (ppb or ppm) | <2.7 ppb |  |  | 21.34 ppb | +- | 184.11 ppb | 527.12 ppb | +- | 199.94 ppb | 147.43 ppm | +- | 3.46 ppm | 562.12 ppm | +- | 30.03 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <1.79E7 |  |  | 3007.2 | +- | 395 | 2659.2 | +- | 146.1 | 213800 | +- | 5309
[6 empty] | (ppb or ppm) | <1449.9 ppm | [8 empty] | 52.59 ppm | +- | 1.31 ppm
PICO V09 | Molex wire ring connectors | 100.6 g | 8.757 | 240305 240313 | 2024-03-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Wire Ring connectors | Types: 16-14 MX 22-18 MX |  |  |  | 2024-03-15 | (mBq/kg) | <4.41 |  |  | 99.35 | +- | 249 | <4.27 |  |  | <5.03 |  |  | <64.56 |  |  | <3.73 |  |  | <3.75
[6 empty] | (ppb or ppm) | <357.21 ppt |  |  | 8.05 ppb | +- | 20.17 ppb | <7.52 ppb |  |  | <1.24 ppb |  |  | <2.09 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 371410 | +- | 337700 | 46.919 | +- | 30.78 | 7.5284 | +- | 3.965 | <7.22
[6 empty] | (ppb or ppm) | 30.08 ppm | +- | 27.35 ppm | [6 empty] | <1.78 ppb
PICO V10 |  | 383.0 g | 10.291 | 240920 24092001 240830 | 2024-09-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Plate |  |  |  |  | 2024-10-01 | (mBq/kg) | 0.5482 | +- | 0.9116 | <27.33 |  |  | <0.64 |  |  | 1.471 | +- | 1.076 | <06.69 |  |  | <0.68 |  |  | 6.163 | +- | 0.7925
[6 empty] | (ppb or ppm) | 44.4 ppt | +- | 73.84 ppt | <2.21 ppb |  |  | <1.13 ppb |  |  | 0.36 ppb | +- | 0.26 ppb | <216.09 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 4459.9 | +- | 23260 | 0.4475 | +- | 3.544 | 2.3468 | +- | 0.6015 | 3.2 | +- | 1.638 | 2.426 | +- | 0.7925 | 0.40873 | +- | 0.4017
[6 empty] | (ppb or ppm) | 361.25 ppb | +- | 1884.06 ppb | [6 empty] | 0.79 ppb | +- | 0.4 ppb
PICO V11 |  | 763.7 g | 7.554 | 241001 241002 24100201 | 2024-10-01 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Welded Stainless Steel Plate |  |  |  |  | 2024-10-09 | (mBq/kg) | <1.32 |  |  | 51.96 | +- | 30.57 | <0.59 |  |  | <1.17 |  |  | <9.45 |  |  | <0.60 |  |  | 5.266 | +- | 0.6924
[6 empty] | (ppb or ppm) | <106.92 ppt |  |  | 4.21 ppb | +- | 2.48 ppb | <1.04 ppb |  |  | <0.29 ppb |  |  | <0.31 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 15959 | +- | 34140 | 2.4727 | +- | 3.251 | 3.4827 | +- | 0.6545 | 3.532 | +- | 1.458 | 1.187 | +- | 1.444 | 0.58289 | +- | 0.3819
[6 empty] | (ppb or ppm) | 1.29 ppm | +- | 2.77 ppm | [6 empty] | 0.87 ppb | +- | 0.36 ppb
PICO V12 |  | 444.3 g | 14.732 | 241009 | 2024-10-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Disk |  |  |  |  | 2024-10-24 | (mBq/kg) | <0.43 |  |  | <51.63 |  |  | 0.904 | +- | 0.3434 | 7.862 | +- | 1.03 | <7.06 |  |  | <0.53 |  |  | 6.736 | +- | 0.6606
[6 empty] | (ppb or ppm) | <0.03 ppb |  |  | <4.18 ppb |  |  | 1.59 ppb | +- | 0.6 ppb | 1.93 ppb | +- | 0.25 ppb | <0.23 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | <17240.00 |  |  | <3.94 |  |  | 1.2795 | +- | 0.5304 | 4.564 | +- | 1.347 | <0.70 |  |  | 0.16835 | +- | 0.3615
[6 empty] | (ppb or ppm) | <1.4 ppm | [8 empty] | 1.12 ppb | +- | 0.33 ppb
PICO V13 |  | 249.6 g | 13.72 | 241024 241030 | 2024-10-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Ring |  |  |  |  | 2024-11-07 | (mBq/kg) | <0.88 |  |  | 18.07 | +- | 40.55 | <0.56 |  |  | <2.68 |  |  | <7.83 |  |  | <1.13 |  |  | 5.561 | +- | 0.8653
[6 empty] | (ppb or ppm) | <71.28 ppt |  |  | 1.46 ppb | +- | 3.28 ppb | <0.99 ppb |  |  | <0.66 ppb |  |  | <252.91 ppb
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 3004.7 | +- | 12910 | <5.74 |  |  | 1.0764 | +- | 0.6557 | <2.61 |  |  | 3.221 | +- | 1.582 | <1.25
[6 empty] | (ppb or ppm) | 243.38 ppb | +- | 1045.71 ppb | [6 empty] | <0.64 ppb
In Progress Sample Measurements Counted on the Vue des Alpes Detector
In Progress and To Be Measured:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Next Sample | [5 empty] | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
(mBq/kg)
(ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)