## Collected Ge Detector Sample Re
Some Useful Information Concerning the results:
The Conversion factors for the primordial nuclides are given by: | [9 empty] | 1 Bq 238U/kg = 81 ppb U (81 x 10-9 gU/g) | [5 empty] | The 238U decay chain gammas used are: | [5 empty] | 226Ra: 186.1 keV
214Pb: 295.21 and 351.92 keV
214Bi: 609.31, 1120.29, 1764.49 and 2204.21 keV
The relationships are valid for any daughters in the 238U, 235U or 232Th chain only if the chain is in equilibrium. | [9 empty] | 1 Bq 232Th/kg = 246 ppb Th (246 x 10-9 gTh/g) | [5 empty] | The 232Th decay chain gammas used are: | [5 empty] | 212Pb: 238.63 and 300.09 keV
208Tl: 583.19 and 2614.53 keV,
228Ac: 911.21 keV
[10 empty] | 1 Bq 40K/kg = 32300 ppb K (32300 x 10-6 gK/g) | [5 empty] | The 40K decay chain gamma used is: | [5 empty] | 40K: 1460.83 keV
[10 empty] | 1 Bq 235U/kg = 1.76 ppm U (1.76 x 10-6 gU/g) | [5 empty] | The 235U decay chain gammas used are: | [5 empty] | 235U: 143.76, 163.33 and 205.31 keV
Background Measurements of the Vue des Alpes Detector
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Background 1 | Efficiency assumes no samples |  | 36.187 | 171213 171218 | 2017-12-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-01-19 | (mBq) | 1.873 | +- | 0.1674 | 5.039 | +- | 1.612 | 0.0612 | +- | 0.05173 | 1.354 | +- | 0.1679 | 15.262 | +- | 1.927 | 0.2008 | +- | 0.096 | 0.06328 | +- | 0.05422
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 76.664 | +- | 14.04 | <0.67 |  |  | <0.081 |  |  | 0.8836 | +- | 0.2577
Background 2 | Efficiency assumes no samples |  | 17.497 | 180212 | 2018-02-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-03-02 | (mBq) | 1.322 | +- | 0.2181 | 3.725 | +- | 2.318 | 0.2334 | +- | 0.07634 | 0.7936 | +- | 0.2178 | 18.386 | +- | 2.838 | 0.39164 | +- | 0.1591 | 0.02449 | +- | 0.08185
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
Background 3 | Efficiency assumes no samples |  | 51.75 | 180319 180417 | 2018-03-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-05-12 | (mBq) | 2.225 | +- | 0.1474 | 4.95 | +- | 1.385 | 0.14 | +- | 0.0455 | 1.274 | +- | 0.1334 | 17.109 | +- | 1.737 | 0.43336 | +- | 0.09 | <0.068
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.15 | +- | 12.63 | 0.5036 | +- | 0.55 | 0.075 | +- | 0.065 | 1.212 | +- | 0.2177
Background 4 | Efficiency assumes no samples |  | 82.542 | 180522 18053001 180801 | 2018-05-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-08-12 | (mBq) | 2.053 | +- | 0.1186 | 6.239 | +- | 1.172 | 0.1237 | +- | 0.03392 | 1.328 | +- | 0.1141 | 15.52 | +- | 1.393 | 0.31949 | +- | 0.0674 | 0.1219 | +- | 0.03733
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 79.586 | +- | 9.876 | 0.731 | +- | 0.4578 | 0.0398 | +- | 0.04739 | 1.13 | +- | 0.18
Background 5 | Efficiency assumes no samples |  | 76.532 | 190910 190914 190915 | 2019-09-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2019-11-28 | (mBq) | 2.04 | +- | 0.122 | 6.373 | +- | 1.193 | 0.133 | +- | 0.03459 | 1.164 | +- | 0.1159 | 16.218 | +- | 1.452 | 0.38528 | +- | 0.0757 | 0.097 | +- | 0.0385
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 82.073 | +- | 10.15 | 0.48395 | +- | 0.443 | <0.052 |  |  | 1.477 | +- | 0.1925
Background 6 | Efficiency assumes no samples |  | 55.304 | 200324 200430 | 2020-03-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2020-05-19 | (mBq) | 3.397 | +- | 0.1763 | 7.801 | +- | 1.57 | 0.1535 | +- | 0.04259 | 2.404 | +- | 0.1694 | 29.088 | +- | 2.378 | 0.24865 | +- | 0.0871 | 0.03165 | +- | 0.05048
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 191.42 | +- | 21.95 | 0.48686 | +- | 0.5507 | 0.043088 | +- | 0.06009 | 2.092 | +- | 0.2595
Background 7 | Efficiency assumes no samples |  | 7.257 | 210111 210119 | 2021-01-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2021-01-20 | (mBq) | 2.063 | +- | 0.3767 | 8.573 | +- | 3.769 | 0.2188 | +- | 0.1077 | 1.139 | +- | 0.3709 | 29.316 | +- | 5.234 | 0.36096 | +- | 0.2337 | <0.12
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 172.52 | +- | 44.42 | <1.55 |  |  | <0.21 |  |  | 1.727 | +- | 0.6023
Background 8 |  | N/A | 67.574 | 210517 210524 210627 210630 | 2021-05-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2021-07-25 | (mBq) | 1.906 | +- | 0.1261 | 4.782 | +- | 1.185 | 0.179 | +- | 0.03537 | 1.278 | +- | 0.1224 | 19.493 | +- | 1.701 | 0.38762 | +- | 0.08147 | 0.03217 | +- | 0.04114
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 141.39 | +- | 16.91 | <1.25 |  |  | <0.061 |  |  | 1.183 | +- | 0.2005
Background 11 |  | N/A | 35.762 | 23092101 230928 230929 23092901 230930 23093001 231001 231003 231008 231012 231014 231018 23101801 231020 231023 231024 231025 231116 | Sep 21 – Oct 26, 2023 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | Nov 16- 22, 2023 | (mBq) | 1.784 | +- | 0.1798 | 4.947 | +- | 2.399 | 0.066 | +- | 0.054 | 1.522 | +- | 0.185 | 13.043 | +- | 1.742 | 0.35176 | +- | 0.1182 | <0.16
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 153.23 | +- | 30.36 | 1.1502 | +- | 0.7831 | <0.16 |  |  | 1.045 | +- | 0.3198
Background 12 |  | N/A | 39.306 | 240613 240707 240710 | 2024-06-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2024-07-23 | (mBq) | 1.64 | +- | 0.1529 | 10.95 | +- | 2.147 | 0.12 | +- | 0.04643 | 1.415 | +- | 0.1524 | 14.151 | +- | 1.648 | 0.285 | +- | 0.1002 | <0.079
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 177.62 | +- | 27.67 | 1.1773 | +- | 0.687 | 0.05895 | +- | 0.06813 | 0.9347 | +- | 0.2503
Combined Background | Combined Backgrounds of Runs 1+2+3+4+5 |  | 238.322 |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector | Efficiency assumes no samples |  |  |  |  | (mBq) | 2.414 | +- | 0.094 | 6.61 | +- | 0.85 | 0.333 | +- | 0.06 | 1.534 | +- | 0.082 | 18.6526 | +- | 1.2214 | 0.4053 | +- | 0.0522 | 0.151 | +- | 0.026
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 99.2049 | +- | 7.7138 | 0.3194 | +- | 0.3263 | 0.0063 | +- | 0.0342 | 1.394 | +- | 0.041
Completed Sample Measurements Counted on the Vue des Alpes Detector
The measurements of the samples below take into account the background measurement shown above. If a measurement is below the background then the upper bound shown is the 90% confidence limit.
SNOLAB Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
SNOLAB V01 | Background is NOT subtracted |  | 20.583 | 170901 | 2017-09-01 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Acrylic Sample Beaker H | Efficiency assumes no samples |  |  |  | 2017-09-22 | (mBq) | 1.439 | +- | 0.145 | 0.182 | +- | 0.729 | 0.354 | +- | 0.104 | 0.798 | +- | 0.114 | 7.886 | +- | 1.318 | 0.197 | +- | 0.082 | 0.04 | +- | 0.048
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 32.615 | +- | 5.256 | 0.38 | +- | 0.563 | <0.052 |  |  | 0.455 | +- | 0.199
SNOLAB V01A | Background Subtracted |  | 20.853 | 170901 | 2017-09-01 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Acrylic Sample Beaker H | Efficiency assumes no samples |  |  |  | 2017-09-22 | (mBq) | 0.294 | +- | 0.142 | <0.095 |  |  | 0.149 | +- | 0.105 | 0.007 | +- | 0.115 | <0.26 |  |  | <0.096 |  |  | 0.012 | +- | 0.05
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 12.805 | +- | 5.104 | 0.38 | +- | 0.563 | <0.052 |  |  | <0.082
SNOLAB V02 | Lot No. A4B8 02/20 | 1871.5 g | 6.979 | 180511 | 2018-05-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Alconox | Alconox Inc. |  |  |  | 2018-05-18 | (mBq/kg) | 184.6 | +- | 4.642 | 104.5 | +- | 18.43 | 5.728 | +- | 0.54 | 139.5 | +- | 4.386 | 8664 | +- | 446.7 | <0.65 |  |  | 0.8054 | +- | 0.7874
[6 empty] | (ppb or ppm) | 14.95 ppb | +- | 0.38 ppb | 8.46 ppb | +- | 1.49 ppb | 10.08 ppb | +- | 0.95 ppb | 34.32 ppb | +- | 1.08 ppb | 279.85 ppm | +- | 14.43 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <79.93 |  |  | 5.6059 | +- | 5.455 | <1.21 |  |  | 61.58 | +- | 3.484
[6 empty] | (ppb or ppm) | <6.47 ppb | [8 empty] | 15.15 ppb | +- | 0.86 ppb
SNOLAB V03 | Sterling AP 308/3-8L, AWS E308-16, E308L-16 | 2294.9 g | 3.91 | 180518 | 2018-05-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Premium Welding Electrodes --- Calibration | McKay Welding Products, Troy, Ohio, Lot No 24L056, Heat No. 738228 |  |  |  | 2018-05-22 | (mBq/kg) | 32330 | +- | 564.6 | 15260 | +- | 854.5 | 1110 | +- | 24.75 | 4810 | +- | 116.1 | 261860 | +- | 13380 | <1.68 |  |  | <6.29
[6 empty] | (ppb or ppm) | 2.62 ppm | +- | 0.05 ppm | 1.24 ppm | +- | 0.07 ppm | 1.95 ppm | +- | 0.04 ppm | 1.18 ppm | +- | 0.03 ppm | 8458.08 ppm | +- | 432.17 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 414.69 | +- | 1085 |  | NM |  |  | NM |  | 4376 | +- | 120.2
[6 empty] | (ppb or ppm) | 33.59 ppb | +- | 87.89 ppb | [6 empty] | 1.08 ppm | +- | 0.03 ppm
SNOLAB V04 | Sources prepared on Nov 21, 2019 | 4.4 g | 12.638 | 191129 | 2019-11-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Calibration Source Wipes | The Kimwipes are from sources SRS-19-004 |  | Standard Background Subtraction Using Combined Backgrounds |  | 2019-12-11 | (mBq/kg) | 455.6 | +- | 79.18 | 1335 | +- | 733.5 | 12.51 | +- | 22.94 | 300.5 | +- | 76.03 | 5534.6 | +- | 1118 | <48.00 |  |  | 11.39 | +- | 31.69
[6 empty] | (ppb or ppm) | 36.9 ppb | +- | 6.41 ppb | 108.14 ppb | +- | 59.41 ppb | 22.02 ppb | +- | 40.37 ppb | 73.92 ppb | +- | 18.7 ppb | 178.77 ppm | +- | 36.11 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 986.42 | +- | 6241 | <229.50 |  |  | 28.188 | +- | 26.26 | 99.8 | +- | 116.5
[6 empty] | (ppb or ppm) | 79.9 ppb | +- | 505.52 ppb | [6 empty] | 24.55 ppb | +- | 28.66 ppb
SNOLAB V04A | Sources prepared on Nov 21, 2019 | 4.4 g | 12.638 | 191129 | 2019-11-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Calibration Source Wipes | The Kimwipes are from sources SRS-19-004 |  | Background Subtraction Using The Results from SNOLAB 05 |  | 2019-12-11 | (mBq/kg) | 8.961 | +- | 114.2 | 264.9 | +- | 989.3 | <23.57 |  |  | 18.4 | +- | 111.7 | 0.00306 | +- | 1567 | <137.40 |  |  | 0.0228 | +- | 38.3
[6 empty] | (ppb or ppm) | 0.73 ppb | +- | 9.25 ppb | 21.46 ppb | +- | 80.13 ppb | <41.48 ppb |  |  | 4.53 ppb | +- | 27.48 ppb | 0 ppm | +- | 50.61 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 0.0168 | +- | 24810 | <944.30 |  |  | 0.025 | +- | 47.83 | 6.347 | +- | 165.8
[6 empty] | (ppb or ppm) | 0 ppb | +- | 2009.61 ppb | [6 empty] | 1.56 ppb | +- | 40.79 ppb
SNOLAB V05 | Sample prepared on Dec 05, 2019 | 4.4 g | 7.882 | 191211 | 2019-12-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Kimwipes Blank | These are clean Kimwipes |  | Standard Background Subtraction Using Combined Backgrounds |  | 2019-12-19 | (mBq/kg) | 40.49 | +- | 84.56 | 801.4 | +- | 823.4 | <14.59 |  |  | 55.51 | +- | 88.79 | 635.3 | +- | 990.9 | <83.35 |  |  | <43.88
[6 empty] | (ppb or ppm) | 3.28 ppb | +- | 6.85 ppb | 64.91 ppb | +- | 66.7 ppb | <25.68 ppb |  |  | 13.66 ppb | +- | 21.84 ppb | 20.52 ppm | +- | 32.01 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 20521 | +- | 11650 | 212.32 | +- | 352.6 | 17.376 | +- | 29.66 | 115.1 | +- | 142.6
[6 empty] | (ppb or ppm) | 1662.2 ppb | +- | 943.65 ppb | [6 empty] | 28.31 ppb | +- | 35.08 ppb
SNOLAB V06 | Vial cap is downwards | 0.832 g | 17.701 | 191219 191220 | 2019-12-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-K2 nEXO Calibration Sample | Specific Activity: 14.0 +- 0.09 Bq of 40K / [g RGK-1] |  |  |  | 2020-01-06 | (mBq/kg) | 4315 | +- | 513.7 | 8816 | +- | 15340 | 146.6 | +- | 180.2 | 297.1 | +- | 420.2 | 15233000 | +- | 801100 | 615.97 | +- | 484.9 | <981.60
| [11.57 – 11.72] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 349.52 ppb | +- | 41.61 ppb | 714.1 ppb | +- | 1242.54 ppb | 258.02 ppb | +- | 317.15 ppb | 73.09 ppb | +- | 103.37 ppb | 492025.9 ppm | +- | 25875.53 ppm
[6 empty] | (Bq) | [12 empty] | 12.6739 | +- | 0.50149
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 795110 | +- | 2797000 | 1644 | +- | 3001 | 594.49 | +- | 511.2 | <1864.00
[6 empty] | (ppb or ppm) | 64.4 ppm | +- | 226.56 ppm | [6 empty] | <458.54 ppb
SNOLAB V07 | Vial cap is upwards | 2.046 g | 1.937 | 200106 | 2020-01-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-U1 nEXO Calibration Sample | Specific Activity: 4.94 +- 0.03 Bq of 238U / [g RGU-1] |  |  |  | 2020-01-08 | (mBq/kg) | 5114000 | +- | 93980 | 6168000 | +- | 377900 | 231300 | +- | 5404 | 2079 | +- | 1486 | 4783.5 |  | 28500 | <637.00 |  |  | <3318.00
| [10.04 – 10.17] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 414.23 ppm | +- | 7.61 ppm | 499.61 ppm | +- | 30.61 ppm | 407.09 ppm | +- | 9.51 ppm | 511.43 ppb | +- | 365.56 ppb | 154.51 ppm | +- | 920.55 ppm
[6 empty] | (Bq) | 10.4632 | +- | 0.19228 | 12.6197 | +- | 0.77318 |  | Average: |  | 10.5888 | +- | 0.1866
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 66658000 | +- | 9102000 | <10990.00 |  |  | <960.20 |  |  | <8204.00
[6 empty] | (ppb or ppm) | 5399.3 ppm | +- | 737.26 ppm | [6 empty] | <2.02 ppm
[6 empty] | (Bq) | 136.3823 | +- | 18.6227
SNOLAB V08 | Vial cap is upwards | 3.327 g | 5 | 200108 | 2020-01-08 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-T1 nEXO Calibration Sample | Specific Activity: 3.25 +- 0.07 Bq of 232Th / [g RGTh-1] |  |  |  | 2020-01-13 | (mBq/kg) | 50150 | +- | 2512 | 511500 | +- | 73360 | 3829 | +- | 540.7 | 2961000 | +- | 68880 | 233380 | +- | 15970 | <1441.00 |  |  | <729.70
| [10.58 – 11.05] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 4.06 ppm | +- | 0.2 ppm | 41.43 ppm | +- | 5.94 ppm | 6.74 ppm | +- | 0.95 ppm | 728.41 ppm | +- | 16.94 ppm | 7538.17 ppm | +- | 515.83 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 9.8512 | +- | 0.22916
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <1159000.00 |  |  | 23356 | +- | 7448 | 26100 | +- | 1789 | 2804000 | +- | 64000
[6 empty] | (ppb or ppm) | <93.88 ppm | [8 empty] | 689.78 ppm | +- | 15.74 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 9.3289 | +- | 0.21293 |  | Average: |  | 9.5709 | +- | 0.15599
SNOLAB V09 | Vial cap is upwards | 0.832 g | 6.896 | 200113 | 2020-01-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-K2 nEXO Calibration Sample | Specific Activity: 14.0 +- 0.09 Bq of 40K / [g RGK-1] |  |  |  | 2020-01-20 | (mBq/kg) | 494.3 | +- | 686.5 | 1890 | +- | 24090 | <437.50 |  |  | 918.8 | +- | 767.2 | 15495000 | +- | 818300 | 763.17 | +- | 785.3 | <1728.00
| [11.57 – 11.72 Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 40.04 ppb | +- | 55.61 ppb | 153.09 ppb | +- | 1951.29 ppb | <0.77 ppm |  |  | 226.02 ppb | +- | 188.73 ppb | 500488.5 ppm | +- | 26431.09 ppm
[6 empty] | (Bq) | [12 empty] | 12.8918 | +- | 0.68083
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 9895400 | +- | 4772000 | 2661.3 | +- | 4678 | 180.42 | +- | 825.5 | <1627.00
[6 empty] | (ppb or ppm) | 801.53 ppm | +- | 386.53 ppm | [6 empty] | <400.24 ppb
SNOLAB V10 | G.F. Thompson Co Ltd. | 78.7 g | 2.104 | 200120 | 2020-01-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Masters Pro-Dope with Teflon | Newmarket ON |  |  |  | 2020-01-22 | (mBq/kg) | 10880 | +- | 279.9 | 23870 | +- | 2191 | 555 | +- | 34.31 | 12220 | +- | 361.1 | 53843 | +- | 3100 | <64.63 |  |  | 20.91 | +- | 27.51
Pipe Sealant | [5 empty] | (ppb or ppm) | 881.28 ppb | +- | 22.67 ppb | 1933.47 ppb | +- | 177.47 ppb | 976.8 ppb | +- | 60.39 ppb | 3006.12 ppb | +- | 88.83 ppb | 1739.13 ppm | +- | 100.13 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 113630 | +- | 135500 | 340.15 | +- | 308.3 | 82.984 | +- | 43.57 | 13140 | +- | 422.8
[6 empty] | (ppb or ppm) | 9204.03 ppb | +- | 10975.5 ppb | [6 empty] | 3232.44 ppb | +- | 104.01 ppb
SNOLAB V11 | Vial cap is upwards | 2.102 g | 4.791 | 200122 | 2020-01-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-U2 nEXO Calibration Sample | Specific Activity: 4.94 +- 0.03 Bq of 238U / [g RGU-1] |  |  |  | 2020-01-27 | (mBq/kg) | 4911000 | +- | 88240 | 5713000 | +- | 311500 | 221400 | +- | 4801 | 2384 | +- | 993.7 | <15330.00 |  |  | <336.50 |  |  | <1583.00
| [10.32 – 10.45] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 397.79 ppm | +- | 7.15 ppm | 462.75 ppm | +- | 25.23 ppm | 389.66 ppm | +- | 8.45 ppm | 586.46 ppb | +- | 244.45 ppb | <495.16 ppm
[6 empty] | (Bq) | 10.3229 | +- | 0.18548 | 12.0087 | +- | 0.62923 |  | Average: |  | 10.4577 | +- | 0.17791
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 71438000 | +- | 6386000 | <472.00 |  |  | <461.00 |  |  | 7337 | +- | 4053
[6 empty] | (ppb or ppm) | 5786.48 ppm | +- | 517.27 ppm | [6 empty] | 1804.9 ppb | +- | 997.04 ppb
[6 empty] | (Bq) | 146.1621 | +- | 13.0658
SNOLAB V12 | Vial cap is upwards | 3.180 g | 2.958 | 200127 | 2020-01-27 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-T2 nEXO Calibration Sample | Specific Activity: 3.25 +- 0.07 Bq of 232Th / [g RGTh-1] |  |  |  | 2020-01-30 | (mBq/kg) | 50670 | +- | 2986 | 787600 | +- | 99220 | 3309 | +- | 723.5 | 3087000 | +- | 72390 | 224340 | +- | 18380 | <1361.00 |  |  | <610.70
| [10.11 – 10.56] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 4.1 ppm | +- | 0.24 ppm | 63.8 ppm | +- | 8.04 ppm | 5.82 ppm | +- | 1.27 ppm | 759.4 ppm | +- | 17.81 ppm | 7246.18 ppm | +- | 593.67 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 10.2704 | +- | 0.24084
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <453000.00 |  |  | 52273 | +- | 10520 | 27111 | +- | 2143 | 2936000 | +- | 67660
[6 empty] | (ppb or ppm) | <36.69 ppm | [8 empty] | 722.26 ppm | +- | 16.64 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 9.7681 | +- | 0.2251 |  | Average: |  | 10.0023 | +- | 0.16446
SNOLAB V13 | Vial cap is upwards | 0.626 g | 4.91 | 200130 | 2020-01-30 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-K1 nEXO Calibration Sample | Specific Activity: 14.0 +- 0.09 Bq of 40K / [g RGK-1] |  |  |  | 2020-02-04 | (mBq/kg) | 2371 | +- | 1068 | 68810 | +- | 33160 | <311.40 |  |  | <1200.00 |  |  | 15444000 | +- | 836630 | <1641.00 |  |  | 104.3 | +- | 1429
| [ 8.70 – 8.82] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 192.05 ppb | +- | 86.51 ppb | 5.57 ppm | +- | 2.69 ppm | <0.55 ppm |  |  | <295.2 ppb |  |  | 498841.2 ppm | +- | 27023.15 ppm
[6 empty] | (Bq) | [12 empty] | 9.6679 | +- | 0.52373
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <562400.00 |  |  | <928.20 |  |  | <933.30 |  |  | <4932.00
[6 empty] | (ppb or ppm) | <45.55 ppm | [8 empty] | <1213.27 ppb
SNOLAB V14 | Part No: FG-p | 63.9 g | 6.021 | 200204 | 2020-02-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Plastic Process Equipment Inc. Food Grade Anti-Seize Thread Compound | Macedonia, OH, 44056 |  |  |  | 2020-02-10 | (mBq/kg) | 1519 | +- | 48.29 | 1630 | +- | 268.8 | 75.37 | +- | 7.183 | 601.7 | +- | 30.42 | 189.63 | +- | 118.3 | 7.4677 | +- | 8.364 | <4.27
Pipe Sealant Fortified with PTFE | [5 empty] | (ppb or ppm) | 123.04 ppb | +- | 3.91 ppb | 132.03 ppb | +- | 21.77 ppb | 132.65 ppb | +- | 12.64 ppb | 148.02 ppb | +- | 7.48 ppb | 6.13 ppm | +- | 3.82 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 46194 | +- | 25930 | <70.64 |  |  | <9.71 |  |  | 430.3 | +- | 38.22
[6 empty] | (ppb or ppm) | 3741.71 ppb | +- | 2100.33 ppb | [6 empty] | 105.85 ppb | +- | 9.4 ppb
SNOLAB V15 | Type: 801-16-BLU | 357.4 g | 2.819 | 200210 | 2020-02-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Push Lok Plus Muli-Purpose Hose | Parker Hannifin Corp |  |  |  | 2020-02-13 | (mBq/kg) | 1803 | +- | 51.43 | 5141 | +- | 534.4 | 106 | +- | 8.728 | 3102 | +- | 91.69 | 9333.1 | +- | 568.8 | 5.009 | +- | 9.986 | <5.85
[6 empty] | (ppb or ppm) | 146.04 ppb | +- | 4.17 ppb | 416.42 ppb | +- | 43.29 ppb | 186.56 ppb | +- | 15.36 ppb | 763.09 ppb | +- | 22.56 ppb | 301.46 ppm | +- | 18.37 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 5826 | +- | 7988 | <124.20 |  |  | 22.169 | +- | 10.2 | 3004 | +- | 101.9
[6 empty] | (ppb or ppm) | 471.91 ppb | +- | 647.03 ppb | [6 empty] | 738.98 ppb | +- | 25.07 ppb
SNOLAB V16 | ¾” Natural polypropylene | 623.5 g | 7.646 | 200213 200214 | 2020-02-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
ENPURE Polypropylene Tubing | Type: SCH-80 Natural |  |  |  | 2020-02-21 | (mBq/kg) | 6.106 | +- | 1.706 | 17.12 | +- | 23.48 | <1.34 |  |  | 12.55 | +- | 1.969 | 4.02 | +- | 15.06 | <1.20 |  |  | 1.021 | +- | 0.5066
| Lot: JJ05-B-12021321352 |  |  |  |  | (ppb or ppm) | 0.49 ppb | +- | 0.14 ppb | 1.39 ppb | +- | 1.9 ppb | <2.36 ppb |  |  | 3.09 ppb | +- | 0.48 ppb | 129.85 ppb | +- | 486.44 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1062.5 | +- | 569.2 | <6.47 |  |  | 0.87879 | +- | 0.6374 | <3.34
[6 empty] | (ppb or ppm) | 86.06 ppb | +- | 46.11 ppb | [6 empty] | <0.82 ppb
SNOLAB V17 | Clear Non-Toxic PVC with Polyester Braid and Rigid PVC Helix | 240.5 g | 3.847 | 200221 | 2020-02-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Reinforced FDA PVC Suction Hose |  |  |  |  | 2020-02-25 | (mBq/kg) | 494.3 | +- | 19.11 | 739.3 | +- | 127.2 | 26.56 | +- | 3.282 | 60.84 | +- | 8.144 | 678.07 | +- | 100.8 | <3.93 |  |  | <3.08
[6 empty] | (ppb or ppm) | 40.04 ppb | +- | 1.55 ppb | 59.88 ppb | +- | 10.3 ppb | 46.75 ppb | +- | 5.78 ppb | 14.97 ppb | +- | 2 ppb | 21.9 ppm | +- | 3.26 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2178.2 | +- | 1735 | 60.512 | +- | 28.93 | <2.38 |  |  | 48.45 | +- | 12.81
[6 empty] | (ppb or ppm) | 176.43 ppb | +- | 140.54 ppb | [6 empty] | 11.92 ppb | +- | 3.15 ppb
SNOLAB V18 | NIST Sample | 28.825 g | 12.736 | 200225 200302 | 2020-02-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Powdered Ocean Sediment Calibration Sample | Cross Comparison Sample A |  |  |  | 2020-03-09 | (mBq/kg) | 11640 | +- | 255.7 | 21440 | +- | 2071 | 588.4 | +- | 23 | 12650 | +- | 331.7 | 255240 | +- | 13160 | 7109 | +- | 366.4 | <29.11
| Sample placed directly on top of the detector |  |  |  |  | (ppb or ppm) | 942.84 ppb | +- | 20.71 ppb | 1736.64 ppb | +- | 167.75 ppb | 1035.58 ppb | +- | 40.48 ppb | 3.11 ppm | +- | 0.08 ppm | 8244.25 ppm | +- | 425.07 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 241Am
[6 empty] | (mBq/kg) | <15060.00 |  |  | <452.60 |  |  | 91.497 | +- | 29.67 | 12730 | +- | 352 | 20580 | +- | 1569
[6 empty] | (ppb or ppm) | <1219.86 ppb | [8 empty] | 3.13 ppm | +- | 0.09 ppm
SNOLAB V19 | NIST Sample | 33.454 g | 21.617 | 200519 | 2020-05-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Powdered Ocean Sediment Calibration Sample | Cross Comparison Sample B |  |  |  | 2020-06-10 | (mBq/kg) | 10980 | +- | 223 | 17630 | +- | 1702 | 571.6 | +- | 18.5 | 12060 | +- | 299.1 | 247900 | +- | 12720 | 6397.8 | +- | 352.7 | <25.82
| Sample placed directly on top of the detector |  |  |  |  | (ppb or ppm) | 889.38 ppb | +- | 18.06 ppb | 1428.03 ppb | +- | 137.86 ppb | 1006.02 ppb | +- | 32.56 ppb | 2.97 ppm | +- | 0.07 ppm | 8007.17 ppm | +- | 410.86 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 241Am
[6 empty] | (mBq/kg) | <8362.00 |  |  | <272.30 |  |  | 74.544 | +- | 20.54 | 12000 | +- | 307.1 | 28410 | +- | 1509
[6 empty] | (ppb or ppm) | <677.32 ppb | [8 empty] | 2.95 ppm | +- | 0.08 ppm
SNOLAB V20 | Lake in City of Greater Sudbury, Ontario | 993.9 g | 14.666 | 200309 | 2020-03-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Long Lake Water | Sample collected on 2019-10-14 at 16:54 |  |  |  | 2020-03-24 | (mBq/kg) | 2.151 | +- | 0.5584 | <12.40 |  |  | <0.27 |  |  | <0.61 |  |  | 52.142 | +- | 8.263 | 1.003 | +- | 0.383 | <0.15
| Sample placed in beaker “∆” |  |  |  |  | (ppb or ppm) | 174.23 ppt | +- | 45.23 ppt | <1 ppb |  |  | <0.48 ppb |  |  | <0.15 ppb |  |  | 1684.19 ppb | +- | 266.89 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1151.5 | +- | 456.7 | 6.8076 | +- | 2.466 | <0.38 |  |  | 1.243 | +- | 0.8064
[6 empty] | (ppb or ppm) | 93.27 ppb | +- | 36.99 ppb | [6 empty] | 0.31 ppb | +- | 0.2 ppb
SNOLAB V21 | Lake in City of Greater Sudbury, Ontario | 981.4 g | 17.117 | 200610 200624 | 2020-06-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Simon Lake Water | Sample collected on 2019-11-01 at 18:15 |  |  |  | 2020-06-30 | (mBq/kg) | 1.599 | +- | 0.506 | 12.4 | +- | 13.22 | <0.44 |  |  | <0.82 |  |  | 402 | +- | 25.3 | <0.46 |  |  | 0.1133 | +- | 0.2314
| Sample placed in beaker “H” |  |  |  |  | (ppb or ppm) | 129.52 ppt | +- | 40.99 ppt | 1004.4 ppt | +- | 1070.82 ppt | <0.77 ppb |  |  | <0.2 ppb |  |  | 12984.6 ppb | +- | 817.19 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1144.6 | +- | 434.2 | <2.80 |  |  | <0.27 |  |  | 4.993 | +- | 1.066
[6 empty] | (ppb or ppm) | 92.71 ppb | +- | 35.17 ppb | [6 empty] | 1.23 ppb | +- | 0.26 ppb
SNOLAB V22 | Lake in City of Greater Sudbury, Ontario | 935.2 g | 21.708 | 200630 | 2020-06-30 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Robinson Lake Water | Sample collected on 2019-10-22 at 12:43 |  |  |  | 2020-07-22 | (mBq/kg) | 5.117 | +- | 0.5798 | 21.36 | +- | 12.7 | <0.33 |  |  | 0.3155 | +- | 0.4927 | 95.614 | +- | 9.55 | 0.23 | +- | 0.2964 | <0.14
| Sample placed in beaker “□” |  |  |  |  | (ppb or ppm) | 414.48 ppt | +- | 46.96 ppt | 1.73 ppb | +- | 1.03 ppb | <0.58 ppb |  |  | 0.08 ppb | +- | 0.12 ppb | 3088.33 ppb | +- | 308.47 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 536.35 | +- | 394.1 | 0.5047 | +- | 2.041 | 0.23 | +- | 0.2964 | 1.943 | +- | 0.7858
[6 empty] | (ppb or ppm) | 43.44 ppb | +- | 31.92 ppb | [6 empty] | 0.48 ppb | +- | 0.19 ppb
SNOLAB V23 | Lake in City of Greater Sudbury, Ontario | 986.0 g | 16.659 | 200722 | 2020-07-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Meatbird Lake Water | Sample collected on 2019-10-08 at 10:50 |  |  |  | 2020-08-11 | (mBq/kg) | 0.19 | +- | 0.4465 | <13.83 |  |  | 0.3666 | +- | 0.2156 | <0.32 |  |  | 124.59 | +- | 11.09 | <0.39 |  |  | <0.21
| Sample placed in beaker “∆” |  |  |  |  | (ppb or ppm) | 15.39 ppt | +- | 36.17 ppt | <1.12 ppb |  |  | 0.65 ppb | +- | 0.38 ppb | <0.08 ppb |  |  | 4024.26 ppb | +- | 358.21 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <567.10 |  |  | 0.975 | +- | 1.909 | 0.02689 | +- | 0.1934 | 0.792 | +- | 0.7894
[6 empty] | (ppb or ppm) | <45.94 ppb | [8 empty] | 0.19 ppb | +- | 0.19 ppb
SNOLAB V24 | Lake in City of Greater Sudbury, Ontario | 842.7 g | 13.5 | 200811 20081101 | 2020-08-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Kelly Lake Water | Sample collected on 2019-10-7 at 10:30 |  |  |  | 2020-08-25 | (mBq/kg) | 3.308 | +- | 0.7805 | 11.09 | +- | 30.51 | <0.60 |  |  | 0.3551 | +- | 0.7696 | 501.76 | +- | 33.74 | 0.5963 | +- | 0.4861 | <0.24
[6 empty] | (ppb or ppm) | 267.95 ppt | +- | 63.22 ppt | 898.29 ppt | +- | 2471.31 ppt | <1.06 ppb |  |  | 0.09 ppb | +- | 0.19 ppb | 16.21 ppm | +- | 1.09 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2300.5 | +- | 1355 | 2.4308 | +- | 3.233 | <0.43 |  |  | 6.605 | +- | 1.395
[6 empty] | (ppb or ppm) | 186.34 ppb | +- | 109.76 ppb | [6 empty] | 1.62 ppb | +- | 0.34 ppb
SNOLAB V25 | Lake in City of Greater Sudbury, Ontario | 1032.2 g | 13.395 | 201006 | 2020-10-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Robinson Lake Water | Sample collected on 2020-09-19 at 14:30 |  |  |  | 2020-10-20 | (mBq/kg) | 1.353 | +- | 0.5589 | <29.58 |  |  | 0.417 | +- | 0.2379 | <1.04 |  |  | 100.12 | +- | 11.32 | 0.4959 | +- | 0.3941 | <0.16
[6 empty] | (ppb or ppm) | 109.59 ppt | +- | 45.27 ppt | <2.4 ppb |  |  | 0.73 ppb | +- | 0.42 ppb | <0.26 ppb |  |  | 3233.88 ppb | +- | 365.64 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 881.22 | +- | 989.9 | 10.405 | +- | 3.225 | 0.33892 | +- | 0.2337 | 1.877 | +- | 0.9873
[6 empty] | (ppb or ppm) | 71.38 ppb | +- | 80.18 ppb | [6 empty] | 0.46 ppb | +- | 0.24 ppb
SNOLAB V26 | Lake in City of Greater Sudbury, Ontario | 985.3 g | 10.743 | 201023 | 2020-10-23 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
St. Charles Lake Water | Sample collected on 2020-09-19 at 14:53 |  |  |  | 2020-11-03 | (mBq/kg) | 1.5 | +- | 0.62 | 27.39 | +- | 23.77 | 0.279 | +- | 0.2828 | 0.3211 | +- | 0.6553 | 56.582 | +- | 9.839 | 1.2519 | +- | 0.4744 | 0.1718 | +- | 0.2352
[6 empty] | (ppb or ppm) | 0.12 ppb | +- | 0.05 ppb | 2.22 ppb | +- | 1.93 ppb | 0.49 ppb | +- | 0.5 ppb | 0.08 ppb | +- | 0.16 ppb | 1.83 ppm | +- | 0.32 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 23258 | +- | 29210 | 13.734 | +- | 3.761 | <0.24 |  |  | 0.159 | +- | 0.9299
[6 empty] | (ppb or ppm) | 1883.9 ppb | +- | 2366.01 ppb | [6 empty] | 0.04 ppb | +- | 0.23 ppb
SNOLAB V27 | Lake in City of Greater Sudbury, Ontario | 956.4 g | 11.146 | 201106 | 2020-11-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Long Lake Water | Sample collected on 2020-09-27 at 15:00 |  |  |  | 2020-11-17 | (mBq/kg) | <0.98 |  |  | <16.85 |  |  | 0.4023 | +- | 0.2632 | <0.63 |  |  | 35.529 | +- | 8.902 | 0.61398 | +- | 0.4425 | <0.31
[6 empty] | (ppb or ppm) | <79.38 ppt |  |  | <1.36 ppb |  |  | 0.71 ppb | +- | 0.46 ppb | <0.15 ppb |  |  | 1147.59 ppb | +- | 287.53 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1542.4 | +- | 1182 | 14.09 | +- | 3.542 | <0.39 |  |  | <0.84
[6 empty] | (ppb or ppm) | 124.93 ppb | +- | 95.74 ppb | [6 empty] | <0.21 ppb
SNOLAB V28 | Lake Bed Sand from Lohi Lake in City of Greater Sudbury, Ontario | 600.4 g | 2.93 | 201020 | 2020-10-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Lohi Lake Sand | Sample collected on 2020-09-19 at 17:10 |  |  |  | 2020-10-23 | (mBq/kg) | 7917 | +- | 172.9 | 17980 | +- | 1656 | 422.1 | +- | 18.84 | 6592 | +- | 179.8 | 381970 | +- | 19550 | N/D – No observable signal due to U & Th Interference | +- | 508.4 | <27.73
[6 empty] | (ppb or ppm) | 641.28 ppb | +- | 14 ppb | 1456.38 ppb | +- | 134.14 ppb | 742.9 ppb | +- | 33.16 ppb | 1621.63 ppb | +- | 44.23 ppb | 12337.63 ppm | +- | 631.47 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 455740 | +- | 198800 | N/D – No observable signal due to U & Th Interference | +- | 710.5 | N/D – No observable signal due to U & Th Interference | +- | 18.84 | 6194 | +- | 190.9
[6 empty] | (ppb or ppm) | 36914.94 ppb | +- | 16102.8 ppb | [6 empty] | 1523.72 ppb | +- | 46.96 ppb
SNOLAB V29 | Lake Bed Sand from Ramsey Lake in City of Greater Sudbury, Ontario | 605.1 g | 2.93 | 201103 | 2020-11-03 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Ramsey Lake Sand | Sample collected on 2020-09-15 at 18:50 |  |  |  | 2020-11-06 | (mBq/kg) | 14950 | +- | 305.7 | 28000 | +- | 2515 | 765.4 | +- | 27.2 | 16230 | +- | 407.2 | 513760 | +- | 26280 | N/D – No observable signal due to U & Th Interference |  |  | <34.32
[6 empty] | (ppb or ppm) | 1210.95 ppb | +- | 24.76 ppb | 2268 ppb | +- | 203.72 ppb | 1347.1 ppb | +- | 47.87 ppb | 3992.58 ppb | +- | 100.17 ppb | 16594.45 ppm | +- | 848.84 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 156490 | +- | 184600 | N/D – No observable signal due to U & Th Interference |  |  | N/D – No observable signal due to U & Th Interference | +- | 24.88 | 14800 | +- | 392.3
[6 empty] | (ppb or ppm) | 12675.69 ppb | +- | 14952.6 ppb | [6 empty] | 3640.8 ppb | +- | 96.51 ppb
SNOLAB V30 | Lake Bed Sand from St Charles Lake in City of Greater Sudbury, Ontario | 591.6 g | 2.534 | 201117 | 2020-11-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
St Charles Lake Sand | Sample collected on 2020-09-19 at 14:55 |  |  |  | 2020-11-20 | (mBq/kg) | 12370 | +- | 263.6 | 29950 | +- | 2642 | 606.5 | +- | 25.82 | 16700 | +- | 422.7 | 572120 | +- | 29270 | N/D – No observable signal due to U & Th Interference |  |  | 17.15 | +- | 25.66
[6 empty] | (ppb or ppm) | 1001.97 ppb | +- | 21.35 ppb | 2425.95 ppb | +- | 214 ppb | 1067.44 ppb | +- | 45.44 ppb | 4108.2 ppb | +- | 103.98 ppb | 18479.48 ppm | +- | 945.42 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <266100.00 |  |  | N/D – No observable signal due to U & Th Interference |  |  | N/D – No observable signal due to U & Th Interference | +- | 30.21 | 16120 | +- | 429.6
[6 empty] | (ppb or ppm) | <21.55 ppm | [8 empty] | 3965.52 ppb | +- | 105.68 ppb
SNOLAB V31 | Lake in City of Greater Sudbury, Ontario | 1023.10 g | 13.687 | 201120 | 2020-11-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Lake Water | Sample collected on 2020-10-23 at 13:00 |  |  |  | 2020-12-04 | (mBq/kg) | 2.203 | +- | 0.5796 | 15.37 | +- | 20.24 | <0.42 |  |  | 0.9665 | +- | 0.6031 | 63.512 | +- | 9.205 | <0.38 |  |  | <0.22
[6 empty] | (ppb or ppm) | 178.44 ppt | +- | 46.95 ppt | 1.24 ppb | +- | 1.64 ppb | <0.74 ppb |  |  | 0.24 ppb | +- | 0.15 ppb | 2051.44 ppb | +- | 297.32 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <1381.00 |  |  | 57.033 | +- | 5.721 | <0.24 |  |  | 3.169 | +- | 0.9504
[6 empty] | (ppb or ppm) | <111.86 ppb | [8 empty] | 0.78 ppb | +- | 0.23 ppb
SNOLAB V32 | Lake in City of Greater Sudbury, Ontario | 961.0 g | 13.777 | 201204 | 2020-12-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Meatbird Lake Water | Sample collected on October 05, 2020 @ 14:00 |  |  |  | 2020-12-18 | (mBq/kg) | 0.8306 | +- | 0.5845 | <24.71 |  |  | <0.30 |  |  | 0.064 | +- | 0.5975 | 167.44 | +- | 15.2 | 0.18371 | +- | 0.3529 | 0.2729 | +- | 0.2459
[6 empty] | (ppb or ppm) | 67.28 ppt | +- | 47.34 ppt | <2 ppb |  |  | <0.53 ppb |  |  | 0.02 ppb | +- | 0.15 ppb | 5.41 ppm | +- | 0.49 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 977.51 | +- | 1067 | 5.128 | +- | 2.804 | 0.29759 | +- | 0.2745 | 4.56 | +- | 1.18
[6 empty] | (ppb or ppm) | 79.18 ppb | +- | 86.43 ppb | [6 empty] | 1.12 ppb | +- | 0.29 ppb
SNOLAB V33 | Lucas Cell Coating | 68.3 g | 23.652 | 201218 | 2020-12-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
ZnS(Ag) Powder | Saint-Gobain |  |  | The sample is also counted in the Well detector as SNOLAB CW10 | 2021-01-11 | (mBq/kg) | <3.30 |  |  | <98.38 |  |  | 0.8626 | +- | 0.9022 | <4.37 |  |  | 21.744 | +- | 34.92 | <3.09 |  |  | <0.89
[6 empty] | (ppb or ppm) | <0.27 ppb |  |  | <7.97 ppb |  |  | 1.52 ppb | +- | 1.59 ppb | <1.08 ppb |  |  | 702.33 ppb | +- | 1127.92 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 53891 | +- | 34430 | 5.0645 | +- | 11.94 | <1.33 |  |  | <5.33
[6 empty] | (ppb or ppm) | 4.37 ppm | +- | 2.79 ppm | [6 empty] | <1.31 ppb
SNOLAB V34 | 72% Cocoa | 202.9 g | 4.98 | 240425 | 2024-04-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Merci (Stork) Dark Chocolate | Made in Germany |  |  |  | 2024-04-30 | (mBq/kg) | 453.3 | +- | 23.5 | 952.5 | +- | 379.5 | 5.613 | +- | 6.602 | 285.7 | +- | 21.85 | 20507 | +- | 10540 | 62.385 | +- | 14.87 | <19.40
[6 empty] | (ppb or ppm) | 36.72 ppb | +- | 1.9 ppb | 77.15 ppb | +- | 30.74 ppb | 9.88 ppb | +- | 11.62 ppb | 70.28 ppb | +- | 5.38 ppb | 662.38 ppm | +- | 340.44 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <29620.00 |  |  | 227.13 | +- | 95.34 | 22.85 | +- | 14.15 | 630.1 | +- | 51.65
[6 empty] | (ppb or ppm) | <2.4 ppm | [8 empty] | 155 ppb | +- | 12.71 ppb
SNOLAB V35 | Medium Grey | 107.4 g | 6.894 | 240430 | 2024-04-30 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Sikafloor 22NA PurCem | Type: SS-0721 |  |  |  | 2024-05-07 | (mBq/kg) | 5466 | +- | 126.7 | 15130 | +- | 1128 | 251.3 | +- | 12.68 | 6028 | +- | 165.6 | 347280 | +- | 17800 | 101.34 | +- | 22.99 | <29.50
[6 empty] | (ppb or ppm) | 442.75 ppb | +- | 10.26 ppb | 1225.53 ppb | +- | 91.37 ppb | 442.29 ppb | +- | 22.32 ppb | 1.48 ppm | +- | 0.04 ppm | 11217.14 ppm | +- | 574.94 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 43451 | +- | 41950 | <207.70 |  |  | 52.065 | +- | 22.07 | 5739 | +- | 173.9
[6 empty] | (ppb or ppm) | 3519.53 ppb | +- | 3397.95 ppb | [6 empty] | 1.41 ppm | +- | 0.04 ppm
SNOLAB V36 |  | 200.5 g | 2.934 | 240507 | 2024-05-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Cadbury Dark Chocolate | Made in Canada |  |  |  | 2024-05-10 | (mBq/kg) | 373.7 | +- | 23.06 | 809.9 | +- | 372.5 | 14.04 | +- | 6.638 | 194.7 | +- | 20.7 | 105060 | +- | 5483 | 39.007 | +- | 14.36 | 7.814 | +- | 15.99
[6 empty] | (ppb or ppm) | 30.27 ppb | +- | 1.87 ppb | 65.6 ppb | +- | 30.17 ppb | 24.71 ppb | +- | 11.68 ppb | 47.9 ppb | +- | 5.09 ppb | 3393.44 ppm | +- | 177.1 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <33080.00 |  |  | <157.70 |  |  | <27.03 |  |  | 370.4 | +- | 47.45
[6 empty] | (ppb or ppm) | <2.68 ppm | [8 empty] | 91.12 ppb | +- | 11.67 ppb
PICO Measurements:
Sample Description |  | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
PICO V01 | Model No. S1328PD3T72 Omega | 60.1 g | 15.659 | 200825 | 2020-08-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
White RTD Wire Leads |  |  |  |  | 2020-09-10 | (mBq/kg) | 5.375 | +- | 7.215 | <127.80 |  |  | 2.542 | +- | 2.395 | <9.16 |  |  | 69.281 | +- | 72.04 | <5.89 |  |  | <1.78
[6 empty] | (ppb or ppm) | 0.44 ppb | +- | 0.58 ppb | <10.35 ppb |  |  | 4.47 ppb | +- | 4.22 ppb | <2.25 ppb |  |  | 2237.78 ppb | +- | 2326.89 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 140700 | +- | 38520 | <53.17 |  |  | <1.90 |  |  | <14.98
[6 empty] | (ppb or ppm) | 11.4 ppm | +- | 3.12 ppm | [6 empty] | <3.69 ppb
PICO V02 | Model No. S1328PD3T72 Omega | 23.9 g | 11.639 | 200910 | 2020-09-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Red RTD Wire Leads |  |  |  |  | 2020-09-22 | (mBq/kg) | 4.669 | +- | 20.77 | <422.90 |  |  | 6.185 | +- | 6.586 | <27.98 |  |  | 74.988 | +- | 232.8 | <17.53 |  |  | <3.64
[6 empty] | (ppb or ppm) | 0.38 ppb | +- | 1.68 ppb | <34.25 ppb |  |  | 10.89 ppb | +- | 11.59 ppb | <6.88 ppb |  |  | 2.42 ppm | +- | 7.52 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 148160 | +- | 69770 | 132.14 | +- | 84.27 | <14.75 |  |  | <42.97
[6 empty] | (ppb or ppm) | 12 ppm | +- | 5.65 ppm | [6 empty] | <10.57 ppb
PICO V03 | Hard Copper Wire | 45.2 g | 13.687 | 200922 | 2020-09-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
PICO Heater Cables |  |  |  |  | 2020-10-06 | (mBq/kg) | <13.25 |  |  | <127.70 |  |  | <3.89 |  |  | 6.695 | +- | 9.819 | <148.60 |  |  | 9.1666 | +- | 6.88 | <5.33
[6 empty] | (ppb or ppm) | <1.07 ppb |  |  | <10.34 ppb |  |  | <6.85 ppb |  |  | 1.65 ppb | +- | 2.42 ppb | <4.8 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 58Co |  |  | 57Co
[6 empty] | (mBq/kg) | 143110 | +- | 62600 | 3.3293 | +- | 41.87 | 2.5467 | +- | 4.666 | 21.23 | +- | 16.07 | 0.54315 | +- | 3.654 | 7.049 | +- | 6.27
[6 empty] | (ppb or ppm) | 11.59 ppm | +- | 5.07 ppm | [6 empty] | 5.22 ppb | +- | 3.95 ppb
PICO V04 | Preamps for Piezos for the PICO and SBC experiments | 21.2 g | 25.61 | 210421 210429 | 2021-04-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Kapton Preamp Boards |  | Quantity: 28 |  |  | 2021-05-17 | (mBq/kg) | 1504 | +- | 45.52 | 4544 | +- | 459.1 | 89.11 | +- | 7.122 | 2226 | +- | 71.56 | 4239.8 | +- | 321.9 | <9.34 |  |  | <5.98
[6 empty] | (ppb or ppm) | 121.82 ppb | +- | 3.69 ppb | 368.06 ppb | +- | 37.19 ppb | 156.83 ppb | +- | 12.53 ppb | 547.6 ppb | +- | 17.6 ppb | 136.95 ppm | +- | 10.4 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 259030 | +- | 88590 | <108.00 |  |  | 12.595 | +- | 9.45 | 2254 | +- | 80.58
[6 empty] | (ppb or ppm) | 20.98 ppm | +- | 7.18 ppm | [6 empty] | 554.48 ppb | +- | 19.82 ppb
PICO V05 | BLU-25751706 Lot 230328 3/32 “ x 7” | 276.2 g | 5.088 | 231026 23102601 231028 231029 | 2023-10-26 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  | 231030 | 2023-10-31 | (mBq/kg) | <186.40 |  |  | 9577 | +- | 4406 | <193.20 |  |  | 570400 | +- | 13220 | 14541 | +- | 902.7 | 61.79 | +- | 71.9 | <50.83
| Sample contains 20 electrodes |  |  |  |  | (ppb or ppm) | <15.1 ppb |  |  | 775.74 ppb | +- | 356.89 ppb | <340.03 ppb |  |  | 140.32 ppm | +- | 3.25 ppm | 469.67 ppm | +- | 29.16 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <2.62E6 |  |  | 2903 | +- | 633.1 | 2985.6 | +- | 182.1 | 207300 | +- | 4975
[6 empty] | (ppb or ppm) | <212.22 ppm | [8 empty] | 51 ppm | +- | 1.22 ppm
PICO V06 | BLU-25751706 Lot 220713 3/32 “ x 7” | 422.2 g | 4.687 | 231031 23103101 231101 231102 231103 | 2023-10-31 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  |  | 2023-11-05 | (mBq/kg) | <68.47 |  |  | 927.8 | +- | 1908 | 64.68 | +- | 53.28 | 93870 | +- | 2196 | 7938.6 | +- | 475.8 | <66.87 |  |  | <10.82
| Sample contains 30 electrodes |  |  |  |  | (ppb or ppm) | <5.55 ppb |  |  | 75.15 ppb | +- | 154.55 ppb | 113.84 ppb | +- | 93.77 ppb | 23.09 ppm | +- | 0.54 ppm | 256.42 ppm | +- | 15.37 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2324000 | +- | 2368000 | 1866.6 | +- | 317.7 | 1527 | +- | 84.04 | 105300 | +- | 2557
[6 empty] | (ppb or ppm) | 188.24 ppm | +- | 191.81 ppm | [6 empty] | 25.9 ppm | +- | 0.63 ppm
PICO V07 | BLU-25751708 Lot 230204 1/8 “ x 7” | 251.5 g | 3.126 | 231105 | 2023-11-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  |  | 2023-11-08 | (mBq/kg) | <115.70 |  |  | <7092.00 |  |  | <235.50 |  |  | 577200 | +- | 13410 | 16792 | +- | 1114 | <111.70 |  |  | <59.50
| Sample contains 10 electrodes |  |  |  |  | (ppb or ppm) | <9.37 ppb |  |  | <574.45 ppb |  |  | <414.48 ppb |  |  | 141.99 ppm | +- | 3.3 ppm | 542.38 ppm | +- | 35.98 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <7.92E6 |  |  | 3190.9 | +- | 949.6 | 3172.3 | +- | 181.6 | 222400 | +- | 5483
[6 empty] | (ppb or ppm) | <641.52 ppm | [8 empty] | 54.71 ppm | +- | 1.35 ppm
PICO V08 | BLU-25751708 Lot 230515 1/8 “ x 7” | 1004.4 g | 7.786 | 231108 | 2023-11-08 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Tungsten Ceriated Welding Electrodes | Air Liquide, Blueshield, Made in China |  |  |  | 2023-11-16 | (mBq/kg) | <33.37 |  |  | 263.4 | +- | 2273 | 299.5 | +- | 113.6 | 599300 | +- | 14060 | 17403 | +- | 929.6 | <59.25 |  |  | <14.68
| Sample contains 40 electrodes |  |  |  |  | (ppb or ppm) | <2.7 ppb |  |  | 21.34 ppb | +- | 184.11 ppb | 527.12 ppb | +- | 199.94 ppb | 147.43 ppm | +- | 3.46 ppm | 562.12 ppm | +- | 30.03 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <1.79E7 |  |  | 3007.2 | +- | 395 | 2659.2 | +- | 146.1 | 213800 | +- | 5309
[6 empty] | (ppb or ppm) | <1449.9 ppm | [8 empty] | 52.59 ppm | +- | 1.31 ppm
PICO V09 | Molex wire ring connectors | 100.6 g | 8.757 | 240305 240313 | 2024-03-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Wire Ring connectors | Types: 16-14 MX 22-18 MX |  |  |  | 2024-03-15 | (mBq/kg) | <4.41 |  |  | 99.35 | +- | 249 | <4.27 |  |  | <5.03 |  |  | <64.56 |  |  | <3.73 |  |  | <3.75
[6 empty] | (ppb or ppm) | <357.21 ppt |  |  | 8.05 ppb | +- | 20.17 ppb | <7.52 ppb |  |  | <1.24 ppb |  |  | <2.09 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 371410 | +- | 337700 | 46.919 | +- | 30.78 | 7.5284 | +- | 3.965 | <7.22
[6 empty] | (ppb or ppm) | 30.08 ppm | +- | 27.35 ppm | [6 empty] | <1.78 ppb
NEXO Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
3 samples have been counted for nEXO results are only available to nEXO collaboration members.
CUTE Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
CUTE V01 | CUTE Source SRS-19-003 | 194.0 g | 6.791 | 210120 | 2021-01-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
CUTE Fe-55 Calibration Sample | 55-Fe source inside a copper holder |  |  |  | 2021-01-27 | (mBq/kg) | <1.44 |  |  | <247.60 |  |  | 6.192 | +- | 3.943 | 7.232 | +- | 2.439 | 7.669 | +- | 20.01 | <1.56 |  |  | 175.3 | +- | 8.401
| The source strength is 222 kBq |  |  |  |  | (ppb or ppm) | <0.12 ppb |  |  | <20.06 ppb |  |  | 10.9 ppb | +- | 6.94 ppb | 1.78 ppb | +- | 0.6 ppb | 247.71 ppb | +- | 646.32 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | <639400.00 | High limit due to large number of 55-Fe X-rays |  | 5.8584 | +- | 9.093 | <2.18 |  |  | 4.19 | +- | 4.472 | <53.66 |  |  | <1.79 |  |  | + large number of X-Rays from 55-Fe
[6 empty] | (ppb or ppm) | <51.79 ppm | [8 empty] | 1.03 ppb | +- | 1.1 ppb
CUTE V02 | Loctite Epoxy by Henkel | 52.3 g | 5.861 | 210415 | 2021-04-15 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Loctite EA-0151 Epoxy | PN: 398454 |  |  |  | 2021-04-21 | (mBq/kg) | <8.65 |  |  | <64.10 |  |  | <3.01 |  |  | <9.01 |  |  | <78.76 |  |  | <4.56 |  |  | <6.37
| Batch No: LB0KAA3587 21/043 |  |  |  |  | (ppb or ppm) | <0.7 ppb |  |  | <5.19 ppb |  |  | <5.3 ppb |  |  | <2.22 ppb |  |  | <2.54 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 897.72 | +- | 981.5 | <74.36 |  |  | <6.39 |  |  | <27.55
[6 empty] | (ppb or ppm) | 72.72 ppb | +- | 79.5 ppb | [6 empty] | <6.78 ppb
CUTE V03 |  | 51.6 g | 7.883 | 240411 | 2024-04-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
G10 Plastic Tubes & Inserts |  |  |  |  | 2024-04-19 | (mBq/kg) | 8100 | +- | 184.2 | 25000 | +- | 1808 | 397.2 | +- | 19.02 | 20790 | +- | 512.1 | 16364 | +- | 950 | <26.87 |  |  | <14.42
[6 empty] | (ppb or ppm) | 656.1 ppb | +- | 14.92 ppb | 2.03 ppm | +- | 0.15 ppm | 699.07 ppb | +- | 33.48 ppb | 5.11 ppm | +- | 0.13 ppm | 528.56 ppm | +- | 30.69 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 30602 | +- | 83100 | 354.71 | +- | 202 | 236.28 | +- | 30.01 | 18730 | +- | 476.9
[6 empty] | (ppb or ppm) | 2.48 ppm | +- | 6.73 ppm | [6 empty] | 4.61 ppm | +- | 0.12 ppm
CUTE V04 |  | 16.9 g | 5.778 | 240419 240425 | 2024-04-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
G10 Plastic Tubes |  |  |  |  | 2024-04-25 | (mBq/kg) | 10710 | +- | 323.5 | 39930 | +- | 3861 | 527.3 | +- | 51.78 | 27360 | +- | 793.2 | 22372 | +- | 1778 | <86.57 |  |  | <78.20
[6 empty] | (ppb or ppm) | 867.51 ppb | +- | 26.2 ppb | 3.23 ppm | +- | 0.31 ppm | 928.05 ppb | +- | 91.13 ppb | 6.73 ppm | +- | 0.2 ppm | 722.62 ppm | +- | 57.43 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 100830 | +- | 253700 | <1163.00 |  |  | 415.95 | +- | 80.19 | 25420 | +- | 812.2
[6 empty] | (ppb or ppm) | 8.17 ppm | +- | 20.55 ppm | [6 empty] | 6.25 ppm | +- | 0.2 ppm
SBC Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
SBC V01 | Stainless Steel | 364.72 g | 8.777 | 210127 210202 | 2021-01-27 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Vessel Body |  |  |  | 2021-02-05 | (mBq/kg) | 1.909 | +- | 1.127 | <61.70 |  |  | 0.4688 | +- | 0.4291 | <2.64 |  |  | <23.78 |  |  | 0.70548 | +- | 0.6779 | 5.059 | +- | 0.8153
| PV Drawing Item #3 |  |  |  |  | (ppb or ppm) | 154.63 ppt | +- | 91.29 ppt | <5 ppb |  |  | 0.83 ppb | +- | 0.76 ppb | <0.65 ppb |  |  | <0.77 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 148260 | +- | 118900 | <7.53 |  |  | 2.8376 | +- | 0.7004 | <1.76 |  |  | 4.727 | +- | 1.904 | <0.86
[6 empty] | (ppb or ppm) | 12.01 ppm | +- | 9.63 ppm | [6 empty] | <0.43 ppb
SBC V02 | Stainless Steel | 254.35 g | 6.791 | 210205 200207 | 2021-02-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Feedthrough Tube for the Pressure Vessel |  |  |  | 2021-02-12 | (mBq/kg) | 3.86 | +- | 2.028 | <67.98 |  |  | 0.3404 | +- | 0.7203 | 4.659 | +- | 2.153 | <34.73 |  |  | 1.0881 | +- | 1.447 | 1.016 | +- | 0.8182
| PV Drawing Item #8 |  |  |  |  | (ppb or ppm) | 312.66 ppt | +- | 164.27 ppt | <5.51 ppb |  |  | 0.6 ppb | +- | 1.27 ppb | 1.15 ppb | +- | 0.53 ppb | <1.12 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 141470 | +- | 129400 | 7.4193 | +- | 8.756 | 2.3993 | +- | 1.235 | <7.34 |  |  | <2.44 |  |  | <1.17
[6 empty] | (ppb or ppm) | 11.46 ppm | +- | 10.48 ppm | [6 empty] | <1.81 ppb
SBC V03 | Stainless Steel | 189.45 g | 11.541 | 210212 210216 | 2021-02-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Vessel Spool |  |  |  | 2021-02-24 | (mBq/kg) | 0.9399 | +- | 1.763 | 126.8 | +- | 71.45 | 2.376 | +- | 0.7457 | <3.89 |  |  | <32.52 |  |  | 1.5692 | +- | 1.16 | 75.03 | +- | 4.169
| PV Drawing Item #6 |  |  |  |  | (ppb or ppm) | 0.08 ppb | +- | 0.14 ppb | 10.27 ppb | +- | 5.79 ppb | 4.18 ppb | +- | 1.31 ppb | <0.96 ppb |  |  | <1.05 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 122400 | +- | 106900 | <12.46 |  |  | 6.4552 | +- | 1.469 | 2.611 | +- | 3.374 | <3.72 |  |  | <2.38
[6 empty] | (ppb or ppm) | 9.91 ppm | +- | 8.66 ppm | [6 empty] | 0.64 ppb | +- | 0.83 ppb
SBC V04 | Stainless Steel | 134.65 g | 8.639 | 210224 210301 | 2021-02-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Vessel Dome |  |  |  | 2021-03-05 | (mBq/kg) | <3.93 |  |  | <108.50 |  |  | <0.74 |  |  | 4.873 | +- | 2.867 | <42.76 |  |  | <1.27 |  |  | <1.33
| PV Drawing Item #1 |  |  |  |  | (ppb or ppm) | <0.32 ppb |  |  | <8.79 ppb |  |  | <1.3 ppb |  |  | 1.2 ppb | +- | 0.71 ppb | <1.38 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 74615 | +- | 120200 | 3.12 | +- | 8.982 | 4.1339 | +- | 1.37 | 12.3 | +- | 4.232 | 3.749 | +- | 2.117 | <1.59
[6 empty] | (ppb or ppm) | 6.04 ppm | +- | 9.74 ppm | [6 empty] | 3.03 ppb | +- | 1.04 ppb
SBC V05 | Stainless Steel | 372.40 g | 14.722 | 210305 | 2021-03-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Relief Pipe |  |  |  | 2021-03-22 | (mBq/kg) | 3.971 | +- | 1.466 | 82.29 | +- | 53.49 | 0.8206 | +- | 0.6175 | 6.322 | +- | 1.587 | <5.66 |  |  | <0.80 |  |  | 3.23 | +- | 0.8053
[6 empty] | (ppb or ppm) | 0.32 ppb | +- | 0.12 ppb | 6.67 ppb | +- | 4.33 ppb | 1.44 ppb | +- | 1.09 ppb | 1.56 ppb | +- | 0.39 ppb | <0.18 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 130530 | +- | 136300 | <6.77 |  |  | 1.6401 | +- | 0.7656 | <2.38 |  |  | <2.28 |  |  | 0.54838 | +- | 0.6217
[6 empty] | (ppb or ppm) | 10.57 ppm | +- | 11.04 ppm | [6 empty] | <0.59 ppb
SBC V06 | Silver Plated Bolts | 744.9 g | 7.182 | 231228 | 2023-12-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Silter Plated Bolts | Label: S30400 THE |  |  |  | 2024-01-05 | (mBq/kg) | <1.40 |  |  | 26.97 | +- | 36.14 | 0.4047 | +- | 0.3929 | 3.417 | +- | 1.103 | <9.62 |  |  | <0.50 |  |  | 23.25 | +- | 1.558
| Units: 6 |  |  |  |  | (ppb or ppm) | <0.11 ppb |  |  | 2.18 ppb | +- | 2.93 ppb | 0.71 ppb | +- | 0.69 ppb | 0.84 ppb | +- | 0.27 ppb | <0.31 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 18039 | +- | 20200 | 10.836 | +- | 5.093 | 0.8271 | +- | 0.6252 | 2.577 | +- | 1.833
[6 empty] | (ppb or ppm) | 1.46 ppm | +- | 1.64 ppm | [6 empty] | 0.63 ppb | +- | 0.45 ppb
SBC V07 | PTFE | 48.5 g | 11.708 | 240105 | 2024-01-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
IA PTFE inner Tower |  |  |  |  | 2024-01-17 | (mBq/kg) | <4.70 |  |  | 181.1 | +- | 111.8 | <4.56 |  |  | <5.05 |  |  | <63.58 |  |  | <11.23 |  |  | <2.16
[6 empty] | (ppb or ppm) | <0.38 ppb |  |  | 14.67 ppb | +- | 9.06 ppb | <8.03 ppb |  |  | <1.24 ppb |  |  | <2.05 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 23168 | +- | 13680 | <40.04 |  |  | <8.61 |  |  | <9.43
[6 empty] | (ppb or ppm) | 1.88 ppm | +- | 1.11 ppm | [6 empty] | <2.32 ppb
SBC V08 |  | 1695.2 g | 7.854 | 240117 | 2024-01-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Bellows Flange | 8 Pieces |  |  |  | 2024-01-25 | (mBq/kg) | 4.943 | +- | 0.8595 | 15.12 | +- | 27.78 | 0.1853 | +- | 0.3833 | 12.42 | +- | 1.2 | <7.91 |  |  | <0.59 |  |  | 10.29 | +- | 0.7999
[6 empty] | (ppb or ppm) | 0.4 ppb | +- | 0.07 ppb | 1.22 ppb | +- | 2.25 ppb | 0.33 ppb | +- | 0.67 ppb | 3.06 ppb | +- | 0.3 ppb | <0.26 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 14818 | +- | 21610 | 5.0672 | +- | 3.814 | 0.83976 | +- | 0.4803 | 15.84 | +- | 1.965 | <1.30 |  |  | 0.97213 | +- | 0.4483
[6 empty] | (ppb or ppm) | 1.2 ppm | +- | 1.75 ppm | [6 empty] | 3.9 ppb | +- | 0.48 ppb
SBC V09 | Standard Stainles Steel Tubes which are welded together | 150.8 g | 6.73 | 240209 240210 24021001 | 2024-02-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Welded Tubes |  |  |  |  | 2024-02-16 | (mBq/kg) | <1.61 |  |  | <73.78 |  |  | <0.93 |  |  | 4.495 | +- | 2.51 | <21.85 |  |  | <1.57 |  |  | 6.462 | +- | 1.508
[6 empty] | (ppb or ppm) | <0.13 ppb |  |  | <5.98 ppb |  |  | <1.64 ppb |  |  | 1.11 ppb | +- | 0.62 ppb | <705.76 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | <14390.00 |  |  | <12.37 |  |  | 2.8532 | +- | 1.386 | 2.074 | +- | 3.812 | <1.65 |  |  | <2.51
[6 empty] | (ppb or ppm) | <1.17 ppm | [8 empty] | 0.51 ppb | +- | 0.94 ppb
PICO V10 |  | 383.0 g | 10.291 | 240920 24092001 240830 | 2024-09-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Plate |  |  |  |  | 2024-10-01 | (mBq/kg) | 0.5482 | +- | 0.9116 | <27.33 |  |  | <0.64 |  |  | 1.471 | +- | 1.076 | <06.69 |  |  | <0.68 |  |  | 6.163 | +- | 0.7925
[6 empty] | (ppb or ppm) | 44.4 ppt | +- | 73.84 ppt | <2.21 ppb |  |  | <1.13 ppb |  |  | 0.36 ppb | +- | 0.26 ppb | <216.09 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 4459.9 | +- | 23260 | 0.4475 | +- | 3.544 | 2.3468 | +- | 0.6015 | 3.2 | +- | 1.638 | 2.426 | +- | 0.7925 | 0.40873 | +- | 0.4017
[6 empty] | (ppb or ppm) | 361.25 ppb | +- | 1884.06 ppb | [6 empty] | 0.79 ppb | +- | 0.4 ppb
PICO V11 |  | 763.7 g | 7.554 | 241001 241002 24100201 | 2024-10-01 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Welded Stainless Steel Plate |  |  |  |  | 2024-10-09 | (mBq/kg) | <1.32 |  |  | 51.96 | +- | 30.57 | <0.59 |  |  | <1.17 |  |  | <9.45 |  |  | <0.60 |  |  | 5.266 | +- | 0.6924
[6 empty] | (ppb or ppm) | <106.92 ppt |  |  | 4.21 ppb | +- | 2.48 ppb | <1.04 ppb |  |  | <0.29 ppb |  |  | <0.31 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 15959 | +- | 34140 | 2.4727 | +- | 3.251 | 3.4827 | +- | 0.6545 | 3.532 | +- | 1.458 | 1.187 | +- | 1.444 | 0.58289 | +- | 0.3819
[6 empty] | (ppb or ppm) | 1.29 ppm | +- | 2.77 ppm | [6 empty] | 0.87 ppb | +- | 0.36 ppb
PICO V12 |  | 444.3 g | 14.732 | 241009 | 2024-10-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Disk |  |  |  |  | 2024-10-24 | (mBq/kg) | <0.43 |  |  | <51.63 |  |  | 0.904 | +- | 0.3434 | 7.862 | +- | 1.03 | <7.06 |  |  | <0.53 |  |  | 6.736 | +- | 0.6606
[6 empty] | (ppb or ppm) | <0.03 ppb |  |  | <4.18 ppb |  |  | 1.59 ppb | +- | 0.6 ppb | 1.93 ppb | +- | 0.25 ppb | <0.23 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | <17240.00 |  |  | <3.94 |  |  | 1.2795 | +- | 0.5304 | 4.564 | +- | 1.347 | <0.70 |  |  | 0.16835 | +- | 0.3615
[6 empty] | (ppb or ppm) | <1.4 ppm | [8 empty] | 1.12 ppb | +- | 0.33 ppb
PICO V13 |  | 249.6 g | 13.72 | 241024 241030 | 2024-10-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Ring |  |  |  |  | 2024-11-07 | (mBq/kg) | <0.88 |  |  | 18.07 | +- | 40.55 | <0.56 |  |  | <2.68 |  |  | <7.83 |  |  | <1.13 |  |  | 5.561 | +- | 0.8653
[6 empty] | (ppb or ppm) | <71.28 ppt |  |  | 1.46 ppb | +- | 3.28 ppb | <0.99 ppb |  |  | <0.66 ppb |  |  | <252.91 ppb
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 3004.7 | +- | 12910 | <5.74 |  |  | 1.0764 | +- | 0.6557 | <2.61 |  |  | 3.221 | +- | 1.582 | <1.25
[6 empty] | (ppb or ppm) | 243.38 ppb | +- | 1045.71 ppb | [6 empty] | <0.64 ppb
DEAP Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
DEAP V01 | Carlisle Brushes | 151.7 g | 8.819 | 210322 | 2021-03-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Carlisle Brushes | Part No: 4039002 | Quantity: 5 |  |  | 2021-03-31 | (mBq/kg) | 90.62 | +- | 6.712 | 237.9 | +- | 99.48 | 5.367 | +- | 1.655 | 16.09 | +- | 4.774 | 648.34 | +- | 82.23 | <1.48 |  |  | <2.27
[6 empty] | (ppb or ppm) | 7.34 ppb | +- | 0.54 ppb | 19.27 ppb | +- | 8.06 ppb | 9.45 ppb | +- | 2.91 ppb | 3.96 ppb | +- | 1.17 ppb | 20.94 ppm | +- | 2.66 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 31077 | +- | 23210 | <20.17 |  |  | 2.5718 | +- | 2.074 | 31.42 | +- | 8.244
[6 empty] | (ppb or ppm) | 2.52 ppm | +- | 1.88 ppm | [6 empty] | 7.73 ppb | +- | 2.03 ppb
DEAP V02 | Gordon Brushes | 207.0 g | 14.715 | 210331 210405 | 2021-03-31 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Gordon Brushes | Part No: 906389 | Quantity: 5 |  |  | 2021-04-15 | (mBq/kg) | 0.4271 | +- | 1.656 | <67.28 |  |  | <1.63 |  |  | 10.54 | +- | 2.125 | <24.83 |  |  | <1.70 |  |  | 1.944 | +- | 0.8329
[6 empty] | (ppb or ppm) | 0.03 ppb | +- | 0.13 ppb | <5.45 ppb |  |  | <2.87 ppb |  |  | 2.59 ppb | +- | 0.52 ppb | <0.8 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 53660 | +- | 41610 | <13.76 |  |  | 1.8481 | +- | 0.8673 | 10.81 | +- | 3.288
[6 empty] | (ppb or ppm) | 4.35 ppm | +- | 3.37 ppm | [6 empty] | 2.66 ppb | +- | 0.81 ppb
CNL Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
CNL V01 |  | 49.2919 g | 4.443 | 230202 | 2023-02-02 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Pembroke Site 3 10-15 cm |  |  |  |  | 2023-02-07 | (mBq/kg) | 10880 | +- | 279.7 | 39520 | +- | 2966 | 501.2 | +- | 33.79 | 13540 | +- | 396 | 669460 | +- | 34440 | 4596.9 | +- | 254.7 | <80.57
[6 empty] | (ppb or ppm) | 881.28 ppb | +- | 22.66 ppb | 3201.12 ppb | +- | 240.25 ppb | 882.11 ppb | +- | 59.47 ppb | 3330.84 ppb | +- | 97.42 ppb | 21623.56 ppm | +- | 1112.41 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 57367 | +- | 108600 | <724.90 |  |  | 202.97 | +- | 60.98 | 13100 | +- | 432.4
[6 empty] | (ppb or ppm) | 4.65 ppm | +- | 8.8 ppm | [6 empty] | 3222.6 ppb | +- | 106.37 ppb
CNL V02 |  | 41.7266 g | 3.002 | 230207 | 2023-02-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Pembroke Site 2 10-15 cm |  |  |  |  | 2023-02-10 | (mBq/kg) | 10760 | +- | 302.5 | 26190 | +- | 2540 | 598.7 | +- | 41.24 | 10610 | +- | 348.3 | 705960 | +- | 36490 | 6556 | +- | 366.6 | <127.10
[6 empty] | (ppb or ppm) | 871.56 ppb | +- | 24.5 ppb | 2121.39 ppb | +- | 205.74 ppb | 1053.71 ppb | +- | 72.58 ppb | 2610.06 ppb | +- | 85.68 ppb | 22802.51 ppm | +- | 1178.63 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 256440 | +- | 123600 | <999.10 |  |  | 93.624 | +- | 66.66 | 10820 | +- | 418.8
[6 empty] | (ppb or ppm) | 20.77 ppm | +- | 10.01 ppm | [6 empty] | 2661.72 ppb | +- | 103.02 ppb
CNL V03 |  | 48.4052 g | 2.841 | 230210 | 2023-02-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Pembroke Site 1 10-15 cm |  |  |  |  | 2023-02-13 | (mBq/kg) | 9014 | +- | 367.5 | 29300 | +- | 2834 | 450.9 | +- | 40.66 | 12620 | +- | 403.2 | 807110 | +- | 41600 | 4355.5 | +- | 257 | <51.11
[6 empty] | (ppb or ppm) | 730.13 ppb | +- | 29.77 ppb | 2373.3 ppb | +- | 229.55 ppb | 793.58 ppb | +- | 71.56 ppb | 3104.52 ppb | +- | 99.19 ppb | 26069.65 ppm | +- | 1343.68 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 128320 | +- | 137500 | <655.20 |  |  | 139.87 | +- | 76.46 | 12320 | +- | 461
[6 empty] | (ppb or ppm) | 10.39 ppm | +- | 11.14 ppm | [6 empty] | 3030.72 ppb | +- | 113.41 ppb
CNL V04 |  | 51.0620 g | 3.899 | 230213 | 2023-02-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Lemieux Island 10-15 cm |  |  |  |  | 2023-02-17 | (mBq/kg) | 12050 | +- | 309.5 | 45830 | +- | 3314 | 746.3 | +- | 40.25 | 19810 | +- | 552.8 | 641560 | +- | 30300 | 2681.6 | +- | 166.6 | <71.39
[6 empty] | (ppb or ppm) | 976.05 ppb | +- | 25.07 ppb | 3712.23 ppb | +- | 268.43 ppb | 1313.49 ppb | +- | 70.84 ppb | 4873.26 ppb | +- | 135.99 ppb | 20722.39 ppm | +- | 978.69 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 641560 | +- | 33030 | 715.55 | +- | 458 | 361.53 | +- | 166.6 | 18830 | +- | 576.3
[6 empty] | (ppb or ppm) | 51.97 ppm | +- | 2.68 ppm | [6 empty] | 4632.18 ppb | +- | 141.77 ppb
CNL V05 |  | 39.3869 g | 3.871 | 230217 | 2023-02-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-004 10-15 cm |  |  |  |  | 2023-02-21 | (mBq/kg) | 26070 | +- | 598.9 | 77410 | +- | 5281 | 1596 | +- | 64.14 | 34040 | +- | 912.8 | 590410 | +- | 30510 | 972.83 | +- | 110.3 | 43.7 | +- | 77.43
[6 empty] | (ppb or ppm) | 2111.67 ppb | +- | 48.51 ppb | 6270.21 ppb | +- | 427.76 ppb | 2808.96 ppb | +- | 112.89 ppb | 8373.84 ppb | +- | 224.55 ppb | 19070.24 ppm | +- | 985.47 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <291400.00 |  |  | 160.19 | +- | 600 | 705.75 | +- | 94.7 | 32860 | +- | 933.7
[6 empty] | (ppb or ppm) | <23.6 ppm | [8 empty] | 8083.56 ppb | +- | 229.69 ppb
CNL V06 |  | 43.9680 g | 2.957 | 230221 | 2023-02-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-003 10-15 cm |  |  |  |  | 2023-02-24 | (mBq/kg) | 19080 | +- | 487.5 | 56920 | +- | 4889 | 1207 | +- | 62.66 | 41210 | +- | 1090 | 622680 | +- | 32210 | 4544.2 | +- | 274.3 | <85.05
[6 empty] | (ppb or ppm) | 1545.48 ppb | +- | 39.49 ppb | 4610.52 ppb | +- | 396.01 ppb | 2124.32 ppb | +- | 110.28 ppb | 10137.66 ppb | +- | 268.14 ppb | 20112.56 ppm | +- | 1040.38 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 286110 | +- | 189600 | <1059.00 |  |  | 715.21 | +- | 105.4 | 40500 | +- | 1125
[6 empty] | (ppb or ppm) | 23.17 ppm | +- | 15.36 ppm | [6 empty] | 9963 ppb | +- | 276.75 ppb
CNL V07 |  | 39.8567 g | 3.766 | 230224 230228 | 2023-02-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-002 10-15 cm |  |  |  |  | 2023-02-28 | (mBq/kg) | 11660 | +- | 310 | 32310 | +- | 2711 | 611 | +- | 38.36 | 12530 | +- | 383.4 | 371710 | +- | 19330 | 5223.6 | +- | 293.2 | 64.58 | +- | 58.44
[6 empty] | (ppb or ppm) | 944.46 ppb | +- | 25.11 ppb | 2617.11 ppb | +- | 219.59 ppb | 1075.36 ppb | +- | 67.51 ppb | 3082.38 ppb | +- | 94.32 ppb | 12006.23 ppm | +- | 624.36 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 78726 | +- | 115200 | 231.36 | +- | 444.7 | 206.38 | +- | 62.46 | 12400 | +- | 430.4
[6 empty] | (ppb or ppm) | 6.38 ppm | +- | 9.33 ppm | [6 empty] | 3050.4 ppb | +- | 105.88 ppb
CNL V08 |  | 41.9112 g | 2.915 | 29022801 | 2023-02-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-005 10-15 cm |  |  |  |  | 2023-03-03 | (mBq/kg) | 23300 | +- | 564.1 | 64860 | +- | 4750 | 1431 | +- | 65.68 | 29120 | +- | 814.2 | 333920 | +- | 27700 | 7712.2 | +- | 427.6 | <167.10
[6 empty] | (ppb or ppm) | 1887.3 ppb | +- | 45.69 ppb | 5253.66 ppb | +- | 384.75 ppb | 2518.56 ppb | +- | 115.6 ppb | 7163.52 ppb | +- | 200.29 ppb | 10785.62 ppm | +- | 894.71 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 119830 | +- | 180100 | 1389.2 | +- | 687.1 | 294.54 | +- | 91.79 | 28480 | +- | 856.4
[6 empty] | (ppb or ppm) | 9706.23 ppb | +- | 14588.1 ppb | [6 empty] | 7006.08 ppb | +- | 210.67 ppb
CNL V09 |  | 41.9332 g | 3.853 | 230303 | 2023-03-03 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
22-ES-006 10-15 cm |  |  |  |  | 2023-03-07 | (mBq/kg) | 25500 | +- | 581.2 | 68240 | +- | 4574 | 1435 | +- | 58.89 | 27010 | +- | 739.8 | 539400 | +- | 27910 | 1702.6 | +- | 131.9 | <93.04
[6 empty] | (ppb or ppm) | 2065.5 ppb | +- | 47.08 ppb | 5527.44 ppb | +- | 370.49 ppb | 2525.6 ppb | +- | 103.65 ppb | 6644.46 ppb | +- | 181.99 ppb | 17422.62 ppm | +- | 901.49 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 19616 | +- | 155700 | <1074.00 |  |  | 403.79 | +- | 82.89 | 26630 | +- | 780.6
[6 empty] | (ppb or ppm) | 1.59 ppm | +- | 12.61 ppm | [6 empty] | 6550.98 ppb | +- | 192.03 ppb
CNL V10 |  | 49.9195 g | 2.724 | 230307 | 2023-03-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Princess Sodalite Mine 10-15 cm |  |  |  |  | 2023-03-10 | (mBq/kg) | 22930 | +- | 552.9 | 63690 | +- | 4739 | 1049 | +- | 57.99 | 26130 | +- | 746.8 | 828250 | +- | 42690 | 62.474 | +- | 93.55 | <150.30
[6 empty] | (ppb or ppm) | 1857.33 ppb | +- | 44.78 ppb | 5158.89 ppb | +- | 383.86 ppb | 1846.24 ppb | +- | 102.06 ppb | 6427.98 ppb | +- | 183.71 ppb | 26752.48 ppm | +- | 1378.89 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <252500.00 |  |  | 1041 | +- | 655.6 | 377.08 | +- | 98.05 | 27410 | +- | 831
[6 empty] | (ppb or ppm) | <20.45 ppm | [8 empty] | 6742.86 ppb | +- | 204.43 ppb
CNL V11 |  | 43.4597 g | 6.3 | 230310 230314 | -619747 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Silver Crater Mine 10-15 cm |  |  |  |  | 2023-03-17 | (mBq/kg) | 99480 | +- | 1861 | 341400 | +- | 15640 | 8150 | +- | 186.3 | 45590 | +- | 1156 | 599190 | +- | 30810 | 5185.4 | +- | 287.6 | <169.50
[6 empty] | (ppb or ppm) | 8057.88 ppb | +- | 150.74 ppb | 27653.4 ppb | +- | 1266.84 ppb | 14344 ppb | +- | 327.89 ppb | 11215.14 ppb | +- | 284.38 ppb | 19353.84 ppm | +- | 995.16 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 349250 | +- | 196200 | 592.52 | +- | 637.6 | 730.47 | +- | 95.68 | 39990 | +- | 1103
[6 empty] | (ppb or ppm) | 28.29 ppm | +- | 15.89 ppm | [6 empty] | 9837.54 ppb | +- | 271.34 ppb
CNL V12 |  | 29.3726 g | 3.897 | 230317 | 2023-03-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Silver Crate Mine 5-10 cm |  |  |  |  | 2023-03-21 | (mBq/kg) | 164900 | +- | 3116 | 495300 | +- | 23130 | 11780 | +- | 279.8 | 52230 | +- | 1402 | 592680 | +- | 30740 | 9499.8 | +- | 527.6 | <237.60
[6 empty] | (ppb or ppm) | 13356.9 ppb | +- | 252.4 ppb | 40119.3 ppb | +- | 1873.53 ppb | 20732.8 ppb | +- | 492.45 ppb | 12848.58 ppb | +- | 344.89 ppb | 19143.56 ppm | +- | 992.9 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 117790 | +- | 306300 | <2320.00 |  |  | 849.02 | +- | 151.5 | 47310 | +- | 1409
[6 empty] | (ppb or ppm) | 9.54 ppm | +- | 24.81 ppm | [6 empty] | 11638.26 ppb | +- | 346.61 ppb
CNL V13 |  | 35.6399 g | 2.878 | 230321 | 2023-03-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-2 5-10 cm |  |  |  |  | 2023-03-24 | (mBq/kg) | 16170 | +- | 434.5 | 42940 | +- | 3753 | 733.1 | +- | 53.65 | 14460 | +- | 467.7 | 719350 | +- | 37270 | 5218.8 | +- | 313.4 | <188.70
[6 empty] | (ppb or ppm) | 1309.77 ppb | +- | 35.19 ppb | 3478.14 ppb | +- | 303.99 ppb | 1290.26 ppb | +- | 94.42 ppb | 3557.16 ppb | +- | 115.05 ppb | 23235.01 ppm | +- | 1203.82 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 287490 | +- | 175600 | <962.90 |  |  | 260.74 | +- | 95.67 | 13920 | +- | 543.2
[6 empty] | (ppb or ppm) | 23.29 ppm | +- | 14.22 ppm | [6 empty] | 3424.32 ppb | +- | 133.63 ppb
CNL V14 |  | 17.3276 g | 3.869 | 230324 | 2023-03-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-8 5-10 cm |  |  |  |  | 2023-03-28 | (mBq/kg) | 7913 | +- | 272.5 | 25650 | +- | 2906 | 485.3 | +- | 48.01 | 8285 | +- | 325.4 | 450450 | +- | 23810 | 16529 | +- | 889.6 | <136.30
[6 empty] | (ppb or ppm) | 640.95 ppb | +- | 22.07 ppb | 2077.65 ppb | +- | 235.39 ppb | 854.13 ppb | +- | 84.5 ppb | 2038.11 ppb | +- | 80.05 ppb | 14549.54 ppm | +- | 769.06 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 286810 | +- | 154000 | 1499.8 | +- | 717.5 | 96.874 | +- | 85.97 | 7417 | +- | 410.4
[6 empty] | (ppb or ppm) | 23.23 ppm | +- | 12.47 ppm | [6 empty] | 1824.58 ppb | +- | 100.96 ppb
CNL V15 |  | 40.0602 g | 2.916 | 230328 | 2023-03-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-5 5-10 cm |  |  |  |  | 2023-03-31 | (mBq/kg) | 16440 | +- | 425.1 | 35480 | +- | 3479 | 764.6 | +- | 50.13 | 10810 | +- | 366.7 | 704010 | +- | 36420 | 4833.4 | +- | 288.7 | <138.00
[6 empty] | (ppb or ppm) | 1331.64 ppb | +- | 34.43 ppb | 2873.88 ppb | +- | 281.8 ppb | 1345.7 ppb | +- | 88.23 ppb | 2659.26 ppb | +- | 90.21 ppb | 22739.52 ppm | +- | 1176.37 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 175180 | +- | 155600 | 716.91 | +- | 608.1 | 130.04 | +- | 84.14 | 10500 | +- | 441.8
[6 empty] | (ppb or ppm) | 14.19 ppm | +- | 12.6 ppm | [6 empty] | 2583 ppb | +- | 108.68 ppb
CNL V16 |  | 45.9848 g | 3.923 | 230331 | 2023-03-31 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-9 10-15 cm |  |  |  |  | 2023-04-04 | (mBq/kg) | 12370 | +- | 319.5 | 32200 | +- | 2829 | 582.3 | +- | 37.65 | 11900 | +- | 367.9 | 663670 | +- | 34190 | 3926.3 | +- | 227.9 | <65.56
[6 empty] | (ppb or ppm) | 1001.97 ppb | +- | 25.88 ppb | 2608.2 ppb | +- | 229.15 ppb | 1024.85 ppb | +- | 66.26 ppb | 2927.4 ppb | +- | 90.5 ppb | 21436.54 ppm | +- | 1104.34 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 41530 | +- | 119000 | <709.80 |  |  | 165.29 | +- | 65 | 11400 | +- | 412
[6 empty] | (ppb or ppm) | 3363.93 ppb | +- | 9639 ppb | [6 empty] | 2804.4 ppb | +- | 101.35 ppb
CNL V17 |  | 5.6586 g | 6.755 | 230404 | 2023-04-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-7 0-5 cm |  |  |  |  | 2023-04-11 | (mBq/kg) | 3382 | +- | 206.7 | 17150 |  | 2924 | 243.2 | +- | 52.88 | 6710 | +- | 322.6 | 170070 | +- | 9840 | 21146 | +- | 1156 | <73.91
[6 empty] | (ppb or ppm) | 273.94 ppb | +- | 16.74 ppb | 1389.15 ppb | +- | 236.84 ppb | 428.03 ppb | +- | 93.07 ppb | 1650.66 ppb | +- | 79.36 ppb | 5493.26 ppm | +- | 317.83 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 3012500 | +- | 393900 | <2021.00 |  |  | 250.25 | +- | 85.25 | 6325 | +- | 433.4
[6 empty] | (ppb or ppm) | 244.01 ppm | +- | 31.91 ppm | [6 empty] | 1555.95 ppb | +- | 106.62 ppb
CNL V18 |  | 42.6733 g | 2.64 | 230411 | 2023-04-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-9 10-15 cm |  |  |  |  | 2023-04-14 | (mBq/kg) | 8319 | +- | 256.6 | 26760 | +- | 2724 | 417.3 | +- | 41.61 | 8318 | +- | 298.6 | 618890 | +- | 32080 | 8730.7 | +- | 479.1 | <177.70
[6 empty] | (ppb or ppm) | 673.84 ppb | +- | 20.78 ppb | 2167.56 ppb | +- | 220.64 ppb | 734.45 ppb | +- | 73.23 ppb | 2046.23 ppb | +- | 73.46 ppb | 19990.15 ppm | +- | 1036.18 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 108290 | +- | 137400 | 311.44 | +- | 465.1 | 184.44 | +- | 77.41 | 7448 | +- | 367.6
[6 empty] | (ppb or ppm) | 8.77 ppm | +- | 11.13 ppm | [6 empty] | 1832.21 ppb | +- | 90.43 ppb
CNL V19 |  | 41.4027 g | 4.92 | 230414 | 2023-04-14 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 2-6 10-15 cm |  |  |  |  | 2023-04-19 | (mBq/kg) | 9160 | +- | 247.7 | 22370 | +- | 2222 | 459.4 | +- | 32.23 | 9486 | +- | 298.2 | 647360 | +- | 33310 | 4011.6 | +- | 227.8 | <56.49
[6 empty] | (ppb or ppm) | 741.96 ppb | +- | 20.06 ppb | 1811.97 ppb | +- | 179.98 ppb | 808.54 ppb | +- | 56.72 ppb | 2333.56 ppb | +- | 73.36 ppb | 20909.73 ppm | +- | 1075.91 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 124390 | +- | 105800 | 540.29 | +- | 412.5 | 173.6 | +- | 58.73 | 9330 | +- | 347.1
[6 empty] | (ppb or ppm) | 10.08 ppm | +- | 8.57 ppm | [6 empty] | 2295.18 ppb | +- | 85.39 ppb
CNL V20 |  | 29.3445 g | 5.894 | 230419 | 2023-04-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-8 5-10 cm |  |  |  |  | 2023-04-25 | (mBq/kg) | 12290 | +- | 313.1 | 27040 | +- | 2374 | 532 | +- | 35.45 | 10870 | +- | 335.9 | 564900 | +- | 29100 | 13061 | +- | 677.9 | 66.18 | +- | 63.58
[6 empty] | (ppb or ppm) | 995.49 ppb | +- | 25.36 ppb | 2190.24 ppb | +- | 192.29 ppb | 936.32 ppb | +- | 62.39 ppb | 2674.02 ppb | +- | 82.63 ppb | 18246.27 ppm | +- | 939.93 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 145200 | +- | 107100 | <601.50 |  |  | 203.01 | +- | 60.09 | 9552 | +- | 371.8
[6 empty] | (ppb or ppm) | 11.76 ppm | +- | 8.68 ppm | [6 empty] | 2349.79 ppb | +- | 91.46 ppb
CNL V21 |  | 27.2959 g | 6.926 | 230425 | 2023-04-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Forest 1-1 0-5 cm |  |  |  |  | 2023-05-02 | (mBq/kg) | 13010 | +- | 325 | 35750 | +- | 2724 | 602.8 | +- | 36.84 | 12570 | +- | 376.4 | 663360 | +- | 34110 | 18312 | +- | 938.5 | 63.93 | +- | 67.23
[6 empty] | (ppb or ppm) | 1053.81 ppb | +- | 26.33 ppb | 2895.75 ppb | +- | 220.64 ppb | 1060.93 ppb | +- | 64.84 ppb | 3092.22 ppb | +- | 92.59 ppb | 21426.53 ppm | +- | 1101.75 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 343740 | +- | 119000 | 1077 | +- | 490.8 | 71.448 | +- | 60.62 | 11610 | +- | 417.3
[6 empty] | (ppb or ppm) | 27.84 ppm | +- | 9.64 ppm | [6 empty] | 2856.06 ppb | +- | 102.66 ppb
LEGEND Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
LEGEND V01 |  | Copper: 58.01 g | 22.462 | 240125 | 2024-01-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Majorana Copper Embedded into Teflon Holder |  | Teflon: 275.132 g |  |  | 2024-02-09 | (mBq/kg) | <0.75 |  |  | 8.16 | +- | 27.66 | <0.88 |  |  | <1.58 |  |  | <9.16 |  |  | <0.97 |  |  | <0.52
[5 empty] | 2024-02-26 | (ppb or ppm) | <60.75 ppt |  |  | 0.66 ppb | +- | 2.24 ppb | <1.55 ppb |  |  | <0.39 ppb |  |  | <295.87 ppb
[5 empty] | March 5,2024 | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 5230.7 | +- | 2585 | <9.68 |  |  | 0.562 | +- | 0.5315 | <2.17 |  |  | <1.55 |  |  | <1.06
[6 empty] | (ppb or ppm) | 0.42 ppm | +- | 0.21 ppm | [6 empty] | <0.53 ppb
DarT Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
DArT V01 | McMaster-Carr | 823.8 g | 26.57 | 240315 24031801 | 2024-03-15 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Acrylic Sheet | PN; 8560K222 |  |  |  | 2024-04-11 | (mBq/kg) | 0.5951 | +- | 0.534 | <19.40 |  |  | <0.31 |  |  | <0.51 |  |  | <6.36 |  |  | <0.35 |  |  | <0.26
| Units: 50 |  |  |  |  | (ppb or ppm) | 48.2 ppt | +- | 43.25 ppt | <1.57 ppb |  |  | <0.55 ppb |  |  | <125.46 ppt |  |  | <205.43 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 3874.5 | +- | 1856 | <3.71 |  |  | 0.21949 | +- | 0.2633 | <0.75
[6 empty] | (ppb or ppm) | 313.83 ppb | +- | 150.34 ppb | [6 empty] | <184.5 ppt
QBITS-CUTE Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
QBITS-CUTE V01 | Chalmers University of Technology | 220.0 g | 13.882 | 240510 | 2024-05-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Sample Holder | Sample #1 |  |  |  | 2024-05-24 | (mBq/kg) | 21.21 | +- | 1.945 | 698.2 | +- | 83.83 | 23.1 | +- | 1.279 | 38.39 | +- | 2.713 | 61.419 | +- | 18.59 | <1.74 |  |  | <0.64
[6 empty] | (ppb or ppm) | 1.72 ppb | +- | 0.16 ppb | 56.55 ppb | +- | 6.79 ppb | 40.66 ppb | +- | 2.25 ppb | 9.44 ppb | +- | 0.67 ppb | 1.98 ppm | +- | 0.6 ppm
| Also Counted as PGT QBITS-CUTE 10 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <60310.00 |  |  | <9.80 |  |  | 1.161 | +- | 0.8752 | 42.86 | +- | 3.814
[6 empty] | (ppb or ppm) | <4.89 ppm | [8 empty] | 10.54 ppb | +- | 0.94 ppb
QBITS-CUTE V02 | Chalmers University of Technology | 8.164 g | 9.664 | 240524 250528 | 2024-05-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Gold Plated Printed Circuit Boards | Two Units |  |  |  | 2024-06-03 | (mBq/kg) | 1402 | +- | 65.96 | 4541 | +- | 888.7 | 55.72 | +- | 12.27 | 1707 | +- | 85.84 | 762.71 | +- | 354.1 | <22.81 |  |  | <21.11
[6 empty] | (ppb or ppm) | 113.56 ppb | +- | 5.34 ppb | 367.82 ppb | +- | 71.98 ppb | 98.07 ppb | +- | 21.6 ppb | 419.92 ppb | +- | 21.12 ppb | 24.64 ppm | +- | 11.44 ppm
| Also Counted as PGT QBITS-CUTE 08 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <100700.00 |  |  | 192.58 | +- | 150.1 | 47.064 | +- | 19.42 | 1765 | +- | 111.5
[6 empty] | (ppb or ppm) | <8.16 ppm | [8 empty] | 434.19 ppb | +- | 27.43 ppb
QBITS-CUTE V03 | Chalmers University of Technology | 7.835 g | 9.983 | 240603 | 2024-06-03 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
PCB ELCO | Bluefors: Elco BV, Hofsttetter, EPIG Plating |  |  |  | 2024-06-13 | (mBq/kg) | 1364 | +- | 66.03 | 3541 | +- | 687.6 | 66.78 | +- | 12.65 | 2130 | +- | 98.61 | 1408.1 | +- | 390.1 | 12.078 | +- | 22.57 | <17.10
| Sample #3 2 units |  |  |  |  | (ppb or ppm) | 110.48 ppb | +- | 5.35 ppb | 286.82 ppb | +- | 55.7 ppb | 117.53 ppb | +- | 22.26 ppb | 523.98 ppb | +- | 24.26 ppb | 45.48 ppm | +- | 12.6 ppm
| Also Counted as PGT QBITS-CUTE 11 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 46030 | +- | 49530 | <134.80 |  |  | 46.201 | +- | 20.03 | 1799 | +- | 114.1
[6 empty] | (ppb or ppm) | 3.73 ppm | +- | 4.01 ppm | [6 empty] | 442.55 ppb | +- | 28.07 ppb
U Calgary Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
U Calgary Sediment V01 | Airdrie Residence Soil Samples | 308.9 g | 3.029 | 24072301 | 2024-07-23 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Airdrie Residence AR15-30A | Depth 15 to 30 cm |  |  |  | 2024-07-26 | (mBq/kg) | 29970 | +- | 591.5 | 51310 | +- | 3493 | 1451 | +- | 46.12 | 26740 | +- | 669.4 | 456400 | +- | 23390 | 1468.21 | +- | 47.33 | 20.44 | +- | 35.53
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.43 ppm | +- | 0.05 ppm | 4.16 ppm | +- | 0.28 ppm | 2.55 ppm | +- | 0.08 ppm | 6.58 ppm | +- | 0.16 ppm | 14741.72 ppm | +- | 755.5 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (mBq/kg) | 129792.86 | +- | 51714.38 | 537.74 | +- | 343.2 | 377.19 | +- | 48.1 | 25830 | +- | 668
[6 empty] | (ppm / ppb / ppt) | 10.51 ppm | +- | 4.19 ppm | [6 empty] | 6.35 ppm | +- | 0.16 ppm
U Calgary Sediment V02 | Airdrie Residence Soil Samples | 345.5 g | 3.832 | 240726 | 2024-07-26 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Airdrie Residence AR30-43A | Depth 30 to 43 cm |  |  |  | 2024-07-30 | (mBq/kg) | 27530 | +- | 528.5 | 50350 | +- | 2967 | 1228 | +- | 36.8 | 21140 | +- | 522.8 | 347210 | +- | 17790 | 112.86 | +- | 27.01 | <20.31
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.23 ppm | +- | 0.04 ppm | 4.08 ppm | +- | 0.24 ppm | 2.16 ppm | +- | 0.06 ppm | 5.2 ppm | +- | 0.13 ppm | 11214.88 ppm | +- | 574.62 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (mBq/kg) | 101317.67 | +- | 37745.8 | 811.77 | +- | 256.7 | 326.09 | +- | 36.35 | 19630 | +- | 508.9
[6 empty] | (ppm / ppb / ppt) | 8.21 ppm | +- | 3.06 ppm | [6 empty] | 4.83 ppm | +- | 0.13 ppm
U Calgary Sediment V03 | Airdrie Residence Soil Samples | 148.9 g | 3.014 | 240730 | 2024-07-30 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Airdrie Residence AR43-50A | Depth 43 to 50 cm |  |  |  | 2024-08-02 | (mBq/kg) | 27490 | +- | 558.1 | 50580 | +- | 3575 | 1309 | +- | 44.61 | 24380 | +- | 622.2 | 428500 | +- | 22070 | 1460.9 | +- | 97.22 | <64.73
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.23 ppm | +- | 0.05 ppm | 4.1 ppm | +- | 0.29 ppm | 2.3 ppm | +- | 0.08 ppm | 6 ppm | +- | 0.15 ppm | 13840.55 ppm | +- | 712.86 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (mBq/kg) | 166511.36 | +- | 100457.05 | 2421.1 | +- | 409 | 500.25 | +- | 57.75 | 23150 | +- | 621.9
[6 empty] | (ppm / ppb / ppt) | 13.49 ppm | +- | 8.14 ppm | [6 empty] | 5.69 ppm | +- | 0.15 ppm
U Calgary Sediment V04 | Charleswood Residence Soil Samples | 27.0 g | 3.9 | 240802 | 2024-08-02 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Charleswood Residence CR0-10 | Depth 0 to 10 cm |  |  |  | 2024-08-06 | (mBq/kg) | 20940 | +- | 509.7 | 53210 | +- | 3936 | 1102 | +- | 53.26 | 17510 | +- | 530.7 | 357150 | +- | 18740 | 1998.1 | +- | 149.8 | <92.66
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.7 ppm | +- | 0.04 ppm | 4.31 ppm | +- | 0.32 ppm | 1.94 ppm | +- | 0.09 ppm | 4.31 ppm | +- | 0.13 ppm | 11535.95 ppm | +- | 605.3 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (mBq/kg) | 132800 | +- | 163700 | 985.66 | +- | 568.5 | 246.15 | +- | 78.6 | 16390 | +- | 561.6
[6 empty] | (ppm / ppb / ppt) | 10.76 ppm | +- | 13.26 ppm | [6 empty] | 4.03 ppm | +- | 0.14 ppm
U Calgary Sediment V05 | Charleswood Residence Soil Samples | 39.1 g | 2.935 | 240806 | 2024-08-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Charleswood Residence CR10-20 | Depth 10 to 20 cm |  |  |  | 2024-08-09 | (mBq/kg) | 18920 | +- | 467.5 | 46840 | +- | 3442 | 967 | +- | 49.72 | 15750 | +- | 481.1 | 302000 | +- | 15880 | 1252.1 | +- | 116.3 | 62.03 | +- | 63.19
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.53 ppm | +- | 0.04 ppm | 3.79 ppm | +- | 0.28 ppm | 1.7 ppm | +- | 0.09 ppm | 3.87 ppm | +- | 0.12 ppm | 9754.6 ppm | +- | 512.92 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (mBq/kg) | 278890 | +- | 155400 | <557.50 |  |  | 259.44 | +- | 74.7 | 15130 | +- | 522.4
[6 empty] | (ppm / ppb / ppt) | 22.59 ppm | +- | 12.59 ppm | [6 empty] | 3.72 ppm | +- | 0.13 ppm
U Calgary Sediment V06 | Charleswood Residence Soil Samples | 43.5 g | 3.918 | 240809 | 2024-08-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Charleswood Residence CR20-30 | Depth 20 to 30 cm |  |  |  | 2024-08-13 | (mBq/kg) | 19800 | +- | 455.9 | 40710 | +- | 3048 | 949.1 | +- | 43.61 | 15350 | +- | 450.2 | 289290 | +- | 15090 | 1688.8 | +- | 121.3 | 70.09 | +- | 53.55
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.6 ppm | +- | 0.04 ppm | 3.3 ppm | +- | 0.25 ppm | 1.67 ppm | +- | 0.08 ppm | 3.78 ppm | +- | 0.11 ppm | 9344.07 ppm | +- | 487.41 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (mBq/kg) | 218290 | +- | 125300 | 544.9 | +- | 448.5 | 317.34 | +- | 65.07 | 15410 | +- | 492.1
[6 empty] | (ppm / ppb / ppt) | 17.68 ppm | +- | 10.15 ppm | [6 empty] | 3.79 ppm | +- | 0.12 ppm
U Calgary Sediment V07 | Porcupine Hills Soil Samples | 79.3 g | 2.927 | 240813 | 2024-08-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Porcupine Hills PHfor |  |  |  |  | 2024-08-16 | (mBq/kg) | 13910 | +- | 327.8 | 30690 | +- | 2243 | 609.8 | +- | 31.94 | 8905 | +- | 272.3 | 239350 | +- | 7378 | 76.28 | +- | 29.34 | <57.99
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.13 ppm | +- | 0.03 ppm | 2.49 ppm | +- | 0.18 ppm | 1.07 ppm | +- | 0.06 ppm | 2.19 ppm | +- | 0.07 ppm | 7731.01 ppm | +- | 238.31 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (mBq/kg) | 38875.06 | +- | 28733.74 | 320.73 | +- | 325.8 | 150.39 | +- | 44.69 | 8493 | +_ | 300.6
[6 empty] | (ppm / ppb / ppt) | 3.15 ppm | +- | 2.33 ppm | [6 empty] | 2.09 ppm | +- | 0.07 ppm
U Calgary Sediment V08 | Spy Hill Soil Samples | 68.8 g | 6.909 | 240816 | 2024-08-16 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Spy Hills SHfor |  |  |  |  | 2024-08-23 | (mBq/kg) | 33880 | +- | 669.7 | 63130 | +- | 3956 | 1534 | +- | 47.67 | 29290 | +- | 737 | 490430 | +- | 25180 | 200.83 | +- | 54.63 | 73.838 | +- | 50.82
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.74 ppm | +- | 0.05 ppm | 5.11 ppm | +- | 0.32 ppm | 2.7 ppm | +- | 0.08 ppm | 7.21 ppm | +- | 0.18 ppm | 15840.89 ppm | +- | 813.31 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (mBq/kg) | 15606.14 | +- | 6092.24 | 708.14 | +- | 379.7 | 492.96 | +- | 58.71 | 27410 | +- | 717.1
[6 empty] | (ppm / ppb / ppt) | 1.26 ppm | +- | 0.49 ppm | [6 empty] | 6.74 ppm | +- | 0.18 ppm
U Calgary Sediment V09 | Breathing Well Soil Samples | 54.0 g | 4.008 | 240823 | 2024-08-23 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Breathing well BW0-10 | Depth: 0 to 10 cm |  |  |  | 2024-08-27 | (mBq/kg) | 34500 | +- | 737 | 68030 | +- | 4834 | 1620 | +- | 61.38 | 31040 | +- | 830.3 | 570150 | +- | 29410 | 4449.8 | +- | 255.5 | 61.67 | +- | 67.61
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.79 ppm | +- | 0.06 ppm | 5.51 ppm | +- | 0.39 ppm | 2.85 ppm | +- | 0.11 ppm | 7.64 ppm | +- | 0.2 ppm | 18415.85 ppm | +- | 949.94 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
[6 empty] | (mBq/kg) | 14584.1 | +- | 5722.88 | <842.90 |  |  | 441.35 | +- | 863.48 | 29360 | +- | 829.1
[6 empty] | (ppm / ppb / ppt) | 1.18 ppm | +- | 0.46 ppm | [6 empty] | 7.22 ppm | +- | 0.2 ppm
U Calgary Sediment V10 | Vermillion Core Samples GM1 10-12 | 401.5 g | 2.975 | 250203 | 2025-02-03 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM1 10-12 |  |  |  |  | 2025-02-06 | (mBq/kg) | 20870 | +- | 419.5 | 43010 | +- | 2821 | 1029 | +- | 447.5 | 17530 | +- | 447.5 | 534000 | +- | 27330 |  | NA |  | <32.26
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.69 ppm | +- | 0.03 ppm | 3.48 ppm | +- | 0.23 ppm | 1.81 ppm | +- | 0.79 ppm | 4.31 ppm | +- | 0.11 ppm | 17248.2 ppm | +- | 882.76 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | <112600.00 |  |  | 385.27 | +- | 271 | 347.94 | +- | 40.46 | 16760 | +- | 455.8 | 455.13 | +- | 55.47
[6 empty] | (ppm / ppb / ppt) | <9.12 ppm | [8 empty] | 4.12 ppm | +- | 0.11 ppm
U Calgary Sediment V11 | Vermillion Core Samples GM1 21-23 | 297.0g | 3.78 | 250206 | 2025-02-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM1 21-23 |  |  |  |  | 2025-02-10 | (mBq/kg) | 18610 | +- | 383.4 | 41810 | +- | 2901 | 953 | +- | 33.43 | 16480 | +- | 428 | 649390 | +- | 33230 |  | NA |  | <32.24
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.51 ppm | +- | 0.03 ppm | 3.39 ppm | +- | 0.23 ppm | 1.68 ppm | +- | 0.06 ppm | 4.05 ppm | +- | 0.11 ppm | 20975.3 ppm | +- | 1073.33 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | <155600.00 |  |  | 94.291 | +- | 279.5 | 283.03 | +- | 40.9 | 16120 | +- | 444 | 527.04 | +- | 59.11
[6 empty] | (ppm / ppb / ppt) | <12.6 ppm | [8 empty] | 3.97 ppm | +- | 0.11 ppm
U Calgary Sediment V12 | Vermillion Core Samples GM1 38-40 | 475.3 g | 7.82 | 250210 250211 25021101 | 2025-02-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM1 38-40 |  |  |  |  | 2025-02-18 | (mBq/kg) | 27770 | +- | 517.1 | 38410 | +- | 2257 | 1260 | +- | 32.6 | 20130 | +- | 483.9 | 482170 | +- | 24640 |  | NA |  | <35.05
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.25 ppm | +- | 0.04 ppm | 3.11 ppm | +- | 0.18 ppm | 2.22 ppm | +- | 0.06 ppm | 4.95 ppm | +- | 0.12 ppm | 15574.09 ppm | +- | 795.87 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | 66623 | +- | 63020 | 238.72 | +- | 165.7 | 333.53 | +- | 27.31 | 18410 | +- | 455.9 | 479.85 | +- | 38.73
[6 empty] | (ppm / ppb / ppt) | 5.4 ppm | +- | 5.1 ppm | [6 empty] | 4.53 ppm | +- | 0.11 ppm
U Calgary Sediment V13 | Vermillion Core Samples GM1 49-51 | 422.7 g | 4.378 | 250218 | 2025-02-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM1 49-51 |  |  |  |  | 2025-02-24 | (mBq/kg) | 20550 | +- | 401.8 | 39560 | +- | 2416 | 945.3 | +- | 29.52 | 16390 | +- | 410.2 | 512920 | +- | 26230 |  | NA |  | <34.78
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.66 ppm | +- | 0.03 ppm | 3.2 ppm | +- | 0.2 ppm | 1.66 ppm | +- | 0.05 ppm | 4.03 ppm | +- | 0.1 ppm | 16567.32 ppm | +- | 847.23 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | 53532 | +- | 80710 | 528.82 | +- | 215 | 213.53 | + | 30.49 | 15510 | +- | 408.9 | 391.23 | +- | 44.53
[6 empty] | (ppm / ppb / ppt) | 4.34 ppm | +- | 6.54 ppm | [6 empty] | 3.82 ppm | +- | 0.1 ppm
U Calgary Sediment V14 | Vermillion Core Samples GM1 58-60 | 456.2 g | 7.079 | 290224 | 2025-02-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM1 58-60 |  |  |  |  | 2025-03-03 | (mBq/kg) | 25730 | +- | 501 | 44680 | +- | 2523 | 1271 | +- | 33.58 | 20950 |  | 504.8 | 487670 | +- | 24920 |  | NA |  | <16.77
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.08 ppm | +- | 0.04 ppm | 3.62 ppm | +- | 0.2 ppm | 2.24 ppm | +- | 0.06 ppm | 5.15 ppm | +- | 0.12 ppm | 15751.74 ppm | +- | 804.92 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | 6769.8 | +- | 66660 | 417.9 | +- | 175.9 | 259.37 | +- | 25.75 | 19570 | +- | 483.9 | 520.63 | +- | 41.67
[6 empty] | (ppm / ppb / ppt) | 0.55 ppm | +- | 5.4 ppm | [6 empty] | 4.81 ppm | +- | 0.12 ppm
U Calgary Sediment V15 | Vermillion Core Samples GM2 17-19 | 516.8 g | 3.748 | 250303 | 2025-03-03 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM2 17-19 |  |  |  |  | 2025-03-07 | (mBq/kg) | 22580 | +- | 434.6 | 34220 | +- | 2185 | 953.9 | +- | 29.62 | 14900 | +- | 374.8 | 399950 | +- | 20460 |  | NA |  | <48.14
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.83 ppm | +- | 0.04 ppm | 2.77 ppm | +- | 0.18 ppm | 1.68 ppm | +- | 0.05 ppm | 3.67 ppm | +- | 0.09 ppm | 12918.39 ppm | +- | 660.86 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | <118500.00 |  |  | 350.13 | +- | 206.4 | 197 | +- | 28.47 | 13420 | +- | 365.3 | 343.23 | +- | 41.43
[6 empty] | (ppm / ppb / ppt) | <9.6 ppm | [8 empty] | 3.3 ppm | +- | 0.09 ppm
U Calgary Sediment V16 | Vermillion Core Samples GM2 27-29 | 542.9 g | 2.977 | 250307 | 2025-03-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM2 27-29 |  |  |  |  | 2025-03-10 | (mBq/kg) | 13380 | +- | 275.1 | 31590 | +- | 2014 | 831.4 | +- | 23.91 | 11920 | +- | 307.6 | 451600 | +- | 73110 |  | NA |  | <33.38
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.08 ppm | +- | 0.02 ppm | 2.56 ppm | +- | 0.16 ppm | 1.46 ppm | +- | 0.04 ppm | 2.93 ppm | +- | 0.08 ppm | 14586.68 ppm | +- | 2361.45 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | <100800.00 |  |  | 183.78 | +- | 203.6 | 165.98 | +- | 28.63 | 10910 | +- | 309.2 | 189.73 | +- | 39.01
[6 empty] | (ppm / ppb / ppt) | <8.16 ppm | [8 empty] | 2.68 ppm | +- | 0.08 ppm
U Calgary Sediment V17 | Vermillion Core Samples GM2 37-39 | 245.5 g | 3.925 | 250317 | 2025-03-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM2 37-39 |  |  |  |  | 2025-03-21 | (mBq/kg) | 24850 | +- | 502.9 | 57150 | +- | 3669 | 1331 | +- | 43 | 22210 | +- | 564.7 | 38910 | +- | 19920 |  | NA |  | <52.99
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.01 ppm | +- | 0.04 ppm | 4.63 ppm | +- | 0.3 ppm | 2.34 ppm | +- | 0.08 ppm | 5.46 ppm | +- | 0.14 ppm | 1256.79 ppm | +- | 643.42 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | <146100.00 |  |  | 174.36 | +- | 182.5 | 185.68 | +- | 27.32 | 20970 | +- | 568.9 | 381.74 | +- | 40.31
[6 empty] | (ppm / ppb / ppt) | <11.83 ppm | [8 empty] | 5.16 ppm | +- | 0.14 ppm
U Calgary Sediment V18 | Vermillion Core Samples GM2 47-49 | 418.4 g | 2.678 | 250310 250312 25031201 250313 | 2025-03-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM2 47-49 |  |  |  |  | 2025-03-14 | (mBq/kg) | 22530 | +- | 451.9 | 44080 | +- | 2933 | 1045 | +- | 36.12 | 17960 | +- | 459.1 | 524020 | +- | 26830 |  | NA |  | <42.93
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.82 ppm | +- | 0.04 ppm | 3.57 ppm | +- | 0.24 ppm | 1.84 ppm | +- | 0.06 ppm | 4.42 ppm | +- | 0.11 ppm | 16925.85 ppm | +- | 866.61 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | <102900.00 |  |  | <524.50 |  |  | 297.26 | +- | 40.89 | 16450 | +- | 457.3 | 461.39 | +- | 58.01
[6 empty] | (ppm / ppb / ppt) | <8.33 ppm | [8 empty] | 4.05 ppm | +- | 0.11 ppm
U Calgary Sediment V19 | Vermillion Core Samples GM2 54-56 | 257.5 g | 2.922 | 250314 | 2025-03-14 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM2 54-56 |  |  |  |  | 2025-03-17 | (mBq/kg) | 35890 | +- | 713 | 77530 | +- | 4701 | 1753 | +- | 56.23 | 28800 | +- | 730.1 | 644840 | +- | 33030 |  | NA |  | 25.32 | +- | 46.49
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.91 ppm | +- | 0.06 ppm | 6.28 ppm | +- | 0.38 ppm | 3.09 ppm | +- | 0.1 ppm | 7.08 ppm | +- | 0.18 ppm | 20828.33 ppm | +- | 1066.87 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | <201700.00 |  |  | 648.38 | +- | 410.7 | 483.59 | +- | 60.18 | 28100 | +- | 750.3 | 588.01 | +- | 83.01
[6 empty] | (ppm / ppb / ppt) | <16.34 ppm | [8 empty] | 6.91 ppm | +- | 0.18 ppm
U Calgary Sediment V20 | Vermillion Core Samples GM3 26-28 | 379.6 g | 2.921 | 250321 | 2025-03-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM3 26-28 |  |  |  |  | 2025-03-24 | (mBq/kg) | 17680 | +- | 365.5 | 39410 | +- | 2723 | 885.8 | +- | 32.63 | 15990 | +- | 414.9 | 576610 | +- | 29510 |  | NA |  | <67.21
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.43 ppm | +- | 0.03 ppm | 3.19 ppm | +- | 0.22 ppm | 1.56 ppm | +- | 0.06 ppm | 3.93 ppm | +- | 0.1 ppm | 18624.5 ppm | +- | 953.17 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | <88190.00 |  |  | 278.28 | +- | 275 | 251.88 | +- | 39.42 | 15380 | +- | 429.1 | 467.63 | +- | 57.43
[6 empty] | (ppm / ppb / ppt) | <7.14 ppm | [8 empty] | 3.78 ppm | +- | 0.11 ppm
U Calgary Sediment V21 | Vermillion Core Samples GM3 36-38 | 379.9 g | 3.955 | 250324 | 2025-03-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM3 36-38 |  |  |  |  | 2025-03-28 | (mBq/kg) | 26150 | +- | 507.8 | 53730 | +- | 2128 | 1234 | +- | 37.35 | 19800 | +- | 494.6 | 541700 | +- | 27710 |  | NA |  | <46.08
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.12 ppm | +- | 0.04 ppm | 4.35 ppm | +- | 0.17 ppm | 2.17 ppm | +- | 0.07 ppm | 4.87 ppm | +- | 0.12 ppm | 17496.91 ppm | +- | 895.03 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | 130970 | +- | 100400 | 316.68 | +- | 266.4 | 314.02 | +- | 37.54 | 18530 | +- | 491 | 497.22 | +- | 53.7
[6 empty] | (ppm / ppb / ppt) | 10.61 ppm | +- | 8.13 ppm | [6 empty] | 4.56 ppm | +- | 0.12 ppm
U Calgary Sediment V22 | Vermillion Core Samples GM3 47-48 | 375.4 g | 2.949 | 250328 | 2025-03-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM3 47-49 |  |  |  |  | 2025-03-31 | (mBq/kg) | 25800 |  | 529.1 | 45690 | +- | 3029 | 1308 | +- | 41.61 | 17920 | +- | 460.6 | 484400 | +- | 24810 |  | NA |  | <58.30
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.09 ppm | +- | 0.04 ppm | 3.7 ppm | +- | 0.25 ppm | 2.3 ppm | +- | 0.07 ppm | 4.41 ppm | +- | 0.11 ppm | 15646.12 ppm | +- | 801.36 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | 121920 | +- | 112800 | 671.18 | +- | 296.8 | 226.68 | +- | 39.91 | 16960 | +- | 471.6 | 450.59 | +- | 59.06
[6 empty] | (ppm / ppb / ppt) | 9.88 ppm | +- | 9.14 ppm | [6 empty] | 4.17 ppm | +- | 0.12 ppm
U Calgary Sediment V23 | Vermillion Core Samples GM3 50-52 | 330.6 g | 2.233 | 250331 | 2025-03-31 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM3 50-52 |  |  |  |  | 2025-04-02 | (mBq/kg) | 24420 | +- | 506.4 | 56030 | +- | 3655 | 1296 | +- | 46.54 | 19350 | +- | 510.1 | 537220 | +- | 27550 |  | NA |  | <22.54
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 1.98 ppm | +- | 0.04 ppm | 4.54 ppm | +- | 0.3 ppm | 2.28 ppm | +- | 0.08 ppm | 4.76 ppm | +- | 0.13 ppm | 17352.21 ppm | +- | 889.87 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | 192360 | +- | 143600 | 13.832 | +- | 363.4 | 272.95 | +- | 50.46 | 19220 | +- | 544 | 487.81 | +- | 72.44
[6 empty] | (ppm / ppb / ppt) | 15.58 ppm | +- | 11.63 ppm | [6 empty] | 4.73 ppm | +- | 0.13 ppm
U Calgary Sediment V24 | Vermillion Core Samples GM3 53-55 | 400.5 g | 4.801 | 250402 | 2025-04-02 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM3 53-55 |  |  |  |  | 2025-04-07 | (mBq/kg) | 25830 | +- | 495.4 | 49890 | +- | 2911 | 1273 | +- | 36.26 | 21320 | +- | 523.3 | 523700 | +- | 26780 |  | NA |  | <23.95
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.09 ppm | +- | 0.04 ppm | 4.04 ppm | +- | 0.24 ppm | 2.24 ppm | +- | 0.06 ppm | 5.24 ppm | +- | 0.13 ppm | 16915.51 ppm | +- | 864.99 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | 88745 | +- | 90560 | <326.40 |  |  | 257.34 | +- | 32.94 | 19720 | +- | 506.4 | 504.53 | +- | 48.85
[6 empty] | (ppm / ppb / ppt) | 7.19 ppm | +- | 7.34 ppm | [6 empty] | 4.85 ppm | +- | 0.12 ppm
U Calgary Sediment V25 | Vermillion Core Samples GM3 56-58 | 517.4 g | 3.926 | 250407 | 2025-04-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM3 56-58 |  |  |  |  | 2025-04-11 | (mBq/kg) | 25870 | +- | 494.3 | 49490 | +- | 2808 | 1172 | +- | 34.34 | 21070 | +- | 516.9 | 479500 | +- | 24520 |  | NA |  | <46.11
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.1 ppm | +- | 0.04 ppm | 4.01 ppm | +- | 0.23 ppm | 2.06 ppm | +- | 0.06 ppm | 5.18 ppm | +- | 0.13 ppm | 15487.85 ppm | +- | 792 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | <115800.00 |  |  | 445.77 | +- | 224.9 | 300.5 | +- | 32.94 | 19030 | +- | 492.2 | 522.54 | +- | 49.16
[6 empty] | (ppm / ppb / ppt) | <9.38 ppm | [8 empty] | 4.68 ppm | +- | 0.12 ppm
U Calgary Sediment V26 | Vermillion Core Samples GM3 63-65 | 161.9 g | 2.935 | 250411 | 2025-04-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Core Sample GM3 63-65 |  |  |  |  | 2025-04-14 | (mBq/kg) | 40780 | +- | 830.9 | 90720 | +- | 5907 | 2068 | +- | 70.39 | 34990 | +- | 900 | 765880 | +- | 39280 |  | NA |  | <65.01
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 3.3 ppm | +- | 0.07 ppm | 7.35 ppm | +- | 0.48 ppm | 3.64 ppm | +- | 0.12 ppm | 8.61 ppm | +- | 0.22 ppm | 24737.92 ppm | +- | 1268.74 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | 111570 | +- | 217800 | 480.17 | +- | 545 | 677.9 | +- | 83.38 | 33010 | +- | 909 | 1545 | +- | 99.87
[6 empty] | (ppm / ppb / ppt) | 9.04 ppm | +- | 17.64 ppm | [6 empty] | 8.12 ppm | +- | 0.22 ppm
U Calgary Sediment V27 |  | 406.9 g | 2.151 | 250414 | 2025-04-14 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
RR1 13Ft, 0-32 |  |  |  |  | 2025-04-16 | (mBq/kg) | 30980 | +- | 616.1 | 56620 | +- | 3740 | 1351 | +- | 47.29 | 28100 | +- | 705.6 | 499870 | +- | 25620 | 3609.1 | +- | 195.1 | <46.73
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.51 ppm | +- | 0.05 ppm | 4.59 ppm | +- | 0.3 ppm | 2.38 ppm | +- | 0.08 ppm | 6.91 ppm | +- | 0.17 ppm | 16145.8 ppm | +- | 827.53 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | 6557.5 | +- | 144400 | 434.55 | +- | 374.9 | 400.14 | +- | 52.39 | 26630 | +- | 711.1 | 696.13 | +- | 77.33
[6 empty] | (ppm / ppb / ppt) | 0.53 ppm | +- | 11.7 ppm | [6 empty] | 6.55 ppm | +- | 0.17 ppm
U Calgary Sediment V28 |  | 414.6 g | 6.804 | 250416 25041601 250423 | 2025-04-16 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
RR1 13Ft, 33-58 |  |  |  |  | 2025-04-23 | (mBq/kg) | 33120 | +- | 600.9 | 48120 | +- | 2810 | 1438 | +- | 38.3 | 30540 | +- | 729.5 | 537550 | +- | 27470 | 670.1 | +- | 43.29 | <15.83
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.68 ppm | +- | 0.05 ppm | 3.9 ppm | +- | 0.23 ppm | 2.53 ppm | +- | 0.07 ppm | 7.51 ppm | +- | 0.18 ppm | 17362.87 ppm | +- | 887.28 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | 107060 | +- | 83860 | 634.11 | +- | 213.9 | 418 | +- | 34.77 | 27490 | +- | 670.5 | 759.36 | +- | 54.95
[6 empty] | (ppm / ppb / ppt) | 8.67 ppm | +- | 6.79 ppm | [6 empty] | 6.76 ppm | +- | 0.16 ppm
RAMPS Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
RAMPS V01 | Thorlabs Inc. | 155.9 g | 23.516 | 240827 | 2024-08-27 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Thorlabs PDA045 SiPM Detector | Window Material: Borosilicate Glass |  |  |  | 2024-09-20 | (mBq/kg) | 1600 | +- | 36.71 | 22440 | +- | 1020 | 410.3 | +- | 10 | 3366 | +- | 85.54 | 2576.2 | +- | 157.4 |  | N/A |  | 3.674 | +- | 3.281
| Detector Window Faces the HPGe Detector |  |  |  |  | (ppb or ppm) | 129.6 ppb | +- | 2.97 ppb | 1817.64 ppb | +- | 82.62 ppb | 722.13 ppb | +- | 17.6 ppb | 828.04 ppb | +- | 21.04 ppb | 83.21 ppm | +- | 5.08 ppm | Peak interference from 232Th progeny
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 43311 | +- | 9088 | <56.44 |  |  | 34.521 | +- | 5.293 | 1947 | +- | 59.15
[6 empty] | (ppb or ppm) | 3.51 ppm | +- | 0.74 ppm | [6 empty] | 478.96 ppb | +- | 14.55 ppb
DarkSide Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
5 samples have been counted for DarkSide results are only available to DarkSide collaboration members.
In Progress Sample Measurements Counted on the Vue des Alpes Detector
In Progress and To Be Measured:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Background 13 |  | N/A | 13.961 | 241107 241113 | 2024-11-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2024-11-21 | (mBq) | 1.127 | +- | 0.2345 | 12.08 | +- | 3.725 | 0.1255 | +- | 0.07481 | 1.103 | +- | 0.2432 | 6.0743 | +- | 2.224 | 0.57723 | +- | 0.1685 | <0.14
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 115.11 | +- | 43.71 | 0.82866 | +- | 1.117 | <0.27 |  |  | 0.5149 | +- | 0.4248
U Calgary Sediment V29 |  | 410.6 g | 4.307 | 25042301 | 2025-04-23 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
RR1 13Ft, 58-87 |  |  |  |  | 2025-04-28 | (mBq/kg) | 33290 | +- | 632.8 | 49690 | +- | 3156 | 1392 | +- | 40.53 | 31040 | +- | 750.5 | 538040 | +- | 27520 | 1492.3 | +- | 86.68 | <34.04
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 2.7 ppm | +- | 0.05 ppm | 4.02 ppm | +- | 0.26 ppm | 2.45 ppm | +- | 0.07 ppm | 7.64 ppm | +- | 0.18 ppm | 17378.69 ppm | +- | 888.9 ppm | Interspec Correction
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac: |  | 134Cs
[6 empty] | (mBq/kg) | <81950.00 |  |  | 189.65 | +- | 271.1 | 518.66 | +- | 44.19 | 28310 | +- | 708.5 | 809.98 | +- | 64.91
[6 empty] | (ppm / ppb / ppt) | <6.64 ppm | [8 empty] | 6.96 ppm | +- | 0.17 ppm
U Calgary Sediment V30 |  | 423.5 g |  |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
RR1 13Ft, 87-109 | [5 empty] | (mBq/kg)
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
(mBq/kg)
[6 empty] | (ppm / ppb / ppt) | 0 ppm | +- | 0 ppm | [6 empty] | 0 ppm | +- | 0 ppm
U Calgary Sediment V31 |  | 395.9 g |  |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Soil Sample 1.5S | [5 empty] | (mBq/kg)
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
(mBq/kg)
[6 empty] | (ppm / ppb / ppt) | 0 ppm | +- | 0 ppm | [6 empty] | 0 ppm | +- | 0 ppm
U Calgary Sediment V32 |  | 409.0 g |  |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Soil Sample 5E5S | [5 empty] | (mBq/kg)
Sebastian Champagne | [5 empty] | (ppm / ppb / ppt) | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm | 0 ppm | +- | 0 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn |  |  | 228Ac:
(mBq/kg)
[6 empty] | (ppm / ppb / ppt) | 0 ppm | +- | 0 ppm | [6 empty] | 0 ppm | +- | 0 ppm
Next Sample | [5 empty] | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
(mBq/kg)
(ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)