## Collected Ge Detector Sample Re
Some Useful Information Concerning the results:
The Conversion factors for the primordial nuclides are given by: | [9 empty] | 1 Bq 238U/kg = 81 ppb U (81 x 10-9 gU/g) | [5 empty] | The 238U decay chain gammas used are: | [5 empty] | 226Ra: 186.1 keV
214Pb: 295.21 and 351.92 keV
214Bi: 609.31, 1120.29, 1764.49 and 2204.21 keV
The relationships are valid for any daughters in the 238U, 235U or 232Th chain only if the chain is in equilibrium. | [9 empty] | 1 Bq 232Th/kg = 246 ppb Th (246 x 10-9 gTh/g) | [5 empty] | The 232Th decay chain gammas used are: | [5 empty] | 212Pb: 238.63 and 300.09 keV
208Tl: 583.19 and 2614.53 keV,
228Ac: 911.21 keV
[10 empty] | 1 Bq 40K/kg = 32300 ppb K (32300 x 10-6 gK/g) | [5 empty] | The 40K decay chain gamma used is: | [5 empty] | 40K: 1460.83 keV
[10 empty] | 1 Bq 235U/kg = 1.76 ppm U (1.76 x 10-6 gU/g) | [5 empty] | The 235U decay chain gammas used are: | [5 empty] | 235U: 143.76, 163.33 and 205.31 keV
Background Measurements of the Vue des Alpes Detector
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Background 1 | Efficiency assumes no samples |  | 36.187 | 171213 171218 | 2017-12-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-01-19 | (mBq) | 1.873 | +- | 0.1674 | 5.039 | +- | 1.612 | 0.0612 | +- | 0.05173 | 1.354 | +- | 0.1679 | 15.262 | +- | 1.927 | 0.2008 | +- | 0.096 | 0.06328 | +- | 0.05422
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 76.664 | +- | 14.04 | <0.67 |  |  | <0.081 |  |  | 0.8836 | +- | 0.2577
Background 2 | Efficiency assumes no samples |  | 17.497 | 180212 | 2018-02-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-03-02 | (mBq) | 1.322 | +- | 0.2181 | 3.725 | +- | 2.318 | 0.2334 | +- | 0.07634 | 0.7936 | +- | 0.2178 | 18.386 | +- | 2.838 | 0.39164 | +- | 0.1591 | 0.02449 | +- | 0.08185
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.72 | +- | 20.73 | <0.98 |  |  | <0.12 |  |  | 1.033 | +- | 0.3662
Background 3 | Efficiency assumes no samples |  | 51.75 | 180319 180417 | 2018-03-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-05-12 | (mBq) | 2.225 | +- | 0.1474 | 4.95 | +- | 1.385 | 0.14 | +- | 0.0455 | 1.274 | +- | 0.1334 | 17.109 | +- | 1.737 | 0.43336 | +- | 0.09 | <0.068
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.15 | +- | 12.63 | 0.5036 | +- | 0.55 | 0.075 | +- | 0.065 | 1.212 | +- | 0.2177
Background 4 | Efficiency assumes no samples |  | 82.542 | 180522 18053001 180801 | 2018-05-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-08-12 | (mBq) | 2.053 | +- | 0.1186 | 6.239 | +- | 1.172 | 0.1237 | +- | 0.03392 | 1.328 | +- | 0.1141 | 15.52 | +- | 1.393 | 0.31949 | +- | 0.0674 | 0.1219 | +- | 0.03733
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 79.586 | +- | 9.876 | 0.731 | +- | 0.4578 | 0.0398 | +- | 0.04739 | 1.13 | +- | 0.18
Background 5 | Efficiency assumes no samples |  | 76.532 | 190910 190914 190915 | 2019-09-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2019-11-28 | (mBq) | 2.04 | +- | 0.122 | 6.373 | +- | 1.193 | 0.133 | +- | 0.03459 | 1.164 | +- | 0.1159 | 16.218 | +- | 1.452 | 0.38528 | +- | 0.0757 | 0.097 | +- | 0.0385
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 82.073 | +- | 10.15 | 0.48395 | +- | 0.443 | <0.052 |  |  | 1.477 | +- | 0.1925
Background 6 | Efficiency assumes no samples |  | 55.304 | 200324 200430 | 2020-03-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2020-05-19 | (mBq) | 3.397 | +- | 0.1763 | 7.801 | +- | 1.57 | 0.1535 | +- | 0.04259 | 2.404 | +- | 0.1694 | 29.088 | +- | 2.378 | 0.24865 | +- | 0.0871 | 0.03165 | +- | 0.05048
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 191.42 | +- | 21.95 | 0.48686 | +- | 0.5507 | 0.043088 | +- | 0.06009 | 2.092 | +- | 0.2595
Background 7 | Efficiency assumes no samples |  | 7.257 | 210111 210119 | 2021-01-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2021-01-20 | (mBq) | 2.063 | +- | 0.3767 | 8.573 | +- | 3.769 | 0.2188 | +- | 0.1077 | 1.139 | +- | 0.3709 | 29.316 | +- | 5.234 | 0.36096 | +- | 0.2337 | <0.12
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 172.52 | +- | 44.42 | <1.55 |  |  | <0.21 |  |  | 1.727 | +- | 0.6023
Combined Background | Combined Backgrounds of Runs 1+2+3+4+5 |  | 238.322 |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector | Efficiency assumes no samples |  |  |  |  | (mBq) | 2.414 | +- | 0.094 | 6.61 | +- | 0.85 | 0.333 | +- | 0.06 | 1.534 | +- | 0.082 | 18.6526 | +- | 1.2214 | 0.4053 | +- | 0.0522 | 0.151 | +- | 0.026
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 99.2049 | +- | 7.7138 | 0.3194 | +- | 0.3263 | 0.0063 | +- | 0.0342 | 1.394 | +- | 0.041
Completed Sample Measurements Counted on the Vue des Alpes Detector
The measurements of the samples below take into account the background measurement shown above. If a measurement is below the background then the upper bound shown is the 90% confidence limit.
SBC Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
SBC V01 | Stainless Steel | 364.72 g | 8.777 | 210127 210202 | 2021-01-27 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Vessel Body |  |  |  | 2021-02-05 | (mBq/kg) | 1.909 | +- | 1.127 | <61.70 |  |  | 0.4688 | +- | 0.4291 | <2.64 |  |  | <23.78 |  |  | 0.70548 | +- | 0.6779 | 5.059 | +- | 0.8153
| PV Drawing Item #3 |  |  |  |  | (ppb or ppm) | 154.63 ppt | +- | 91.29 ppt | <5 ppb |  |  | 0.83 ppb | +- | 0.76 ppb | <0.65 ppb |  |  | <0.77 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 148260 | +- | 118900 | <7.53 |  |  | 2.8376 | +- | 0.7004 | <1.76 |  |  | 4.727 | +- | 1.904 | <0.86
[6 empty] | (ppb or ppm) | 12.01 ppm | +- | 9.63 ppm | [6 empty] | <0.43 ppb
SBC V02 | Stainless Steel | 254.35 g | 6.791 | 210205 200207 | 2021-02-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Feedthrough Tube for the Pressure Vessel |  |  |  | 2021-02-12 | (mBq/kg) | 3.86 | +- | 2.028 | <67.98 |  |  | 0.3404 | +- | 0.7203 | 4.659 | +- | 2.153 | <34.73 |  |  | 1.0881 | +- | 1.447 | 1.016 | +- | 0.8182
| PV Drawing Item #8 |  |  |  |  | (ppb or ppm) | 312.66 ppt | +- | 164.27 ppt | <5.51 ppb |  |  | 0.6 ppb | +- | 1.27 ppb | 1.15 ppb | +- | 0.53 ppb | <1.12 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 141470 | +- | 129400 | 7.4193 | +- | 8.756 | 2.3993 | +- | 1.235 | <7.34 |  |  | <2.44 |  |  | <1.17
[6 empty] | (ppb or ppm) | 11.46 ppm | +- | 10.48 ppm | [6 empty] | <1.81 ppb
SBC V03 | Stainless Steel | 189.45 g | 11.541 | 210212 210216 | 2021-02-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Vessel Spool |  |  |  | 2021-02-24 | (mBq/kg) | 0.9399 | +- | 1.763 | 126.8 | +- | 71.45 | 2.376 | +- | 0.7457 | <3.89 |  |  | <32.52 |  |  | 1.5692 | +- | 1.16 | 75.03 | +- | 4.169
| PV Drawing Item #6 |  |  |  |  | (ppb or ppm) | 0.08 ppb | +- | 0.14 ppb | 10.27 ppb | +- | 5.79 ppb | 4.18 ppb | +- | 1.31 ppb | <0.96 ppb |  |  | <1.05 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 122400 | +- | 106900 | <12.46 |  |  | 6.4552 | +- | 1.469 | 2.611 | +- | 3.374 | <3.72 |  |  | <2.38
[6 empty] | (ppb or ppm) | 9.91 ppm | +- | 8.66 ppm | [6 empty] | 0.64 ppb | +- | 0.83 ppb
SBC V04 | Stainless Steel | 134.65 g | 8.639 | 210224 210301 | 2021-02-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Vessel Dome |  |  |  | 2021-03-05 | (mBq/kg) | <3.93 |  |  | <108.50 |  |  | <0.74 |  |  | 4.873 | +- | 2.867 | <42.76 |  |  | <1.27 |  |  | <1.33
| PV Drawing Item #1 |  |  |  |  | (ppb or ppm) | <0.32 ppb |  |  | <8.79 ppb |  |  | <1.3 ppb |  |  | 1.2 ppb | +- | 0.71 ppb | <1.38 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 74615 | +- | 120200 | 3.12 | +- | 8.982 | 4.1339 | +- | 1.37 | 12.3 | +- | 4.232 | 3.749 | +- | 2.117 | <1.59
[6 empty] | (ppb or ppm) | 6.04 ppm | +- | 9.74 ppm | [6 empty] | 3.03 ppb | +- | 1.04 ppb
SBC V05 | Stainless Steel | 372.40 g | 14.722 | 210305 | 2021-03-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel | Pressure Relief Pipe |  |  |  | 2021-03-22 | (mBq/kg) | 3.971 | +- | 1.466 | 82.29 | +- | 53.49 | 0.8206 | +- | 0.6175 | 6.322 | +- | 1.587 | <5.66 |  |  | <0.80 |  |  | 3.23 | +- | 0.8053
[6 empty] | (ppb or ppm) | 0.32 ppb | +- | 0.12 ppb | 6.67 ppb | +- | 4.33 ppb | 1.44 ppb | +- | 1.09 ppb | 1.56 ppb | +- | 0.39 ppb | <0.18 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 130530 | +- | 136300 | <6.77 |  |  | 1.6401 | +- | 0.7656 | <2.38 |  |  | <2.28 |  |  | 0.54838 | +- | 0.6217
[6 empty] | (ppb or ppm) | 10.57 ppm | +- | 11.04 ppm | [6 empty] | <0.59 ppb
SBC V06 | Silver Plated Bolts | 744.9 g | 7.182 | 231228 | 2023-12-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Silter Plated Bolts | Label: S30400 THE |  |  |  | 2024-01-05 | (mBq/kg) | <1.40 |  |  | 26.97 | +- | 36.14 | 0.4047 | +- | 0.3929 | 3.417 | +- | 1.103 | <9.62 |  |  | <0.50 |  |  | 23.25 | +- | 1.558
| Units: 6 |  |  |  |  | (ppb or ppm) | <0.11 ppb |  |  | 2.18 ppb | +- | 2.93 ppb | 0.71 ppb | +- | 0.69 ppb | 0.84 ppb | +- | 0.27 ppb | <0.31 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 18039 | +- | 20200 | 10.836 | +- | 5.093 | 0.8271 | +- | 0.6252 | 2.577 | +- | 1.833
[6 empty] | (ppb or ppm) | 1.46 ppm | +- | 1.64 ppm | [6 empty] | 0.63 ppb | +- | 0.45 ppb
SBC V07 | PTFE | 48.5 g | 11.708 | 240105 | 2024-01-05 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
IA PTFE inner Tower |  |  |  |  | 2024-01-17 | (mBq/kg) | <4.70 |  |  | 181.1 | +- | 111.8 | <4.56 |  |  | <5.05 |  |  | <63.58 |  |  | <11.23 |  |  | <2.16
[6 empty] | (ppb or ppm) | <0.38 ppb |  |  | 14.67 ppb | +- | 9.06 ppb | <8.03 ppb |  |  | <1.24 ppb |  |  | <2.05 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 23168 | +- | 13680 | <40.04 |  |  | <8.61 |  |  | <9.43
[6 empty] | (ppb or ppm) | 1.88 ppm | +- | 1.11 ppm | [6 empty] | <2.32 ppb
SBC V08 |  | 1695.2 g | 7.854 | 240117 | 2024-01-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Bellows Flange | 8 Pieces |  |  |  | 2024-01-25 | (mBq/kg) | 4.943 | +- | 0.8595 | 15.12 | +- | 27.78 | 0.1853 | +- | 0.3833 | 12.42 | +- | 1.2 | <7.91 |  |  | <0.59 |  |  | 10.29 | +- | 0.7999
[6 empty] | (ppb or ppm) | 0.4 ppb | +- | 0.07 ppb | 1.22 ppb | +- | 2.25 ppb | 0.33 ppb | +- | 0.67 ppb | 3.06 ppb | +- | 0.3 ppb | <0.26 ppm
| This is the standard background to be subtracted from samples beginning on May 25, 2018 |  |  |  |  | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | 14818 | +- | 21610 | 5.0672 | +- | 3.814 | 0.83976 | +- | 0.4803 | 15.84 | +- | 1.965 | <1.30 |  |  | 0.97213 | +- | 0.4483
[6 empty] | (ppb or ppm) | 1.2 ppm | +- | 1.75 ppm | [6 empty] | 3.9 ppb | +- | 0.48 ppb
SBC V09 | Standard Stainles Steel Tubes which are welded together | 150.8 g | 6.73 | 240209 240210 24021001 | 2024-02-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Stainless Steel Welded Tubes |  |  |  |  | 2024-02-16 | (mBq/kg) | <1.61 |  |  | <73.78 |  |  | <0.93 |  |  | 4.495 | +- | 2.51 | <21.85 |  |  | <1.57 |  |  | 6.462 | +- | 1.508
[6 empty] | (ppb or ppm) | <0.13 ppb |  |  | <5.98 ppb |  |  | <1.64 ppb |  |  | 1.11 ppb | +- | 0.62 ppb | <705.76 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 57Co |  |  | 58Co
[6 empty] | (mBq/kg) | <14390.00 |  |  | <12.37 |  |  | 2.8532 | +- | 1.386 | 2.074 | +- | 3.812 | <1.65 |  |  | <2.51
[6 empty] | (ppb or ppm) | <1.17 ppm | [8 empty] | 0.51 ppb | +- | 0.94 ppb
In Progress Sample Measurements Counted on the Vue des Alpes Detector
In Progress and To Be Measured:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Next Sample | [5 empty] | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
(mBq/kg)
(ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)