## Collected Ge Detector Sample Re
Some Useful Information Concerning the results:
The Conversion factors for the primordial nuclides are given by: | [9 empty] | 1 Bq 238U/kg = 81 ppb U (81 x 10-9 gU/g) | [5 empty] | The 238U decay chain gammas used are: | [5 empty] | 226Ra: 186.1 keV
214Pb: 295.21 and 351.92 keV
214Bi: 609.31, 1120.29, 1764.49 and 2204.21 keV
The relationships are valid for any daughters in the 238U, 235U or 232Th chain only if the chain is in equilibrium. | [9 empty] | 1 Bq 232Th/kg = 246 ppb Th (246 x 10-9 gTh/g) | [5 empty] | The 232Th decay chain gammas used are: | [5 empty] | 212Pb: 238.63 and 300.09 keV
208Tl: 583.19 and 2614.53 keV,
228Ac: 911.21 keV
[10 empty] | 1 Bq 40K/kg = 32300 ppb K (32300 x 10-6 gK/g) | [5 empty] | The 40K decay chain gamma used is: | [5 empty] | 40K: 1460.83 keV
[10 empty] | 1 Bq 235U/kg = 1.76 ppm U (1.76 x 10-6 gU/g) | [5 empty] | The 235U decay chain gammas used are: | [5 empty] | 235U: 143.76, 163.33 and 205.31 keV
Background Measurements of the Vue des Alpes Detector
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Background 1 | Efficiency assumes no samples |  | 36.187 | 171213 171218 | 2017-12-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-01-19 | (mBq) | 1.873 | +- | 0.1674 | 5.039 | +- | 1.612 | 0.0612 | +- | 0.05173 | 1.354 | +- | 0.1679 | 15.262 | +- | 1.927 | 0.2008 | +- | 0.096 | 0.06328 | +- | 0.05422
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 76.664 | +- | 14.04 | <0.67 |  |  | <0.081 |  |  | 0.8836 | +- | 0.2577
Background 2 | Efficiency assumes no samples |  | 17.497 | 180212 | 2018-02-12 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-03-02 | (mBq) | 1.322 | +- | 0.2181 | 3.725 | +- | 2.318 | 0.2334 | +- | 0.07634 | 0.7936 | +- | 0.2178 | 18.386 | +- | 2.838 | 0.39164 | +- | 0.1591 | 0.02449 | +- | 0.08185
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.72 | +- | 20.73 | <0.98 |  |  | <0.12 |  |  | 1.033 | +- | 0.3662
Background 3 | Efficiency assumes no samples |  | 51.75 | 180319 180417 | 2018-03-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-05-12 | (mBq) | 2.225 | +- | 0.1474 | 4.95 | +- | 1.385 | 0.14 | +- | 0.0455 | 1.274 | +- | 0.1334 | 17.109 | +- | 1.737 | 0.43336 | +- | 0.09 | <0.068
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 101.15 | +- | 12.63 | 0.5036 | +- | 0.55 | 0.075 | +- | 0.065 | 1.212 | +- | 0.2177
Background 4 | Efficiency assumes no samples |  | 82.542 | 180522 18053001 180801 | 2018-05-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2018-08-12 | (mBq) | 2.053 | +- | 0.1186 | 6.239 | +- | 1.172 | 0.1237 | +- | 0.03392 | 1.328 | +- | 0.1141 | 15.52 | +- | 1.393 | 0.31949 | +- | 0.0674 | 0.1219 | +- | 0.03733
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 79.586 | +- | 9.876 | 0.731 | +- | 0.4578 | 0.0398 | +- | 0.04739 | 1.13 | +- | 0.18
Background 5 | Efficiency assumes no samples |  | 76.532 | 190910 190914 190915 | 2019-09-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2019-11-28 | (mBq) | 2.04 | +- | 0.122 | 6.373 | +- | 1.193 | 0.133 | +- | 0.03459 | 1.164 | +- | 0.1159 | 16.218 | +- | 1.452 | 0.38528 | +- | 0.0757 | 0.097 | +- | 0.0385
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 82.073 | +- | 10.15 | 0.48395 | +- | 0.443 | <0.052 |  |  | 1.477 | +- | 0.1925
Background 6 | Efficiency assumes no samples |  | 55.304 | 200324 200430 | 2020-03-24 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2020-05-19 | (mBq) | 3.397 | +- | 0.1763 | 7.801 | +- | 1.57 | 0.1535 | +- | 0.04259 | 2.404 | +- | 0.1694 | 29.088 | +- | 2.378 | 0.24865 | +- | 0.0871 | 0.03165 | +- | 0.05048
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 191.42 | +- | 21.95 | 0.48686 | +- | 0.5507 | 0.043088 | +- | 0.06009 | 2.092 | +- | 0.2595
Background 7 | Efficiency assumes no samples |  | 7.257 | 210111 210119 | 2021-01-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector |  |  |  |  | 2021-01-20 | (mBq) | 2.063 | +- | 0.3767 | 8.573 | +- | 3.769 | 0.2188 | +- | 0.1077 | 1.139 | +- | 0.3709 | 29.316 | +- | 5.234 | 0.36096 | +- | 0.2337 | <0.12
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 172.52 | +- | 44.42 | <1.55 |  |  | <0.21 |  |  | 1.727 | +- | 0.6023
Combined Background | Combined Backgrounds of Runs 1+2+3+4+5 |  | 238.322 |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Completely Empty Detector | Efficiency assumes no samples |  |  |  |  | (mBq) | 2.414 | +- | 0.094 | 6.61 | +- | 0.85 | 0.333 | +- | 0.06 | 1.534 | +- | 0.082 | 18.6526 | +- | 1.2214 | 0.4053 | +- | 0.0522 | 0.151 | +- | 0.026
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 99.2049 | +- | 7.7138 | 0.3194 | +- | 0.3263 | 0.0063 | +- | 0.0342 | 1.394 | +- | 0.041
Completed Sample Measurements Counted on the Vue des Alpes Detector
The measurements of the samples below take into account the background measurement shown above. If a measurement is below the background then the upper bound shown is the 90% confidence limit.
SNOLAB Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
SNOLAB V01 | Background is NOT subtracted |  | 20.583 | 170901 | 2017-09-01 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Acrylic Sample Beaker H | Efficiency assumes no samples |  |  |  | 2017-09-22 | (mBq) | 1.439 | +- | 0.145 | 0.182 | +- | 0.729 | 0.354 | +- | 0.104 | 0.798 | +- | 0.114 | 7.886 | +- | 1.318 | 0.197 | +- | 0.082 | 0.04 | +- | 0.048
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 32.615 | +- | 5.256 | 0.38 | +- | 0.563 | <0.052 |  |  | 0.455 | +- | 0.199
SNOLAB V01A | Background Subtracted |  | 20.853 | 170901 | 2017-09-01 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Acrylic Sample Beaker H | Efficiency assumes no samples |  |  |  | 2017-09-22 | (mBq) | 0.294 | +- | 0.142 | <0.095 |  |  | 0.149 | +- | 0.105 | 0.007 | +- | 0.115 | <0.26 |  |  | <0.096 |  |  | 0.012 | +- | 0.05
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq) | 12.805 | +- | 5.104 | 0.38 | +- | 0.563 | <0.052 |  |  | <0.082
SNOLAB V02 | Lot No. A4B8 02/20 | 1871.5 g | 6.979 | 180511 | 2018-05-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Alconox | Alconox Inc. |  |  |  | 2018-05-18 | (mBq/kg) | 184.6 | +- | 4.642 | 104.5 | +- | 18.43 | 5.728 | +- | 0.54 | 139.5 | +- | 4.386 | 8664 | +- | 446.7 | <0.65 |  |  | 0.8054 | +- | 0.7874
[6 empty] | (ppb or ppm) | 14.95 ppb | +- | 0.38 ppb | 8.46 ppb | +- | 1.49 ppb | 10.08 ppb | +- | 0.95 ppb | 34.32 ppb | +- | 1.08 ppb | 279.85 ppm | +- | 14.43 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <79.93 |  |  | 5.6059 | +- | 5.455 | <1.21 |  |  | 61.58 | +- | 3.484
[6 empty] | (ppb or ppm) | <6.47 ppb | [8 empty] | 15.15 ppb | +- | 0.86 ppb
SNOLAB V03 | Sterling AP 308/3-8L, AWS E308-16, E308L-16 | 2294.9 g | 3.91 | 180518 | 2018-05-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Premium Welding Electrodes --- Calibration | McKay Welding Products, Troy, Ohio, Lot No 24L056, Heat No. 738228 |  |  |  | 2018-05-22 | (mBq/kg) | 32330 | +- | 564.6 | 15260 | +- | 854.5 | 1110 | +- | 24.75 | 4810 | +- | 116.1 | 261860 | +- | 13380 | <1.68 |  |  | <6.29
[6 empty] | (ppb or ppm) | 2.62 ppm | +- | 0.05 ppm | 1.24 ppm | +- | 0.07 ppm | 1.95 ppm | +- | 0.04 ppm | 1.18 ppm | +- | 0.03 ppm | 8458.08 ppm | +- | 432.17 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 414.69 | +- | 1085 |  | NM |  |  | NM |  | 4376 | +- | 120.2
[6 empty] | (ppb or ppm) | 33.59 ppb | +- | 87.89 ppb | [6 empty] | 1.08 ppm | +- | 0.03 ppm
SNOLAB V04 | Sources prepared on Nov 21, 2019 | 4.4 g | 12.638 | 191128 | 2019-11-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Calibration Source Wipes | The Kimwipes are from sources SRS-19-004 |  | Standard Background Subtraction Using Combined Backgrounds |  | 2019-12-11 | (mBq/kg) | 455.6 | +- | 79.18 | 1335 | +- | 733.5 | 12.51 | +- | 22.94 | 300.5 | +- | 76.03 | 5534.6 | +- | 1118 | <48.00 |  |  | 11.39 | +- | 31.69
[6 empty] | (ppb or ppm) | 36.9 ppb | +- | 6.41 ppb | 108.14 ppb | +- | 59.41 ppb | 22.02 ppb | +- | 40.37 ppb | 73.92 ppb | +- | 18.7 ppb | 178.77 ppm | +- | 36.11 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 986.42 | +- | 6241 | <229.50 |  |  | 28.188 | +- | 26.26 | 99.8 | +- | 116.5
[6 empty] | (ppb or ppm) | 79.9 ppb | +- | 505.52 ppb | [6 empty] | 24.55 ppb | +- | 28.66 ppb
SNOLAB V04A | Sources prepared on Nov 21, 2019 | 4.4 g | 12.638 | 191128 | 2019-11-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Calibration Source Wipes | The Kimwipes are from sources SRS-19-004 |  | Background Subtraction Using The Results from SNOLAB 05 |  | 2019-12-11 | (mBq/kg) | 8.961 | +- | 114.2 | 264.9 | +- | 989.3 | <23.57 |  |  | 18.4 | +- | 111.7 | 0.00306 | +- | 1567 | <137.40 |  |  | 0.0228 | +- | 38.3
[6 empty] | (ppb or ppm) | 0.73 ppb | +- | 9.25 ppb | 21.46 ppb | +- | 80.13 ppb | <41.48 ppb |  |  | 4.53 ppb | +- | 27.48 ppb | 0 ppm | +- | 50.61 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 0.0168 | +- | 24810 | <944.30 |  |  | 0.025 | +- | 47.83 | 6.347 | +- | 165.8
[6 empty] | (ppb or ppm) | 0 ppb | +- | 2009.61 ppb | [6 empty] | 1.56 ppb | +- | 40.79 ppb
SNOLAB V05 | Sample prepared on Dec 05, 2019 | 4.4 g | 7.882 | 191211 | 2019-12-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Kimwipes Blank | These are clean Kimwipes |  | Standard Background Subtraction Using Combined Backgrounds |  | 2019-12-19 | (mBq/kg) | 40.49 | +- | 84.56 | 801.4 | +- | 823.4 | <14.59 |  |  | 55.51 | +- | 88.79 | 635.3 | +- | 990.9 | <83.35 |  |  | <43.88
[6 empty] | (ppb or ppm) | 3.28 ppb | +- | 6.85 ppb | 64.91 ppb | +- | 66.7 ppb | <25.68 ppb |  |  | 13.66 ppb | +- | 21.84 ppb | 20.52 ppm | +- | 32.01 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 20521 | +- | 11650 | 212.32 | +- | 352.6 | 17.376 | +- | 29.66 | 115.1 | +- | 142.6
[6 empty] | (ppb or ppm) | 1662.2 ppb | +- | 943.65 ppb | [6 empty] | 28.31 ppb | +- | 35.08 ppb
SNOLAB V06 | Vial cap is downwards | 0.832 g | 17.701 | 191219 191220 | 2019-12-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-K2 nEXO Calibration Sample | Specific Activity: 14.0 +- 0.09 Bq of 40K / [g RGK-1] |  |  |  | 2020-01-06 | (mBq/kg) | 4315 | +- | 513.7 | 8816 | +- | 15340 | 146.6 | +- | 180.2 | 297.1 | +- | 420.2 | 15233000 | +- | 801100 | 615.97 | +- | 484.9 | <981.60
| [11.57 – 11.72] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 349.52 ppb | +- | 41.61 ppb | 714.1 ppb | +- | 1242.54 ppb | 258.02 ppb | +- | 317.15 ppb | 73.09 ppb | +- | 103.37 ppb | 492025.9 ppm | +- | 25875.53 ppm
[6 empty] | (Bq) | [12 empty] | 12.6739 | +- | 0.50149
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 795110 | +- | 2797000 | 1644 | +- | 3001 | 594.49 | +- | 511.2 | <1864.00
[6 empty] | (ppb or ppm) | 64.4 ppm | +- | 226.56 ppm | [6 empty] | <458.54 ppb
SNOLAB V07 | Vial cap is upwards | 2.046 g | 1.937 | 200106 | 2020-01-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-U1 nEXO Calibration Sample | Specific Activity: 4.94 +- 0.03 Bq of 238U / [g RGU-1] |  |  |  | 2020-01-08 | (mBq/kg) | 5114000 | +- | 93980 | 6168000 | +- | 377900 | 231300 | +- | 5404 | 2079 | +- | 1486 | 4783.5 |  | 28500 | <637.00 |  |  | <3318.00
| [10.04 – 10.17] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 414.23 ppm | +- | 7.61 ppm | 499.61 ppm | +- | 30.61 ppm | 407.09 ppm | +- | 9.51 ppm | 511.43 ppb | +- | 365.56 ppb | 154.51 ppm | +- | 920.55 ppm
[6 empty] | (Bq) | 10.4632 | +- | 0.19228 | 12.6197 | +- | 0.77318 |  | Average: |  | 10.5888 | +- | 0.1866
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 66658000 | +- | 9102000 | <10990.00 |  |  | <960.20 |  |  | <8204.00
[6 empty] | (ppb or ppm) | 5399.3 ppm | +- | 737.26 ppm | [6 empty] | <2.02 ppm
[6 empty] | (Bq) | 136.3823 | +- | 18.6227
SNOLAB V08 | Vial cap is upwards | 3.327 g | 5 | 200108 | 2020-01-08 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-T1 nEXO Calibration Sample | Specific Activity: 3.25 +- 0.07 Bq of 232Th / [g RGTh-1] |  |  |  | 2020-01-13 | (mBq/kg) | 50150 | +- | 2512 | 511500 | +- | 73360 | 3829 | +- | 540.7 | 2961000 | +- | 68880 | 233380 | +- | 15970 | <1441.00 |  |  | <729.70
| [10.58 – 11.05] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 4.06 ppm | +- | 0.2 ppm | 41.43 ppm | +- | 5.94 ppm | 6.74 ppm | +- | 0.95 ppm | 728.41 ppm | +- | 16.94 ppm | 7538.17 ppm | +- | 515.83 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 9.8512 | +- | 0.22916
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <1159000.00 |  |  | 23356 | +- | 7448 | 26100 | +- | 1789 | 2804000 | +- | 64000
[6 empty] | (ppb or ppm) | <93.88 ppm | [8 empty] | 689.78 ppm | +- | 15.74 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 9.3289 | +- | 0.21293 |  | Average: |  | 9.5709 | +- | 0.15599
SNOLAB V09 | Vial cap is upwards | 0.832 g | 6.896 | 200113 | 2020-01-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-K2 nEXO Calibration Sample | Specific Activity: 14.0 +- 0.09 Bq of 40K / [g RGK-1] |  |  |  | 2020-01-20 | (mBq/kg) | 494.3 | +- | 686.5 | 1890 | +- | 24090 | <437.50 |  |  | 918.8 | +- | 767.2 | 15495000 | +- | 818300 | 763.17 | +- | 785.3 | <1728.00
| [11.57 – 11.72 Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 40.04 ppb | +- | 55.61 ppb | 153.09 ppb | +- | 1951.29 ppb | <0.77 ppm |  |  | 226.02 ppb | +- | 188.73 ppb | 500488.5 ppm | +- | 26431.09 ppm
[6 empty] | (Bq) | [12 empty] | 12.8918 | +- | 0.68083
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 9895400 | +- | 4772000 | 2661.3 | +- | 4678 | 180.42 | +- | 825.5 | <1627.00
[6 empty] | (ppb or ppm) | 801.53 ppm | +- | 386.53 ppm | [6 empty] | <400.24 ppb
SNOLAB V10 | G.F. Thompson Co Ltd. | 78.7 g | 2.104 | 200120 | 2020-01-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Masters Pro-Dope with Teflon | Newmarket ON |  |  |  | 2020-01-22 | (mBq/kg) | 10880 | +- | 279.9 | 23870 | +- | 2191 | 555 | +- | 34.31 | 12220 | +- | 361.1 | 53843 | +- | 3100 | <64.63 |  |  | 20.91 | +- | 27.51
Pipe Sealant | [5 empty] | (ppb or ppm) | 881.28 ppb | +- | 22.67 ppb | 1933.47 ppb | +- | 177.47 ppb | 976.8 ppb | +- | 60.39 ppb | 3006.12 ppb | +- | 88.83 ppb | 1739.13 ppm | +- | 100.13 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 113630 | +- | 135500 | 340.15 | +- | 308.3 | 82.984 | +- | 43.57 | 13140 | +- | 422.8
[6 empty] | (ppb or ppm) | 9204.03 ppb | +- | 10975.5 ppb | [6 empty] | 3232.44 ppb | +- | 104.01 ppb
SNOLAB V11 | Vial cap is upwards | 2.102 g | 4.791 | 200122 | 2020-01-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-U2 nEXO Calibration Sample | Specific Activity: 4.94 +- 0.03 Bq of 238U / [g RGU-1] |  |  |  | 2020-01-27 | (mBq/kg) | 4911000 | +- | 88240 | 5713000 | +- | 311500 | 221400 | +- | 4801 | 2384 | +- | 993.7 | <15330.00 |  |  | <336.50 |  |  | <1583.00
| [10.32 – 10.45] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 397.79 ppm | +- | 7.15 ppm | 462.75 ppm | +- | 25.23 ppm | 389.66 ppm | +- | 8.45 ppm | 586.46 ppb | +- | 244.45 ppb | <495.16 ppm
[6 empty] | (Bq) | 10.3229 | +- | 0.18548 | 12.0087 | +- | 0.62923 |  | Average: |  | 10.4577 | +- | 0.17791
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 71438000 | +- | 6386000 | <472.00 |  |  | <461.00 |  |  | 7337 | +- | 4053
[6 empty] | (ppb or ppm) | 5786.48 ppm | +- | 517.27 ppm | [6 empty] | 1804.9 ppb | +- | 997.04 ppb
[6 empty] | (Bq) | 146.1621 | +- | 13.0658
SNOLAB V12 | Vial cap is upwards | 3.180 g | 2.958 | 200127 | 2020-01-27 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-T2 nEXO Calibration Sample | Specific Activity: 3.25 +- 0.07 Bq of 232Th / [g RGTh-1] |  |  |  | 2020-01-30 | (mBq/kg) | 50670 | +- | 2986 | 787600 | +- | 99220 | 3309 | +- | 723.5 | 3087000 | +- | 72390 | 224340 | +- | 18380 | <1361.00 |  |  | <610.70
| [10.11 – 10.56] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 4.1 ppm | +- | 0.24 ppm | 63.8 ppm | +- | 8.04 ppm | 5.82 ppm | +- | 1.27 ppm | 759.4 ppm | +- | 17.81 ppm | 7246.18 ppm | +- | 593.67 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 10.2704 | +- | 0.24084
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <453000.00 |  |  | 52273 | +- | 10520 | 27111 | +- | 2143 | 2936000 | +- | 67660
[6 empty] | (ppb or ppm) | <36.69 ppm | [8 empty] | 722.26 ppm | +- | 16.64 ppm
[6 empty] | (Bq) | [7 empty] | Sample Strength: |  | 9.7681 | +- | 0.2251 |  | Average: |  | 10.0023 | +- | 0.16446
SNOLAB V13 | Vial cap is upwards | 0.626 g | 4.91 | 200130 | 2020-01-30 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
19-004-K1 nEXO Calibration Sample | Specific Activity: 14.0 +- 0.09 Bq of 40K / [g RGK-1] |  |  |  | 2020-02-04 | (mBq/kg) | 2371 | +- | 1068 | 68810 | +- | 33160 | <311.40 |  |  | <1200.00 |  |  | 15444000 | +- | 836630 | <1641.00 |  |  | 104.3 | +- | 1429
| [ 8.70 – 8.82] Bq, 3-sigma range |  |  |  |  | (ppb or ppm) | 192.05 ppb | +- | 86.51 ppb | 5.57 ppm | +- | 2.69 ppm | <0.55 ppm |  |  | <295.2 ppb |  |  | 498841.2 ppm | +- | 27023.15 ppm
[6 empty] | (Bq) | [12 empty] | 9.6679 | +- | 0.52373
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <562400.00 |  |  | <928.20 |  |  | <933.30 |  |  | <4932.00
[6 empty] | (ppb or ppm) | <45.55 ppm | [8 empty] | <1213.27 ppb
SNOLAB V14 | Part No: FG-p | 63.9 g | 6.021 | 200204 | 2020-02-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Plastic Process Equipment Inc. Food Grade Anti-Seize Thread Compound | Macedonia, OH, 44056 |  |  |  | 2020-02-10 | (mBq/kg) | 1519 | +- | 48.29 | 1630 | +- | 268.8 | 75.37 | +- | 7.183 | 601.7 | +- | 30.42 | 189.63 | +- | 118.3 | 7.4677 | +- | 8.364 | <4.27
Pipe Sealant Fortified with PTFE | [5 empty] | (ppb or ppm) | 123.04 ppb | +- | 3.91 ppb | 132.03 ppb | +- | 21.77 ppb | 132.65 ppb | +- | 12.64 ppb | 148.02 ppb | +- | 7.48 ppb | 6.13 ppm | +- | 3.82 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 46194 | +- | 25930 | <70.64 |  |  | <9.71 |  |  | 430.3 | +- | 38.22
[6 empty] | (ppb or ppm) | 3741.71 ppb | +- | 2100.33 ppb | [6 empty] | 105.85 ppb | +- | 9.4 ppb
SNOLAB V15 | Type: 801-16-BLU | 357.4 g | 2.819 | 200210 | 2020-02-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Push Lok Plus Muli-Purpose Hose | Parker Hannifin Corp |  |  |  | 2020-02-13 | (mBq/kg) | 1803 | +- | 51.43 | 5141 | +- | 534.4 | 106 | +- | 8.728 | 3102 | +- | 91.69 | 9333.1 | +- | 568.8 | 5.009 | +- | 9.986 | <5.85
[6 empty] | (ppb or ppm) | 146.04 ppb | +- | 4.17 ppb | 416.42 ppb | +- | 43.29 ppb | 186.56 ppb | +- | 15.36 ppb | 763.09 ppb | +- | 22.56 ppb | 301.46 ppm | +- | 18.37 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 5826 | +- | 7988 | <124.20 |  |  | 22.169 | +- | 10.2 | 3004 | +- | 101.9
[6 empty] | (ppb or ppm) | 471.91 ppb | +- | 647.03 ppb | [6 empty] | 738.98 ppb | +- | 25.07 ppb
SNOLAB V16 | ¾” Natural polypropylene | 623.5 g | 7.646 | 200213 200214 | 2020-02-13 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
ENPURE Polypropylene Tubing | Type: SCH-80 Natural |  |  |  | 2020-02-21 | (mBq/kg) | 6.106 | +- | 1.706 | 17.12 | +- | 23.48 | <1.34 |  |  | 12.55 | +- | 1.969 | 4.02 | +- | 15.06 | <1.20 |  |  | 1.021 | +- | 0.5066
| Lot: JJ05-B-12021321352 |  |  |  |  | (ppb or ppm) | 0.49 ppb | +- | 0.14 ppb | 1.39 ppb | +- | 1.9 ppb | <2.36 ppb |  |  | 3.09 ppb | +- | 0.48 ppb | 129.85 ppb | +- | 486.44 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1062.5 | +- | 569.2 | <6.47 |  |  | 0.87879 | +- | 0.6374 | <3.34
[6 empty] | (ppb or ppm) | 86.06 ppb | +- | 46.11 ppb | [6 empty] | <0.82 ppb
SNOLAB V17 | Clear Non-Toxic PVC with Polyester Braid and Rigid PVC Helix | 240.5 g | 3.847 | 200221 | 2020-02-21 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Reinforced FDA PVC Suction Hose |  |  |  |  | 2020-02-25 | (mBq/kg) | 494.3 | +- | 19.11 | 739.3 | +- | 127.2 | 26.56 | +- | 3.282 | 60.84 | +- | 8.144 | 678.07 | +- | 100.8 | <3.93 |  |  | <3.08
[6 empty] | (ppb or ppm) | 40.04 ppb | +- | 1.55 ppb | 59.88 ppb | +- | 10.3 ppb | 46.75 ppb | +- | 5.78 ppb | 14.97 ppb | +- | 2 ppb | 21.9 ppm | +- | 3.26 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2178.2 | +- | 1735 | 60.512 | +- | 28.93 | <2.38 |  |  | 48.45 | +- | 12.81
[6 empty] | (ppb or ppm) | 176.43 ppb | +- | 140.54 ppb | [6 empty] | 11.92 ppb | +- | 3.15 ppb
SNOLAB V18 | NIST Sample | 28.825 g | 12.736 | 200225 200302 | 2020-02-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Powdered Ocean Sediment Calibration Sample | Cross Comparison Sample A |  |  |  | 2020-03-09 | (mBq/kg) | 11640 | +- | 255.7 | 21440 | +- | 2071 | 588.4 | +- | 23 | 12650 | +- | 331.7 | 255240 | +- | 13160 | 7109 | +- | 366.4 | <29.11
| Sample placed directly on top of the detector |  |  |  |  | (ppb or ppm) | 942.84 ppb | +- | 20.71 ppb | 1736.64 ppb | +- | 167.75 ppb | 1035.58 ppb | +- | 40.48 ppb | 3.11 ppm | +- | 0.08 ppm | 8244.25 ppm | +- | 425.07 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 241Am
[6 empty] | (mBq/kg) | <15060.00 |  |  | <452.60 |  |  | 91.497 | +- | 29.67 | 12730 | +- | 352 | 20580 | +- | 1569
[6 empty] | (ppb or ppm) | <1219.86 ppb | [8 empty] | 3.13 ppm | +- | 0.09 ppm
SNOLAB V19 | NIST Sample | 33.454 g | 21.617 | 200519 | 2020-05-19 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Powdered Ocean Sediment Calibration Sample | Cross Comparison Sample B |  |  |  | 2020-06-10 | (mBq/kg) | 10980 | +- | 223 | 17630 | +- | 1702 | 571.6 | +- | 18.5 | 12060 | +- | 299.1 | 247900 | +- | 12720 | 6397.8 | +- | 352.7 | <25.82
| Sample placed directly on top of the detector |  |  |  |  | (ppb or ppm) | 889.38 ppb | +- | 18.06 ppb | 1428.03 ppb | +- | 137.86 ppb | 1006.02 ppb | +- | 32.56 ppb | 2.97 ppm | +- | 0.07 ppm | 8007.17 ppm | +- | 410.86 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac: |  |  | 241Am
[6 empty] | (mBq/kg) | <8362.00 |  |  | <272.30 |  |  | 74.544 | +- | 20.54 | 12000 | +- | 307.1 | 28410 | +- | 1509
[6 empty] | (ppb or ppm) | <677.32 ppb | [8 empty] | 2.95 ppm | +- | 0.08 ppm
SNOLAB V20 | Lake in City of Greater Sudbury, Ontario | 993.9 g | 14.666 | 200309 | 2020-03-09 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Long Lake Water | Sample collected on 2019-10-14 at 16:54 |  |  |  | 2020-03-24 | (mBq/kg) | 2.151 | +- | 0.5584 | <12.40 |  |  | <0.27 |  |  | <0.61 |  |  | 52.142 | +- | 8.263 | 1.003 | +- | 0.383 | <0.15
| Sample placed in beaker “∆” |  |  |  |  | (ppb or ppm) | 174.23 ppt | +- | 45.23 ppt | <1 ppb |  |  | <0.48 ppb |  |  | <0.15 ppb |  |  | 1684.19 ppb | +- | 266.89 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1151.5 | +- | 456.7 | 6.8076 | +- | 2.466 | <0.38 |  |  | 1.243 | +- | 0.8064
[6 empty] | (ppb or ppm) | 93.27 ppb | +- | 36.99 ppb | [6 empty] | 0.31 ppb | +- | 0.2 ppb
SNOLAB V21 | Lake in City of Greater Sudbury, Ontario | 981.4 g | 17.117 | 200610 200624 | 2020-06-10 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Simon Lake Water | Sample collected on 2019-11-01 at 18:15 |  |  |  | 2020-06-30 | (mBq/kg) | 1.599 | +- | 0.506 | 12.4 | +- | 13.22 | <0.44 |  |  | <0.82 |  |  | 402 | +- | 25.3 | <0.46 |  |  | 0.1133 | +- | 0.2314
| Sample placed in beaker “H” |  |  |  |  | (ppb or ppm) | 129.52 ppt | +- | 40.99 ppt | 1004.4 ppt | +- | 1070.82 ppt | <0.77 ppb |  |  | <0.2 ppb |  |  | 12984.6 ppb | +- | 817.19 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1144.6 | +- | 434.2 | <2.80 |  |  | <0.27 |  |  | 4.993 | +- | 1.066
[6 empty] | (ppb or ppm) | 92.71 ppb | +- | 35.17 ppb | [6 empty] | 1.23 ppb | +- | 0.26 ppb
SNOLAB V22 | Lake in City of Greater Sudbury, Ontario | 935.2 g | 21.708 | 200630 | 2020-06-30 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Robinson Lake Water | Sample collected on 2019-10-22 at 12:43 |  |  |  | 2020-07-22 | (mBq/kg) | 5.117 | +- | 0.5798 | 21.36 | +- | 12.7 | <0.33 |  |  | 0.3155 | +- | 0.4927 | 95.614 | +- | 9.55 | 0.23 | +- | 0.2964 | <0.14
| Sample placed in beaker “□” |  |  |  |  | (ppb or ppm) | 414.48 ppt | +- | 46.96 ppt | 1.73 ppb | +- | 1.03 ppb | <0.58 ppb |  |  | 0.08 ppb | +- | 0.12 ppb | 3088.33 ppb | +- | 308.47 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 536.35 | +- | 394.1 | 0.5047 | +- | 2.041 | 0.23 | +- | 0.2964 | 1.943 | +- | 0.7858
[6 empty] | (ppb or ppm) | 43.44 ppb | +- | 31.92 ppb | [6 empty] | 0.48 ppb | +- | 0.19 ppb
SNOLAB V23 | Lake in City of Greater Sudbury, Ontario | 986.0 g | 16.659 | 200722 | 2020-07-22 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Meatbird Lake Water | Sample collected on 2019-10-08 at 10:50 |  |  |  | 2020-08-11 | (mBq/kg) | 0.19 | +- | 0.4465 | <13.83 |  |  | 0.3666 | +- | 0.2156 | <0.32 |  |  | 124.59 | +- | 11.09 | <0.39 |  |  | <0.21
| Sample placed in beaker “∆” |  |  |  |  | (ppb or ppm) | 15.39 ppt | +- | 36.17 ppt | <1.12 ppb |  |  | 0.65 ppb | +- | 0.38 ppb | <0.08 ppb |  |  | 4024.26 ppb | +- | 358.21 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <567.10 |  |  | 0.975 | +- | 1.909 | 0.02689 | +- | 0.1934 | 0.792 | +- | 0.7894
[6 empty] | (ppb or ppm) | <45.94 ppb | [8 empty] | 0.19 ppb | +- | 0.19 ppb
SNOLAB V24 | Lake in City of Greater Sudbury, Ontario | 842.7 g | 13.5 | 200811 20081101 | 2020-08-11 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Kelly Lake Water | Sample collected on 2019-10-7 at 10:30 |  |  |  | 2020-08-25 | (mBq/kg) | 3.308 | +- | 0.7805 | 11.09 | +- | 30.51 | <0.60 |  |  | 0.3551 | +- | 0.7696 | 501.76 | +- | 33.74 | 0.5963 | +- | 0.4861 | <0.24
[6 empty] | (ppb or ppm) | 267.95 ppt | +- | 63.22 ppt | 898.29 ppt | +- | 2471.31 ppt | <1.06 ppb |  |  | 0.09 ppb | +- | 0.19 ppb | 16.21 ppm | +- | 1.09 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 2300.5 | +- | 1355 | 2.4308 | +- | 3.233 | <0.43 |  |  | 6.605 | +- | 1.395
[6 empty] | (ppb or ppm) | 186.34 ppb | +- | 109.76 ppb | [6 empty] | 1.62 ppb | +- | 0.34 ppb
SNOLAB V25 | Lake in City of Greater Sudbury, Ontario | 1032.2 g | 13.395 | 201006 | 2020-10-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Robinson Lake Water | Sample collected on 2020-09-19 at 14:30 |  |  |  | 2020-10-20 | (mBq/kg) | 1.353 | +- | 0.5589 | <29.58 |  |  | 0.417 | +- | 0.2379 | <1.04 |  |  | 100.12 | +- | 11.32 | 0.4959 | +- | 0.3941 | <0.16
[6 empty] | (ppb or ppm) | 109.59 ppt | +- | 45.27 ppt | <2.4 ppb |  |  | 0.73 ppb | +- | 0.42 ppb | <0.26 ppb |  |  | 3233.88 ppb | +- | 365.64 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 881.22 | +- | 989.9 | 10.405 | +- | 3.225 | 0.33892 | +- | 0.2337 | 1.877 | +- | 0.9873
[6 empty] | (ppb or ppm) | 71.38 ppb | +- | 80.18 ppb | [6 empty] | 0.46 ppb | +- | 0.24 ppb
SNOLAB V26 | Lake in City of Greater Sudbury, Ontario | 985.3 g | 10.743 | 201023 | 2020-10-23 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
St. Charles Lake Water | Sample collected on 2020-09-19 at 14:53 |  |  |  | 2020-11-03 | (mBq/kg) | 1.5 | +- | 0.62 | 27.39 | +- | 23.77 | 0.279 | +- | 0.2828 | 0.3211 | +- | 0.6553 | 56.582 | +- | 9.839 | 1.2519 | +- | 0.4744 | 0.1718 | +- | 0.2352
[6 empty] | (ppb or ppm) | 0.12 ppb | +- | 0.05 ppb | 2.22 ppb | +- | 1.93 ppb | 0.49 ppb | +- | 0.5 ppb | 0.08 ppb | +- | 0.16 ppb | 1.83 ppm | +- | 0.32 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 23258 | +- | 29210 | 13.734 | +- | 3.761 | <0.24 |  |  | 0.159 | +- | 0.9299
[6 empty] | (ppb or ppm) | 1883.9 ppb | +- | 2366.01 ppb | [6 empty] | 0.04 ppb | +- | 0.23 ppb
SNOLAB V27 | Lake in City of Greater Sudbury, Ontario | 956.4 g | 11.146 | 201106 | 2020-11-06 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Long Lake Water | Sample collected on 2020-09-27 at 15:00 |  |  |  | 2020-11-17 | (mBq/kg) | <0.98 |  |  | <16.85 |  |  | 0.4023 | +- | 0.2632 | <0.63 |  |  | 35.529 | +- | 8.902 | 0.61398 | +- | 0.4425 | <0.31
[6 empty] | (ppb or ppm) | <79.38 ppt |  |  | <1.36 ppb |  |  | 0.71 ppb | +- | 0.46 ppb | <0.15 ppb |  |  | 1147.59 ppb | +- | 287.53 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 1542.4 | +- | 1182 | 14.09 | +- | 3.542 | <0.39 |  |  | <0.84
[6 empty] | (ppb or ppm) | 124.93 ppb | +- | 95.74 ppb | [6 empty] | <0.21 ppb
SNOLAB V28 | Lake Bed Sand from Lohi Lake in City of Greater Sudbury, Ontario | 600.4 g | 2.93 | 201020 | 2020-10-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Lohi Lake Sand | Sample collected on 2020-09-19 at 17:10 |  |  |  | 2020-10-23 | (mBq/kg) | 7917 | +- | 172.9 | 17980 | +- | 1656 | 422.1 | +- | 18.84 | 6592 | +- | 179.8 | 381970 | +- | 19550 | N/D – No observable signal due to U & Th Interference | +- | 508.4 | <27.73
[6 empty] | (ppb or ppm) | 641.28 ppb | +- | 14 ppb | 1456.38 ppb | +- | 134.14 ppb | 742.9 ppb | +- | 33.16 ppb | 1621.63 ppb | +- | 44.23 ppb | 12337.63 ppm | +- | 631.47 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 455740 | +- | 198800 | N/D – No observable signal due to U & Th Interference | +- | 710.5 | N/D – No observable signal due to U & Th Interference | +- | 18.84 | 6194 | +- | 190.9
[6 empty] | (ppb or ppm) | 36914.94 ppb | +- | 16102.8 ppb | [6 empty] | 1523.72 ppb | +- | 46.96 ppb
SNOLAB V29 | Lake Bed Sand from Ramsey Lake in City of Greater Sudbury, Ontario | 605.1 g | 2.93 | 201103 | 2020-11-03 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Ramsey Lake Sand | Sample collected on 2020-09-15 at 18:50 |  |  |  | 2020-11-06 | (mBq/kg) | 14950 | +- | 305.7 | 28000 | +- | 2515 | 765.4 | +- | 27.2 | 16230 | +- | 407.2 | 513760 | +- | 26280 | N/D – No observable signal due to U & Th Interference |  |  | <34.32
[6 empty] | (ppb or ppm) | 1210.95 ppb | +- | 24.76 ppb | 2268 ppb | +- | 203.72 ppb | 1347.1 ppb | +- | 47.87 ppb | 3992.58 ppb | +- | 100.17 ppb | 16594.45 ppm | +- | 848.84 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 156490 | +- | 184600 | N/D – No observable signal due to U & Th Interference |  |  | N/D – No observable signal due to U & Th Interference | +- | 24.88 | 14800 | +- | 392.3
[6 empty] | (ppb or ppm) | 12675.69 ppb | +- | 14952.6 ppb | [6 empty] | 3640.8 ppb | +- | 96.51 ppb
SNOLAB V30 | Lake Bed Sand from St Charles Lake in City of Greater Sudbury, Ontario | 591.6 g | 2.534 | 201117 | 2020-11-17 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
St Charles Lake Sand | Sample collected on 2020-09-19 at 14:55 |  |  |  | 2020-11-20 | (mBq/kg) | 12370 | +- | 263.6 | 29950 | +- | 2642 | 606.5 | +- | 25.82 | 16700 | +- | 422.7 | 572120 | +- | 29270 | N/D – No observable signal due to U & Th Interference |  |  | 17.15 | +- | 25.66
[6 empty] | (ppb or ppm) | 1001.97 ppb | +- | 21.35 ppb | 2425.95 ppb | +- | 214 ppb | 1067.44 ppb | +- | 45.44 ppb | 4108.2 ppb | +- | 103.98 ppb | 18479.48 ppm | +- | 945.42 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <266100.00 |  |  | N/D – No observable signal due to U & Th Interference |  |  | N/D – No observable signal due to U & Th Interference | +- | 30.21 | 16120 | +- | 429.6
[6 empty] | (ppb or ppm) | <21.55 ppm | [8 empty] | 3965.52 ppb | +- | 105.68 ppb
SNOLAB V31 | Lake in City of Greater Sudbury, Ontario | 1023.10 g | 13.687 | 201120 | 2020-11-20 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Vermillion Lake Water | Sample collected on 2020-10-23 at 13:00 |  |  |  | 2020-12-04 | (mBq/kg) | 2.203 | +- | 0.5796 | 15.37 | +- | 20.24 | <0.42 |  |  | 0.9665 | +- | 0.6031 | 63.512 | +- | 9.205 | <0.38 |  |  | <0.22
[6 empty] | (ppb or ppm) | 178.44 ppt | +- | 46.95 ppt | 1.24 ppb | +- | 1.64 ppb | <0.74 ppb |  |  | 0.24 ppb | +- | 0.15 ppb | 2051.44 ppb | +- | 297.32 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <1381.00 |  |  | 57.033 | +- | 5.721 | <0.24 |  |  | 3.169 | +- | 0.9504
[6 empty] | (ppb or ppm) | <111.86 ppb | [8 empty] | 0.78 ppb | +- | 0.23 ppb
SNOLAB V32 | Lake in City of Greater Sudbury, Ontario | 961.0 g | 13.777 | 201204 | 2020-12-04 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Meatbird Lake Water | Sample collected on October 05, 2020 @ 14:00 |  |  |  | 2020-12-18 | (mBq/kg) | 0.8306 | +- | 0.5845 | <24.71 |  |  | <0.30 |  |  | 0.064 | +- | 0.5975 | 167.44 | +- | 15.2 | 0.18371 | +- | 0.3529 | 0.2729 | +- | 0.2459
[6 empty] | (ppb or ppm) | 67.28 ppt | +- | 47.34 ppt | <2 ppb |  |  | <0.53 ppb |  |  | 0.02 ppb | +- | 0.15 ppb | 5.41 ppm | +- | 0.49 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 977.51 | +- | 1067 | 5.128 | +- | 2.804 | 0.29759 | +- | 0.2745 | 4.56 | +- | 1.18
[6 empty] | (ppb or ppm) | 79.18 ppb | +- | 86.43 ppb | [6 empty] | 1.12 ppb | +- | 0.29 ppb
SNOLAB V33 | Lucas Cell Coating | 68.3 g | 23.652 | 201218 | 2020-12-18 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
ZnS(Ag) Powder | Saint-Gobain |  |  | The sample is also counted in the Well detector as SNOLAB CW10 | 2021-01-11 | (mBq/kg) | <3.30 |  |  | <98.38 |  |  | 0.8626 | +- | 0.9022 | <4.37 |  |  | 21.744 | +- | 34.92 | <3.09 |  |  | <0.89
[6 empty] | (ppb or ppm) | <0.27 ppb |  |  | <7.97 ppb |  |  | 1.52 ppb | +- | 1.59 ppb | <1.08 ppb |  |  | 702.33 ppb | +- | 1127.92 ppb
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 53891 | +- | 34430 | 5.0645 | +- | 11.94 | <1.33 |  |  | <5.33
[6 empty] | (ppb or ppm) | 4.37 ppm | +- | 2.79 ppm | [6 empty] | <1.31 ppb
SNOLAB V34 | 72% Cocoa | 202.9 g | 4.98 | 240425 | 2024-04-25 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Merci (Stork) Dark Chocolate | Made in Germany |  |  |  | 2024-04-30 | (mBq/kg) | 453.3 | +- | 23.5 | 952.5 | +- | 379.5 | 5.613 | +- | 6.602 | 285.7 | +- | 21.85 | 20507 | +- | 10540 | 62.385 | +- | 14.87 | <19.40
[6 empty] | (ppb or ppm) | 36.72 ppb | +- | 1.9 ppb | 77.15 ppb | +- | 30.74 ppb | 9.88 ppb | +- | 11.62 ppb | 70.28 ppb | +- | 5.38 ppb | 662.38 ppm | +- | 340.44 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <29620.00 |  |  | 227.13 | +- | 95.34 | 22.85 | +- | 14.15 | 630.1 | +- | 51.65
[6 empty] | (ppb or ppm) | <2.4 ppm | [8 empty] | 155 ppb | +- | 12.71 ppb
SNOLAB V35 | Medium Grey | 107.4 g | 6.894 | 240430 | 2024-04-30 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Sikafloor 22NA PurCem | Type: SS-0721 |  |  |  | 2024-05-07 | (mBq/kg) | 5466 | +- | 126.7 | 15130 | +- | 1128 | 251.3 | +- | 12.68 | 6028 | +- | 165.6 | 347280 | +- | 17800 | 101.34 | +- | 22.99 | <29.50
[6 empty] | (ppb or ppm) | 442.75 ppb | +- | 10.26 ppb | 1225.53 ppb | +- | 91.37 ppb | 442.29 ppb | +- | 22.32 ppb | 1.48 ppm | +- | 0.04 ppm | 11217.14 ppm | +- | 574.94 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | 43451 | +- | 41950 | <207.70 |  |  | 52.065 | +- | 22.07 | 5739 | +- | 173.9
[6 empty] | (ppb or ppm) | 3519.53 ppb | +- | 3397.95 ppb | [6 empty] | 1.41 ppm | +- | 0.04 ppm
SNOLAB V36 |  | 200.5 g | 2.934 | 240507 | 2024-05-07 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
Cadbury Dark Chocolate | Made in Canada |  |  |  | 2024-05-10 | (mBq/kg) | 373.7 | +- | 23.06 | 809.9 | +- | 372.5 | 14.04 | +- | 6.638 | 194.7 | +- | 20.7 | 105060 | +- | 5483 | 39.007 | +- | 14.36 | 7.814 | +- | 15.99
[6 empty] | (ppb or ppm) | 30.27 ppb | +- | 1.87 ppb | 65.6 ppb | +- | 30.17 ppb | 24.71 ppb | +- | 11.68 ppb | 47.9 ppb | +- | 5.09 ppb | 3393.44 ppm | +- | 177.1 ppm
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
[6 empty] | (mBq/kg) | <33080.00 |  |  | <157.70 |  |  | <27.03 |  |  | 370.4 | +- | 47.45
[6 empty] | (ppb or ppm) | <2.68 ppm | [8 empty] | 91.12 ppb | +- | 11.67 ppb
In Progress Sample Measurements Counted on the Vue des Alpes Detector
In Progress and To Be Measured:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Run Numbers | Counting Dates (if applicable)
Next Sample | [5 empty] | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  |  | 232Th |  |  | 40K |  |  | 137Cs |  |  | 60Co |  | Comments
(mBq/kg)
(ppb or ppm)
[6 empty] | Results: | 210Pb: |  |  |  | 7Be: |  |  | 54Mn: |  |  | 228Ac:
(mBq/kg)
(ppb or ppm)