## Collected Ge Detector Sample Re
Some Useful Information Concerning the results:
The Conversion factors for the primordial nuclides are given by: | [8 empty] | 1 Bq 238U/kg = 81 ppb U (81 x 10-9 gU/g) |  |  |  |  | The 238U decay chain gammas used are: |  |  |  | 226Ra: 186.1 keV
214Pb: 295.21 and 351.92 keV
214Bi: 609.31, 1120.29, 1764.49 and 2204.21 keV
The relationships are valid for any daughters in the 238U, 235U or 232Th chain only if the chain is in equilibrium. | [8 empty] | 1 Bq 232Th/kg = 246 ppb Th (246 x 10-9 gTh/g) |  |  |  |  | The 232Th decay chain gammas used are: |  |  |  | 212Pb: 238.63 and 300.09 keV
208Tl: 583.19 and 2614.53 keV,
228Ac: 911.21 keV
[9 empty] | 1 Bq 40K/kg = 32300 ppb K (32300 x 10-6 gK/g) |  |  |  |  | The 40K decay chain gamma used is: |  |  |  | 40K: 1460.83 keV
[9 empty] | 1 Bq 235U/kg = 1.76 ppm U (1.76 x 10-6 gU/g) |  |  |  |  | The 235U decay chain gammas used are: |  |  |  | 235U: 143.76, 163.33 and 205.31 keV
The measurements of the samples below take into account the background measurements. If a measurement is below the background then the upper bound shown is the 90% confidence limit.
DarT Measurements:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Counting Dates (if applicable)
DArT P01 | McMaster-Carr | 320.0 g | 12.812 | 2024-03-15 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  | 232Th |  | 40K |  | 137Cs |  | 60Co | Comments
Acrylic Tube | PN: 8486K642 |  | Run: 230315 23031801 | 2024-03-28 | (mBq/kg) | <0.80 |  |  | 28.99 | +- | 32.9 | <0.70 |  | <2.78 |  | <21.67 |  | <0.46 |  | <0.25
| 3.3” ID |  |  |  | (ppm / ppb / ppt) | <64.8 ppt |  |  | 2.35 ppb | +- | 2.66 ppb | <1.23 ppb |  | <0.68 ppb |  | <0.7 ppm
[5 empty] | Results: |  | 7Be: |  |  | 54Mn |  |  |  |  | 228Ac:
[5 empty] | (mBq/kg) | <4.22 |  |  | 0.71333 | +- | 0.4997 |  |  | <4.16
[5 empty] | (ppm / ppb / ppt) | [8 empty] | <1.02 ppb
DArT P02 | McMaster-Carr | 595.0 g | 13.736 | 2024-03-28 | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  | 232Th |  | 40K |  | 137Cs |  | 60Co | Comments
Acrylic Tube | PN; 8486K635 |  | Run: 240328 | 2024-04-11 | (mBq/kg) | 0.5838 | +- | 0.6587 | <28.65 |  |  | <0.46 |  | <1.74 |  | <10.27 |  | <0.96 |  | <0.26
| ID: 3.75” |  |  |  | (ppm / ppb / ppt) | 47.29 ppt | +- | 53.35 ppt | <2.32 ppb |  |  | <0.81 ppb |  | <0.43 ppb |  | <331.72 ppb
[5 empty] | Results: |  | 7Be: |  |  | 54Mn |  |  |  |  | 228Ac:
[5 empty] | (mBq/kg) | <3.93 |  |  | <0.30 |  |  |  |  | <1.72
[5 empty] | (ppm / ppb / ppt) | [8 empty] | <0.42 ppb
In Progress and To Be Measured:
Sample Description | Manufacturer | Mass (g) | Live Time (days) | Counting Dates (if applicable)
Next Sample |  |  |  |  | Results: |  | 238U from 226Ra |  |  | 238U from 234Th |  |  | 235U |  | 232Th |  | 40K |  | 137Cs |  | 60Co | Comments
(mBq/kg)
(ppm / ppb / ppt)
[5 empty] | Results: |  | 7Be: |  |  | 54Mn |  |  |  |  | 228Ac:
(mBq/kg)
(ppm / ppb / ppt)